## Sheet1
| Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8175473&pc=GAM | 2210968-2018-77872 |  | Malfunction | ETHICON INC. | 2018-12-18 05:00 | GAM | VICRYL CT UD 8X27IN 0 S/A CT-1 CR | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE OR STAFF OPENED THE STERILE SUTURE PACKET AND A DIFFERENT TYPE OF SUTURE AND NEEDLE WAS INSIDE. SUTURE FROM A DIFFERENT LOT WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 01/15/2019. INVESTIGATION SUMMARY: IT WAS REPORTED ASSEMBLY INCORRECT COMPONENT. THREE EMPTY OPENED FOILS AND A LABELED WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8 OF PRODUCT CODE JJ41, LOT # MGZ466 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF WINDING FORMER, IT WAS OBSERVED EIGHT NEEDLE/SUTURE COMBINATIONS THAT CONTAINED A VICRYL UNDAYED SUTURE EACH NEEDLE, WHICH IS CORRECT TO PRODUCT CODE JJ41. IN ADDITION, THE SWAGE AND ATTACHMENT AREA WERE NOTED AS EXPECTED. THE SUTURES WERE EXAMINED AND NO DEFECTS OR DAMAGED WERE OBSERVED. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, NO ASSEMBLY INCORRECT COMPONENT WERE OBSERVED. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND IT WAS OBSERVED A MISMATCHED IN ONE SAMPLE BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE FIVE SAMPLES WERE EXAMINED AND NO INCORRECT COMPONENT WERE OBSERVED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8175475&pc=GAM | 2210968-2018-77870 |  | Malfunction | ETHICON INC. | 2018-12-18 05:00 | GAM | VICRYL CT UD 8X27IN 0 S/A CT-1 CR | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE OR STAFF OPENED THE STERILE SUTURE PACKET AND A DIFFERENT TYPE OF SUTURE AND NEEDLE WAS INSIDE. SUTURE FROM A DIFFERENT LOT WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 01/15/2019. INVESTIGATION SUMMARY: IT WAS REPORTED ASSEMBLY INCORRECT COMPONENT. AN EMPTY OPENED FOIL OF PRODUCT CODE JJ41, LOT # MGZ466 AND A LABELED WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8 OF PRODUCT CODE J774 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, A MISMATCHED BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF ASSEMBLY INCORRECT COMPONENT IS MIXED PRODUCT CODE. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND IT WAS OBSERVED A MISMATCHED IN ONE SAMPLE BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE FIVE SAMPLES WERE EXAMINED AND NO INCORRECT COMPONENT WERE OBSERVED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). (B)(4) INCORRECT COMPONENT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8175510&pc=GAM | 2210968-2018-77866 |  | Malfunction | ETHICON INC. | 2018-12-18 05:00 | GAM | VICRYL CT UD 8X27IN 0 S/A CT-1 CR | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE OR STAFF OPENED THE STERILE SUTURE PACKET AND A DIFFERENT TYPE OF SUTURE AND NEEDLE WAS INSIDE. SUTURE FROM A DIFFERENT LOT WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 01/15/2019. INVESTIGATION SUMMARY: IT WAS REPORTED ASSEMBLY INCORRECT COMPONENT. AN EMPTY OPENED FOIL OF PRODUCT CODE JJ41, LOT # MGZ466 AND A LABELED WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8 OF PRODUCT CODE J774 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, A MISMATCHED BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF ASSEMBLY INCORRECT COMPONENT IS MIXED PRODUCT CODE. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND IT WAS OBSERVED A MISMATCHED IN ONE SAMPLE BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE FIVE SAMPLES WERE EXAMINED AND NO INCORRECT COMPONENT WERE OBSERVED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8175514&pc=GAM | 2210968-2018-77868 |  | Malfunction | ETHICON INC. | 2018-12-18 05:00 | GAM | VICRYL CT UD 8X27IN 0 S/A CT-1 CR | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE OR STAFF OPENED THE STERILE SUTURE PACKET AND A DIFFERENT TYPE OF SUTURE AND NEEDLE WAS INSIDE. SUTURE FROM A DIFFERENT LOT WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 01/15/2019. INVESTIGATION SUMMARY IT WAS REPORTED ASSEMBLY INCORRECT COMPONENT. AN EMPTY OPENED FOIL OF PRODUCT CODE JJ41, LOT # MGZ466 AND A LABELED WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8 OF PRODUCT CODE J774 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, A MISMATCHED BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF ASSEMBLY INCORRECT COMPONENT IS MIXED PRODUCT CODE. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND IT WAS OBSERVED A MISMATCHED IN ONE SAMPLE BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE FIVE SAMPLES WERE EXAMINED AND NO INCORRECT COMPONENT WERE OBSERVED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8175515&pc=GAM | 2210968-2018-77867 |  | Malfunction | ETHICON INC. | 2018-12-18 05:00 | GAM | VICRYL CT UD 8X27IN 0 S/A CT-1 CR | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE OR STAFF OPENED THE STERILE SUTURE PACKET AND A DIFFERENT TYPE OF SUTURE AND NEEDLE WAS INSIDE. SUTURE FROM A DIFFERENT LOT WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: DATE SENT TO THE FDA: 01/15/2019. INVESTIGATION SUMMARY: IT WAS REPORTED ASSEMBLY INCORRECT COMPONENT. AN EMPTY OPENED FOIL OF PRODUCT CODE JJ41, LOT # MGZ466 AND A LABELED WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8 OF PRODUCT CODE J774 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, A MISMATCHED BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF ASSEMBLY INCORRECT COMPONENT IS MIXED PRODUCT CODE. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND IT WAS OBSERVED A MISMATCHED IN ONE SAMPLE BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE FIVE SAMPLES WERE EXAMINED AND NO INCORRECT COMPONENT WERE OBSERVED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8175521&pc=GAM | 2210968-2018-77869 |  | Malfunction | ETHICON INC. | 2018-12-18 05:00 | GAM | VICRYL CT UD 8X27IN 0 S/A CT-1 CR | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE OR STAFF OPENED THE STERILE SUTURE PACKET AND A DIFFERENT TYPE OF SUTURE AND NEEDLE WAS INSIDE. SUTURE FROM A DIFFERENT LOT WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 01/15/2019. INVESTIGATION SUMMARY: IT WAS REPORTED ASSEMBLY INCORRECT COMPONENT. AN EMPTY OPENED FOIL OF PRODUCT CODE JJ41, LOT # MGZ466 AND A LABELED WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8 OF PRODUCT CODE J774 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, A MISMATCHED BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF ASSEMBLY INCORRECT COMPONENT IS MIXED PRODUCT CODE. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND IT WAS OBSERVED A MISMATCHED IN ONE SAMPLE BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE FIVE SAMPLES WERE EXAMINED AND NO INCORRECT COMPONENT WERE OBSERVED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8175522&pc=GAM | 2210968-2018-77864 |  | Malfunction | ETHICON INC. | 2018-12-18 05:00 | GAM | VICRYL CT UD 8X27IN 0 S/A CT-1 CR | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE OR STAFF OPENED THE STERILE SUTURE PACKET AND A DIFFERENT TYPE OF SUTURE AND NEEDLE WAS INSIDE. SUTURE FROM A DIFFERENT LOT WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 01/15/2019. INVESTIGATION SUMMARY: IT WAS REPORTED ASSEMBLY INCORRECT COMPONENT. SIX UNOPENED SAMPLES OF PRODUCT CODE JJ41, LOT # MGZ466 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND IT WAS OBSERVED A MISMATCHED IN ONE SAMPLE BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE FIVE SAMPLES WERE EXAMINED AND NO INCORRECT COMPONENT WERE OBSERVED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF AN UNOPENED SAMPLE, THE ASSIGNABLE CAUSE OF ASSEMBLY INCORRECT COMPONENT IS MIXED PRODUCT CODE. PER THE CONDITION OF FIVE UNOPENED SAMPLES, NO ASSEMBLY INCORRECT COMPONENT WERE OBSERVED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8175526&pc=GAM | 2210968-2018-77865 |  | Malfunction | ETHICON INC. | 2018-12-18 05:00 | GAM | VICRYL CT UD 8X27IN 0 S/A CT-1 CR | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE OR STAFF OPENED THE STERILE SUTURE PACKET AND A DIFFERENT TYPE OF SUTURE AND NEEDLE WAS INSIDE. SUTURE FROM A DIFFERENT LOT WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 01/15/2019. INVESTIGATION SUMMARY: IT WAS REPORTED ASSEMBLY INCORRECT COMPONENT. AN EMPTY OPENED FOIL OF PRODUCT CODE JJ41, LOT # MGZ466 AND A LABELED WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8 OF PRODUCT CODE J774 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, A MISMATCHED BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF ASSEMBLY INCORRECT COMPONENT IS MIXED PRODUCT CODE. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND IT WAS OBSERVED A MISMATCHED IN ONE SAMPLE BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE FIVE SAMPLES WERE EXAMINED AND NO INCORRECT COMPONENT WERE OBSERVED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8175530&pc=GAM | 2210968-2018-77871 |  | Malfunction | ETHICON INC. | 2018-12-18 05:00 | GAM | VICRYL CT UD 8X27IN 0 S/A CT-1 CR | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE OR STAFF OPENED THE STERILE SUTURE PACKET AND A DIFFERENT TYPE OF SUTURE AND NEEDLE WAS INSIDE. SUTURE FROM A DIFFERENT LOT WERE USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 01/15/2019. INVESTIGATION SUMMARY: IT WAS REPORTED ASSEMBLY INCORRECT COMPONENT. AN EMPTY OPENED FOIL OF PRODUCT CODE JJ41, LOT # MGZ466 AND A LABELED WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8 OF PRODUCT CODE J774 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, A MISMATCHED BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF ASSEMBLY INCORRECT COMPONENT IS MIXED PRODUCT CODE. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND IT WAS OBSERVED A MISMATCHED IN ONE SAMPLE BETWEEN THE PRODUCT CODE JJ41 OF FOIL AND THE PRODUCT CODE J774 OF PAPER LID OF WINDING FORMER WERE OBSERVED. ALSO, THE SAMPLES INTO OF WINDING FORMER WERE EXAMINED AND BELONG TO PRODUCT CODE J774. THE FIVE SAMPLES WERE EXAMINED AND NO INCORRECT COMPONENT WERE OBSERVED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8168937&pc=GAM | 2210968-2018-77804 |  | Malfunction | ETHICON INC. | 2018-12-17 | GAM | VICRYL UND 3-0 8X45 CM CTX | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED PRIOR TO USE ON A PATIENT, IT WAS DISCOVERED THAT THERE HAD BEEN PDS IN THE PACKAGE OF VICRYL. THE LOT NUMBER IS UNKNOWN. FURTHER DETAILS ARE NOT PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8157209&pc=GAM | 2210968-2018-77729 | 2018-11-19 05:00 | Malfunction | ETHICON INC. | 2018-12-12 05:00 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A BOX OF SUTURE WAS MIXED FROM ONE PACKAGE INSIDE OF ANOTHER BOX. THIS OCCURRED PRE-OPERATIVELY AND WAS THEREFORE NOT USED ON A PATIENT. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 01/10/2019. INVESTIGATION SUMMARY: IT WAS REPORTED ASSEMBLY INCORRECT IMPLANTABLE. AN UNOPENED SAMPLE OF PRODUCT CODE VCP311, LOT MHM195 WERE RETURNED FOR ANALYSIS. THE BOX WAS NOT RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLE, IDENTIFIED WITH PRODUCT CODE VCP311, LOT MHM195 NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLE WAS OPENED, AND THE PRODUCT INTO THE PACKET BELONG TO CODE VCP311. NO FURTHER INVESTIGATION COULD BE PERFORMED DUE TO THE BOX WAS DISCARDED BY THE CUSTOMER AND IS NOT POSSIBLE IDENTIFIED THE BATCH NUMBER. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8124678&pc=GAM | 2210968-2018-77446 | 2018-10-15 | Malfunction | ETHICON INC. | 2018-12-03 | GAM | VICRYL CT UD 8X18IN 2-0 S/A CT-1 CR | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS OPENED, HOWEVER THE ITEM INSIDE WAS LABELED AS DIFFERENT PRODUCT CODE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS MEDWATCH REPORT IS IN RESPONSE TO RECEIPT OF MAUDE EVENT REPORT #MW5080600.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8124684&pc=GAM | 2210968-2018-77445 | 2018-10-15 | Malfunction | ETHICON INC. | 2018-12-03 | GAM | VICRYL CT UD 8X18IN 2-0 S/A CT-1 CR | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS OPENED, HOWEVER THE ITEM INSIDE WAS LABELED AS DIFFERENT PRODUCT CODE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS MEDWATCH REPORT IS IN RESPONSE TO RECEIPT OF MAUDE EVENT REPORT #MW5080600.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8111672&pc=GAM | 2210968-2018-77376 | 2018-10-24 04:00 | Malfunction | ETHICON INC. | 2018-11-28 05:00 | GAM | VICRYL CT UD 8X18IN 2-0 S/A CT-1 CR | Delivered as Unsterile Product | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CABG ON (B)(6) 2018, AND SUTURE WAS USED. WHEN THE BOX OF SUTURE WAS OPENED, IT APPEARED THAT ONE OF THE CONTROL RELEASE PACKAGES WAS SOILED INSIDE THE FOIL PACK AND THAT THE STERILITY OF THE PRODUCT WAS COMPROMISED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 12/28/2018. EVALUATION: INITIALLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, A STAIN IN A CONTROL RELEASE SAMPLE OF PRODUCT COULD BE OBSERVED. HOWEVER, NO CONCLUSION COULD BE REACHED AS TO THE ORIGIN OF THE FOREIGN MATTER. AN EMPTY OPENED BOX , AN EMPTY OPENED FOIL, A PAPER LID, A WINDING FORMER WITH EIGHT NEEDLE-SUTURE COMBINATIONS AND TWELVE UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE PAPER LID, A BROWN STAIN WAS NOTED. THE WINDING FORMER WAS EXAMINED AND A BROWN STAIN WERE OBSERVED ON THE BACK OF THE TRAY, THE SUTURE WERE DISPENSED WITHOUT PROBLEMS AND NO FOREIGN MATTER WERE NOTED ON THE STRANDS. IN THE INSPECTION OF THE FOIL THE INTEGRITY OF PACKAGE WAS NOT COMPROMISED. DURING THE VISUAL INSPECTION OF TWELVE UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND NO FOREIGN MATTER WAS OBSERVED; NOR THE INTEGRITY OF PACKAGES WERE COMPROMISED. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE OPENED SAMPLE, FOREIGN MATTER WAS FOUND, BUT THE REPORTED COMPLAINT PER PACKAGING INTEGRITY COULD NOT BE CONFIRMED DUE TO THE INTEGRITY OF PACKAGE OBSERVED NOT BEING COMPROMISED. PER THE CONDITION OF TWELVE UNOPENED SAMPLES, NO PACKAGING FOREIGN MATTER NON-BIOLOGICAL OR PACKAGING INTEGRITY WERE FOUND. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY, WHAT DO YOU MEAN BY " SOILED INSIDE THE FOIL PACK "? WAS THERE ANY FOREIGN MATTER FOUND ON THE INSIDE OF THE SUTURE PACK? OR WAS THERE ANY STAIN / INK SMUDGE ON THE INSIDE OF THE PACKAGE? THE SOIL WAS FOUND ON THE INSIDE OF THE FOIL PACK. WAS THE STERILITY OF SUTURE PRODUCT COMPROMISED IN ANY WAY? (PACKAGE NOT SEALED/ANY HOLE OR TEAR OR PUNCTURE IN PACKAGE) CAUSING THE ISSUE REPORTED IN EVENT? THE STERILITY OF THE PRODUCT WAS COMPROMISED. A PIC OF THE PRODUCT IS ATTACHED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066163&pc=GAM | 2210968-2018-77134 | 2018-10-11 04:00 | Malfunction | ETHICON INC. | 2018-11-13 05:00 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED AN ANIMAL UNDERWENT AN OVARIECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THREAD BROKE AT DIFFERENT PLACES POST OPERATIVELY. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 03/27/2019. ADDITIONAL: EIGHT UNOPENED SAMPLES OF PRODUCT CODE VCP473, LOT # MA8DXMR0 WERE RETURNED FOR ANALYSIS. THE ISSUE SAMPLE WAS NOT RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF EIGHT SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE REPRESENTATIVE SAMPLES, NO BREAKAGE SUTURE WAS FOUND AND THE TESTED SAMPLES MET THE REQUIREMENTS. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8023987&pc=GAM | 2210968-2018-76847 |  | Malfunction | ETHICON, INC. | 2018-10-31 | GAM | VICRYL RAPIDE | Break | No Consequences Or Impact To Patient | K033746 |  | 517 | Event Description: THIS REPORT SUMMARIZES <NOE> 517 </NOE> MALFUNCTION EVENTS, INCLUDING 242 SUTURE BREAKAGE, 111 NEEDLE BREAKAGE AND 164 DETACHMENT OF DEVICE COMPONENT THAT OCCURRED INTRAOPERATIVELY. Manufacturer Narrative: INITIAL SUTURE BREAKAGE MALFUNCTION EVENTS: 226. INITIAL SUTURE BREAKAGE DEVICES RETURNED: 246. 17 USE ERROR CAUSED OR CONTRIBUTED TO EVENT 192 NO DEVICE FAILURE. 21 MANUFACTURING DEFICIENCY. 16 UNABLE TO CONFIRM COMPLAINT. INITIAL NEEDLE BREAKAGE MALFUNCTION EVENTS: 102. INITIAL NEEDLE BREAKAGE DEVICES RETURNED: 161. 20 USE ERROR CAUSED OR CONTRIBUTED TO EVENT. 129 NO DEVICE FAILURE. 3 MANUFACTURING DEFICIENCY. 9 UNABLE TO CONFIRM COMPLAINT. INITIAL NEEDLE PULL OFF MALFUNCTION EVENTS: 147. INITIAL NEEDLE PULL OFF DEVICES RETURNED: 254. 15 USE ERROR CAUSED OR CONTRIBUTED TO EVENT. 159 NO DEVICE FAILURE. 48 MANUFACTURING DEFICIENCY. 32 UNABLE TO CONFIRM COMPLAINT. FOLLOW-UP SUTURE BREAKAGE MALFUNCTION EVENTS: 16. FOLLOW-UP SUTURE BREAKAGE DEVICES RETURNED: 41. 1 USE ERROR CAUSED OR CONTRIBUTED TO EVENT. 39 NO DEVICE FAILURE. 0 MANUFACTURING DEFICIENCY. 1 UNABLE TO CONFIRM COMPLAINT. FOLLOW-UP NEEDLE BREAKAGE MALFUNCTION EVENTS: 9. FOLLOW-UP NEEDLE BREAKAGE DEVICES RETURNED: 7. 1 USE ERROR CAUSED OR CONTRIBUTED TO EVENT 1 NO DEVICE FAILURE. 3 MANUFACTURING DEFICIENCY. 2 UNABLE TO CONFIRM COMPLAINT. FOLLOW-UP NEEDLE PULL OFF MALFUNCTION EVENTS: 17. FOLLOW-UP NEEDLE PULL OFF DEVICES RETURNED: 43. 5 USE ERROR CAUSED OR CONTRIBUTED TO EVENT. 32 NO DEVICE FAILURE. 4 MANUFACTURING DEFICIENCY. 2 UNABLE TO CONFIRM COMPLAINT. ALL DEVICES LABELED FOR SINGLE USE; 0 DEVICES REPROCESSED AND RE-USED 0 REMEDIAL ACTION. COMPLAINT INFORMATION WILL BE INCLUDED IN COMPLAINT TRENDING THAT IS REVIEWED ON A REGULAR BASIS TO DETERMINE IF FURTHER ACTION IS NECESSARY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8010563&pc=GAM | 2210968-2018-76788 |  | Malfunction | ETHICON INC. | 2018-10-26 04:00 | GAM | VICRYL RAP UND 70CM M1.5 | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TESTICULAR BIOPSY PROCEDURE IN 2018 AND A SUTURE WAS USED. AFTER THE TESTICULAR BIOPSY, THE SUTURE IS REABSORBING TOO FAST (AFTER 2-3 DAYS). THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 11/28/2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. EVALUATION: REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF MULTIPLE DEVICE ISSUES CAPTURED IN REPORTS 2210968-2018-76745, 2210968-2018-76785, 2210968-2018-76786, 2210968-2018-76787, AND 2210968-2018-76788. ANALYSIS RESULTS HAVE BEEN INCLUDED IN FOLLOW-UP REPORTS FOR EACH. AN OPENED BOX WITH TWENTY-NINE UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE ISSUE SAMPLE WAS NOT RECEIVED FOR EVALUATION. THIRTEEN REPRESENTATIVE SAMPLES WERE EXAMINED FOR VISUAL DEFECTS. THE PACKETS WERE OPENED AND DURING THE VISUAL INSPECTION, THE SUTURES WERE CORRECTLY PLACED ON THE WINDING FORMERS, THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG THE STRANDS WITH NO DEFECTS OR DAMAGES OBSERVED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES RECEIVED, THE ASSIGNABLE CAUSE OF THE RAPID ABSORPTION OF SUTURE CANNOT BE DETERMINED SINCE NO DEFECTS WERE OBSERVED ON THE PACKETS OR THE SUTURES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT DO YOU MEAN BY "SUTURE IS RESORBING TOO FAST"? SUTURE ABSOPTED WITHIN 2 TO 3 DAYS CONFIRM THE TOTAL QTY OF SUTURES INVOLVED WITH THE REPORTED ISSUE? WERE ALL 5 SUTURES REPORTED TO HAVE ISSUE USED ON SAME PATIENT/SINGLE PROCEDURE? YES
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8010567&pc=GAM | 2210968-2018-76786 |  | Malfunction | ETHICON INC. | 2018-10-26 04:00 | GAM | VICRYL RAP UND 70CM M1.5 | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TESTICULAR BIOPSY PROCEDURE IN 2018 AND A SUTURE WAS USED. AFTER THE TESTICULAR BIOPSY, THE SUTURE IS REABSORBING TOO FAST (AFTER 2-3 DAYS). THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF MULTIPLE DEVICE ISSUES CAPTURED IN REPORTS 2210968-2018-76745, 2210968-2018-76785, 2210968-2018-76786, 2210968-2018-76787, AND 2210968-2018-76788. ANALYSIS RESULTS HAVE BEEN INCLUDED IN FOLLOW-UP REPORTS FOR EACH. AN OPENED BOX WITH TWENTY-NINE UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE ISSUE SAMPLE WAS NOT RECEIVED FOR EVALUATION THIRTEEN REPRESENTATIVE SAMPLES WERE EXAMINED FOR VISUAL DEFECTS. THE PACKETS WERE OPENED AND DURING THE VISUAL INSPECTION, THE SUTURES WERE CORRECTLY PLACED ON THE WINDING FORMERS, THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG THE STRANDS WITH NO DEFECTS OR DAMAGES OBSERVED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES RECEIVED, THE ASSIGNABLE CAUSE OF THE RAPID ABSORPTION OF SUTURE CANNOT BE DETERMINED SINCE NO DEFECTS WERE OBSERVED ON THE PACKETS OR THE SUTURES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT DO YOU MEAN BY "SUTURE IS RESORBING TOO FAST"? SUTURE ABSORBED WITHIN 2 TO 3 DAYS. CONFIRM THE TOTAL QTY OF SUTURES INVOLVED WITH THE REPORTED ISSUE? WERE ALL 5 SUTURES REPORTED TO HAVE ISSUE USED ON SAME PATIENT/SINGLE PROCEDURE? YES
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8010568&pc=GAM | 2210968-2018-76785 |  | Malfunction | ETHICON INC. | 2018-10-26 04:00 | GAM | VICRYL RAP UND 70CM M1.5 | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TESTICULAR BIOPSY PROCEDURE IN 2018 AND A SUTURE WAS USED. AFTER THE TESTICULAR BIOPSY, THE SUTURE IS REABSORBING TOO FAST (AFTER 2-3 DAYS). THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 11/28/2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. EVALUATION: REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF MULTIPLE DEVICE ISSUES CAPTURED IN REPORTS 2210968-2018-76745, 2210968-2018-76785, 2210968-2018-76786, 2210968-2018-76787, AND 2210968-2018-76788. ANALYSIS RESULTS HAVE BEEN INCLUDED IN FOLLOW-UP REPORTS FOR EACH. AN OPENED BOX WITH TWENTY-NINE UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE ISSUE SAMPLE WAS NOT RECEIVED FOR EVALUATION. THIRTEEN REPRESENTATIVE SAMPLES WERE EXAMINED FOR VISUAL DEFECTS. THE PACKETS WERE OPENED AND DURING THE VISUAL INSPECTION, THE SUTURES WERE CORRECTLY PLACED ON THE WINDING FORMERS, THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG THE STRANDS WITH NO DEFECTS OR DAMAGES OBSERVED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES RECEIVED, THE ASSIGNABLE CAUSE OF THE RAPID ABSORPTION OF SUTURE CANNOT BE DETERMINED SINCE NO DEFECTS WERE OBSERVED ON THE PACKETS OR THE SUTURES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT DO YOU MEAN BY "SUTURE IS REABSORBING TOO FAST"? SUTURE ABSORBED WITHIN 2 TO 3 DAYS. CONFIRM THE TOTAL QTY OF SUTURES INVOLVED WITH THE REPORTED ISSUE? WERE ALL 5 SUTURES REPORTED TO HAVE ISSUE USED ON SAME PATIENT/SINGLE PROCEDURE? YES
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8010572&pc=GAM | 2210968-2018-76787 |  | Malfunction | ETHICON INC. | 2018-10-26 04:00 | GAM | VICRYL RAP UND 70CM M1.5 | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TESTICULAR BIOPSY PROCEDURE IN 2018 AND A SUTURE WAS USED. AFTER THE TESTICULAR BIOPSY, THE SUTURE IS REABSORBING TOO FAST (AFTER 2-3 DAYS). THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 11/28/2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. EVALUATION: REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF MULTIPLE DEVICE ISSUES CAPTURED IN REPORTS 2210968-2018-76745, 2210968-2018-76785, 2210968-2018-76786, 2210968-2018-76787, AND 2210968-2018-76788. ANALYSIS RESULTS HAVE BEEN INCLUDED IN FOLLOW-UP REPORTS FOR EACH. AN OPENED BOX WITH TWENTY-NINE UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE ISSUE SAMPLE WAS NOT RECEIVED FOR EVALUATION. THIRTEEN REPRESENTATIVE SAMPLES WERE EXAMINED FOR VISUAL DEFECTS. THE PACKETS WERE OPENED AND DURING THE VISUAL INSPECTION, THE SUTURES WERE CORRECTLY PLACED ON THE WINDING FORMERS, THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG THE STRANDS WITH NO DEFECTS OR DAMAGES OBSERVED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES RECEIVED, THE ASSIGNABLE CAUSE OF THE RAPID ABSORPTION OF SUTURE CANNOT BE DETERMINED SINCE NO DEFECTS WERE OBSERVED ON THE PACKETS OR THE SUTURES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT DO YOU MEAN BY "SUTURE IS REABSORBING TOO FAST"? SUTURE ABSORBED WITHIN 2 TO 3 DAYS. CONFIRM THE TOTAL QTY OF SUTURES INVOLVED WITH THE REPORTED ISSUE? WERE ALL 5 SUTURES REPORTED TO HAVE ISSUE USED ON SAME PATIENT/SINGLE PROCEDURE? YES
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8000587&pc=GAM | 2210968-2018-76729 | 2018-09-12 04:00 | Malfunction | ETHICON INC. | 2018-10-24 04:00 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A HOSPITAL MADE A ROUTINE IN VITRO CULTURE FOR THE STERILE MEDICAL DEVICES. IT WAS FOUND THAT SUTURE CULTURED ON (B)(6) 2018 AND (B)(6) 2018 HAD BACTERIAL GROWTH. THE CULTURING METHOD IS BROTH ENRICHMENT CULTURE, AND CULTURE TIME IS 72 HOURS. AND THEN THE HOSPITAL REQUIRED A DIFFERENT HOSPITAL TO RE-CULTURE FOR THIS LOT SUTURE, AND THAT OTHER HOSPITAL CULTURE RESULT IS NO BACTERIAL GROWTH. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION AS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 12/18/2018. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. Manufacturer Narrative: DATE SENT TO THE FDA: 02/21/2019. INVESTIGATION SUMMARY: IT WAS REPORTED INTEGRITY ISSUES. TWO UNOPENED SAMPLES OF PRODUCT CODE VCP358H WITH LOT# MC7946 WAS RECEIVED FOR EVALUATION. THE SAMPLES WERE VISUALLY INSPECTED, AND NO ANOMALIES WERE FOUND. THE CERTIFICATE OF ANALYSIS OF THE BATCH WAS REVIEW AND RESULTS ARE ABOVE THE IN-VITRO REFERENCE RESULTS. ALONG WITH THE VALIDATION OF THE STERILIZATION PROCESS, CONTROLS TO MONITOR EACH STERILIZATION CYCLE ARE IN PLACE TO ENSURE EVERY STERILIZATION RUN IS EXECUTED WITHIN THE CRITICAL PROCESS PARAMETERS REQUIRED TO DELIVER A STERILITY ASSURANCE LEVEL (SAL) OF 10-6. THE ROUTINE CONTROL OF THE MANUFACTURING PROCESS INCLUDES ENVIRONMENTAL MONITORING OF THE MANUFACTURING AREAS, ALONG WITH THE MONITORING THE QUANTITY OF MICROORGANISM PRESENT ON NON-STERILE PRODUCTS (BIOBURDEN ANALYSIS). THE CONTROLS FOR ROUTINE PRODUCTION AND STERILIZATION OF THE FINISHED DEVICE PROVIDE A GREATER LEVEL OF STERILITY ASSURANCE THAN ACTUAL PRODUCT TESTING. THE METHOD OF STERILITY TESTING DEVICES REQUIRES HUMAN INTERVENTION TO TRANSFER THE PRODUCT FROM THE STERILE BARRIER PACKAGING INTO A TEST MEDIA. THE PROCESS OF TRANSFER INTRODUCES OPPORTUNITIES FOR EXTRINSIC CONTAMINATION OF THE SAMPLE, ALLOWING NO CONCLUSIONS TO BE DRAWN WITH REGARDS TO THE PRODUCT OVERALL STERILITY ASSURANCE. BASED ON THE ROBUSTNESS OF THE STERILIZATION VALIDATION METHODOLOGY, THE ROUTINE MONITORING AND CONTROLS OF THE PRODUCT AND PROCESSES AND THE INHERENT CONTAMINATION RISKS ASSOCIATED WITH DEVICE STERILITY TESTING, NO STERILITY TESTING ON THE DEVICES IS RECOMMENDED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8000598&pc=GAM | 2210968-2018-76732 | 2018-09-12 04:00 | Malfunction | ETHICON INC. | 2018-10-24 04:00 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A HOSPITAL MADE A ROUTINE IN VITRO CULTURE FOR THE STERILE MEDICAL DEVICES. IT WAS FOUND THAT SUTURE CULTURED ON (B)(6) 2018 AND (B)(6) 2018, HAD BACTERIAL GROWTH. THE CULTURING METHOD IS BROTH ENRICHMENT CULTURE, AND CULTURE TIME IS 72 HOURS. AND THEN THE HOSPITAL REQUIRED A DIFFERENT HOSPITAL TO RE-CULTURE FOR THIS LOT SUTURE, AND THAT OTHER HOSPITAL CULTURE RESULT IS NO BACTERIAL GROWTH. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION AS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 12/18/2018. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 02/21/2019. INVESTIGATION SUMMARY: IT WAS REPORTED INTEGRITY ISSUES. TWO UNOPENED SAMPLES OF PRODUCT CODE VCP358H WITH LOT# MC7946 WAS RECEIVED FOR EVALUATION. THE SAMPLES WERE VISUALLY INSPECTED, AND NO ANOMALIES WERE FOUND. THE CERTIFICATE OF ANALYSIS OF THE BATCH WAS REVIEW AND RESULTS ARE ABOVE THE IN-VITRO REFERENCE RESULTS. ALONG WITH THE VALIDATION OF THE STERILIZATION PROCESS, CONTROLS TO MONITOR EACH STERILIZATION CYCLE ARE IN PLACE TO ENSURE EVERY STERILIZATION RUN IS EXECUTED WITHIN THE CRITICAL PROCESS PARAMETERS REQUIRED TO DELIVER A STERILITY ASSURANCE LEVEL (SAL) OF 10-6. THE ROUTINE CONTROL OF THE MANUFACTURING PROCESS INCLUDES ENVIRONMENTAL MONITORING OF THE MANUFACTURING AREAS, ALONG WITH THE MONITORING THE QUANTITY OF MICROORGANISM PRESENT ON NON-STERILE PRODUCTS (BIOBURDEN ANALYSIS). THE CONTROLS FOR ROUTINE PRODUCTION AND STERILIZATION OF THE FINISHED DEVICE PROVIDE A GREATER LEVEL OF STERILITY ASSURANCE THAN ACTUAL PRODUCT TESTING. THE METHOD OF STERILITY TESTING DEVICES REQUIRES HUMAN INTERVENTION TO TRANSFER THE PRODUCT FROM THE STERILE BARRIER PACKAGING INTO A TEST MEDIA. THE PROCESS OF TRANSFER INTRODUCES OPPORTUNITIES FOR EXTRINSIC CONTAMINATION OF THE SAMPLE, ALLOWING NO CONCLUSIONS TO BE DRAWN WITH REGARDS TO THE PRODUCT OVERALL STERILITY ASSURANCE. BASED ON THE ROBUSTNESS OF THE STERILIZATION VALIDATION METHODOLOGY, THE ROUTINE MONITORING AND CONTROLS OF THE PRODUCT AND PROCESSES AND THE INHERENT CONTAMINATION RISKS ASSOCIATED WITH DEVICE STERILITY TESTING, NO STERILITY TESTING ON THE DEVICES IS RECOMMENDED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8000695&pc=GAM | 2210968-2018-76736 | 2018-09-12 04:00 | Malfunction | ETHICON INC. | 2018-10-24 04:00 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A HOSPITAL MADE A ROUTINE IN VITRO CULTURE FOR THE STERILE MEDICAL DEVICES. IT WAS FOUND THAT SUTURE CULTURED ON (B)(6) 2018 AND (B)(6) 2018 HAD BACTERIAL GROWTH. THE CULTURING METHOD IS BROTH ENRICHMENT CULTURE, AND CULTURE TIME IS 72 HOURS. AND THEN THE HOSPITAL REQUIRED A DIFFERENT HOSPITAL TO RE-CULTURE FOR THIS LOT SUTURE, AND THAT OTHER HOSPITAL CULTURE RESULT IS NO BACTERIAL GROWTH. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION AS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 12/18/2018. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. Manufacturer Narrative: DATE SENT TO THE FDA: 02/21/2019. INVESTIGATION SUMMARY: IT WAS REPORTED INTEGRITY ISSUES. TWO UNOPENED SAMPLES OF PRODUCT CODE VCP358H WITH LOT# MC7946 WAS RECEIVED FOR EVALUATION. THE SAMPLES WERE VISUALLY INSPECTED, AND NO ANOMALIES WERE FOUND. THE CERTIFICATE OF ANALYSIS OF THE BATCH WAS REVIEW AND RESULTS ARE ABOVE THE IN-VITRO REFERENCE RESULTS. ALONG WITH THE VALIDATION OF THE STERILIZATION PROCESS, CONTROLS TO MONITOR EACH STERILIZATION CYCLE ARE IN PLACE TO ENSURE EVERY STERILIZATION RUN IS EXECUTED WITHIN THE CRITICAL PROCESS PARAMETERS REQUIRED TO DELIVER A STERILITY ASSURANCE LEVEL (SAL) OF 10-6. THE ROUTINE CONTROL OF THE MANUFACTURING PROCESS INCLUDES ENVIRONMENTAL MONITORING OF THE MANUFACTURING AREAS, ALONG WITH THE MONITORING THE QUANTITY OF MICROORGANISM PRESENT ON NON-STERILE PRODUCTS (BIOBURDEN ANALYSIS). THE CONTROLS FOR ROUTINE PRODUCTION AND STERILIZATION OF THE FINISHED DEVICE PROVIDE A GREATER LEVEL OF STERILITY ASSURANCE THAN ACTUAL PRODUCT TESTING. THE METHOD OF STERILITY TESTING DEVICES REQUIRES HUMAN INTERVENTION TO TRANSFER THE PRODUCT FROM THE STERILE BARRIER PACKAGING INTO A TEST MEDIA. THE PROCESS OF TRANSFER INTRODUCES OPPORTUNITIES FOR EXTRINSIC CONTAMINATION OF THE SAMPLE, ALLOWING NO CONCLUSIONS TO BE DRAWN WITH REGARDS TO THE PRODUCT OVERALL STERILITY ASSURANCE. BASED ON THE ROBUSTNESS OF THE STERILIZATION VALIDATION METHODOLOGY, THE ROUTINE MONITORING AND CONTROLS OF THE PRODUCT AND PROCESSES AND THE INHERENT CONTAMINATION RISKS ASSOCIATED WITH DEVICE STERILITY TESTING, NO STERILITY TESTING ON THE DEVICES IS RECOMMENDED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002316&pc=GAM | 2210968-2018-76745 |  | Malfunction | ETHICON INC. | 2018-10-24 04:00 | GAM | VICRYL RAP UND 70CM M1.5 | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TESTICULAR BIOPSY PROCEDURE IN 2018, AND A SUTURE WAS USED. AFTER THE TESTICULAR BIOPSY, THE SUTURE IS REABSORBING TOO FAST (AFTER 2-3 DAYS). THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 11/28/2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. EVALUATION: REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF MULTIPLE DEVICE ISSUES CAPTURED IN REPORTS 2210968-2018-76745, 2210968-2018-76785, 2210968-2018-76786, 2210968-2018-76787, AND 2210968-2018-76788. ANALYSIS RESULTS HAVE BEEN INCLUDED IN FOLLOW-UP REPORTS FOR EACH. AN OPENED BOX WITH TWENTY-NINE UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE ISSUE SAMPLE WAS NOT RECEIVED FOR EVALUATION. THIRTEEN REPRESENTATIVE SAMPLES WERE EXAMINED FOR VISUAL DEFECTS. THE PACKETS WERE OPENED AND DURING THE VISUAL INSPECTION, THE SUTURES WERE CORRECTLY PLACED ON THE WINDING FORMERS, THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG THE STRANDS WITH NO DEFECTS OR DAMAGES OBSERVED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES RECEIVED, THE ASSIGNABLE CAUSE OF THE RAPID ABSORPTION OF SUTURE CANNOT BE DETERMINED SINCE NO DEFECTS WERE OBSERVED ON THE PACKETS OR THE SUTURES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT DO YOU MEAN BY "SUTURE IS RESORBING TOO FAST"? SUTURE ABSORBED WITHIN 2 TO 3 DAYS CONFIRM THE TOTAL QTY OF SUTURES INVOLVED WITH THE REPORTED ISSUE? WERE ALL 5 SUTURES REPORTED TO HAVE ISSUE USED ON SAME PATIENT/SINGLE PROCEDURE? YES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002491&pc=GAM | 2210968-2018-76713 | 2018-09-20 04:00 | Malfunction | ETHICON INC. | 2018-10-24 04:00 | GAM | VICRYL SUTURE 54"(135CM) 3-0 VIO | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE (B)(6) 2018, AND A SUTURE WAS USED. THE VICRYL SUTURE DIDN'T ABSORB IN AN APPROPRIATE TIME FRAME. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 11/28/2018. INVESTIGATION SUMMARY: IT WAS REPORTED ABSORPTION OF SUTURE. AN UNOPENED SAMPLE OF PRODUCT CODE J531 LOT MBZ730 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE REPRESENTATIVE SAMPLE, WAS EXAMINED FOR VISUAL DEFECTS: APPEARANCE, COLOR PACKAGES, WRINKLES IN THE SEAL AREA, CONTINUOUS SEALS, SEAL MARGINS, OVER-SEALING, DAMAGE (TORN, PUNCTURED). THE PACKET WAS OPENED AND DURING THE VISUAL INSPECTION, THE SWAGE AND ATTACHMENT AREA WERE AS EXPECTED, THE SUTURE STRAND WAS CORRECTLY PLACED ON THE TRAY AND DISPENSED WITHOUT PROBLEMS. NO DEFECTS OR DAMAGED WERE OBSERVED ALONG OF SUTURE. IN ADDITION, THE VICRYL SUTURE RETAINS APPROXIMATELY 75 % OF THE ORIGINAL TENSILE STRENGTH TWO WEEKS POST IMPLANTATION. AT THREE WEEKS, APPROXIMATELY 50% AND THE ABSORPTION TOTAL OF VICRYL SUTURE IS ESSENTIALLY COMPLETE BETWEEN 56 AND 70 DAYS. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE REPRESENTATIVE SAMPLE, NO DEFECTS WERE FOUND. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? UNSURE, AS I DIDN¿T DO THE REPAIRS. I DID REMOVE KNOTS FROM EACH OF THE PATIENTS AND THEN LONG STRANDS OF SUTURE MATERIAL. THEREFORE I¿D ASSUME IT WAS A CONTINUOUS STITCH. WHICH TISSUE LAYER, AT WHICH LOCATION WAS THE SUTURE, WAS USED TO CLOSE? ALL OF THESE PATIENTS HAD 2ND DEGREE TEARS (MUSCLE LAYERS) AND ONE HAD AN EPISIOTOMY (ALSO A MUSCLE LAYER). DID THE SUTURE ABSORB PRIOR TO OR BEYOND THE APPROPRIATE TIME FRAME? THE SUTURE MATERIAL THAT COULD NOT BE REMOVED, AS FAR AS I CAN TELL, WAS ABSORBED IN THE APPROPRIATE TIME FRAME. HOWEVER, IN ONE CASE WE COULD NOT GET ACCESS TO THE KNOT, WE COULD SEE IT BELOW THE SKIN SURFACE, BUT THE SKIN HAD HEALED OVER IT. I¿M NOT SURE IF THAT KNOT EVER DISSOLVED. THE PROBLEM WASN¿T THAT IT WAS ¿NOTICEABLE¿, IT WAS THAT THESE WOMEN WERE IN SO MUCH PAIN AT 2+ WEEKS POSTPARTUM (AFTER THEY HAD STATED THEIR PERINEUM WAS FEELING BETTER IN THE FIRST COUPLE WEEKS), THAT I HAD TO INSPECT THE REPAIR AND REMOVE THE MATERIAL. I¿M UNSURE IF THE MATERIAL WOULD HAVE ABSORBED HAD IT BEEN LEFT IN. I ALSO WOULD NOT HAVE KEPT THE MATERIAL IN TO FIND OUT, AS THESE WOMEN WERE IN SO MUCH PAIN AND HAD DISCOMFORT. WHAT WAS THE SURGEONS EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION¿? I TELL ALL MY PATIENTS IT TAKES 55 - 70 DAYS TO DISSOLVE AND NOT BE SEEN, HOWEVER THE REPAIR SHOULD FEEL BETTER AND BETTER EACH DAY AND BY 2 WEEKS SHE SHOULD HAVE LITTLE TO NO DISCOMFORT. IN THESE CASES, BY 2 WEEKS THEIR PERINEUM AND VAGINA WERE EXTREMELY PAINFUL. DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? I DID NOT DO THE REPAIR, SO I¿M UNSURE OF ANY DEFICIENCY OR ANOMALY. WAS THE SUTURE PACKAGE REPROCESSED OR STERILIZED? ALL OUR SUTURES COME FROM THE MANUFACTURER, THE PACKAGING WAS NOT REPROCESSED WAS THE PACKAGE NOTED TO BE DAMAGED? I DID NOT DO THE REPAIR, BUT I¿D ASSUME IT WOULDN¿T BE USED IF THERE WAS PACKAGE DAMAGE NOTED. THE OBSTETRICIAN STATED HE WOULD NOT USE A PACKAGE THAT WAS DAMAGED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEVICE CODES: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE LOT NUMBER? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? WHICH TISSUE LAYER, AT WHICH LOCATION WAS THE SUTURE, WAS USED TO CLOSE? DID THE SUTURE ABSORB PRIOR TO OR BEYOND THE APPROPRIATE TIME FRAME? WHAT WAS THE SURGEONS EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION"? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? WAS THE SUTURE PACKAGE REPROCESSED OR STERILIZED? WAS THE PACKAGE NOTED TO BE DAMAGED? THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7952397&pc=GAM | 2210968-2018-76447 | 2018-09-12 | Malfunction | ETHICON INC. | 2018-10-10 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 VIO | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2018 AND SUTURE WAS USED. IT WAS FOUND THAT THERE HAD BEEN NO PAPER LABEL IN THE PACKAGE WHEN OPENING THE PACKAGE. FURTHER DETAILS ARE NOT PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: PRODUCT COMPLAINT # PC-(B)(4). AN EMPTY OPENED FOIL AND A WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8, WERE RETURNED FOR ANALYSIS. THIS PRODUCT REQUIRES 8 STRANDS PER PACKET. DURING THE VISUAL INSPECTION OF WINDING FORMER, NO PAPER LID WAS PLACED ON THE WINDING FORMER, SINCE THE INSERTION TABS WERE NOT OBSERVED; THIS INDICATES THAT THE PAPER LID WAS NOT SECURED CORRECTLY. ALSO, THE NEEDLES AND SUTURES WERE EXAMINED AND WERE CORRECT BY THE PRODUCT. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE, THE ASSIGNABLE CAUSE OF THE PACKAGING LABELING IDENTIFICATION IS A WRONG NUMBER OF COMPONENTS. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE MISSING PAPER LABEL PREVENT PROPER IDENTIFICATION OF THE PRODUCT? IT IS SUPPOSED THAT THE SIZE WERE IDENTIFIED WITH ALUMINUM PACKAGE. NO FURTHER INFORMATION WILL BE PROVIDED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7946930&pc=GAM | 2210968-2018-76407 | 2018-09-11 04:00 | Malfunction | ETHICON INC. | 2018-10-09 04:00 | GAM | VICRYL SUTURE 18"(45CM) 5-0 VIO | Component Misassembled | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED WHILE STOCKING THE DEVICE AT FACILITY ON AN UNKNOWN DATE, IT WAS NOTICED THAT THE PRODUCT LISTED ONLINE AND PRODUCT MASTER THE NEEDLE LISTED IS A TF-4 BUT THE PRODUCT CAME WITH A DIFFERENT CONFIGURATION: TF-5 NEEDLE. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 12/14/2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. EVALUATION: THERE WAS NO SAMPLE RECEIVED FOR ANALYSIS. ONLY PICTURES OF THE SAMPLE WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, A BOX OF PRODUCT WITH A NEEDLE DESCRIPTION COULD BE OBSERVED. ACCORDING TO THE MANUFACTURING RECORD, THE COMPONENTS OF THE PRODUCT ARE CORRECT. HOWEVER, THE INFORMATION OF THE PRODUCT AT THE ETHICON WEB PAGE ISN'T ACTUALIZED AS THE DESCRIPTION IS FOR ANOTHER NEEDLE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7929408&pc=GAM | 7929408 | 2018-09-20 | Malfunction | ETHICON, INC. | 2018-10-03 | GAM | VICRYL PLUS UR-6 | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: PATIENT WAS UNDERGOING A LAPAROSCOPIC CHOLECYSTECTOMY. UPON USING A SUTURE 0 VICRYL UR-6 BY THE SURGEON, THE NEEDLE BROKE. XRAY WAS UTILIZED TO FIND THE MISSING NEEDLE PIECE. IT WAS RETRIEVED FROM THE ABDOMINAL TISSUE. THE DETACHED PIECE WAS CONFIRMED TO BE OUTSIDE OF THE PATIENT'S BODY (ABDOMEN) THROUGH XRAY AND VISUAL VERIFICATION BY THE SURGEON, SCRUB TECHNICIAN AND THE CIRCULATING NURSE. BROKEN PIECE DISCARDED. NO PT HARM. THE ONE USED BY THE SURGEON THAT BROKE: ETHICON: 0 VICRYL PLUS UR-6 (REF VCP603); LOT # MDM357 OR # MBM414 (UNSURE WHICH OF THESE 2 LOT#'S IS THE ONE THAT BROKE); EXP 2023.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7931481&pc=GAM | 2210968-2018-76284 |  | Malfunction | ETHICON INC. | 2018-10-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Component Misassembled | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE, THE CUSTOMER INQUIRED ABOUT THE DESCRIPTION IN OUR CATALOG VS. THE SUTURE BOXES. THE DESCRIPTION ON THE WEBSITE AND IN THE CATALOG SAYS IT IS ON AN RB-2 NEEDLE, BUT THE PACKAGE AND THE NEEDLE INSIDE THE PACKAGE IS A TF NEEDLE. NO PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: LOT #? ANY SAMPLE AVAILABLE TO BE RETURNED FOR EVALUATION? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7917279&pc=GAM | 2210968-2018-76127 | 2018-01-19 | Malfunction | ETHICON INC. | 2018-09-28 | GAM | VICRYL SUTURE 27"(70CM) 5-0 UND | Break | Blood Loss | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SURGICAL PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. THE SUTURE IS EASY TO BREAK THAN PREVIOUS ONE DURING SEWING IN ADIPOSE TISSUE. CHANGED NEW SUTURE TO COMPLETE THE SURGERY. THE HOSPITAL CLAIMED THAT IT WILL CAUSE INJURY ON TISSUE DUE TO REPEAT SEWING AND MUCH BLEEDING. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE PROVIDE THE PRODUCT CODE AND LOT NUMBER OF THE PRODUCT INVOLVED. AFTER PHONE CALL WITH CUSTOMER, THE CODE IS (B)(4) AND LOT IS LMM236. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHEN DID THE BLEEDING OCCUR? THE SOURCE AND TRIGGERING EVENT OF BLEEDING AND THE VOLUME OF BLOOD LOSS? HOW THE BLEEDING WAS TREATED? IF A SURGICAL INTERVENTION WAS NEEDED, PLEASE PROVIDE DETAILS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7916238&pc=GAM | 2210968-2018-76121 | 2018-08-13 | Malfunction | ETHICON INC. | 2018-09-27 | GAM | VICRYL RAPIDE SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: A DEVICE WAS RETURNED FOR EVALUATION. THE EVALUATION SHOWED A MULTITUDE OF WRINKLES OVER THE WHOLE FOIL. SEVERAL PINHOLES AT TOP LABEL SIDE AS WELL AS BOTTOM CAVITY SIDE OF THE PACKS WERE OBSERVED. THIS DEVICE WAS NOT USED ON A PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). INVESTIGATION SUMMARY: WE RECEIVED 1 EA OF COMPLAINT SAMPLE FOIL, SBF FOLDER SUTURE PIECES CODE (B)(4). SUTURE RECEIVED IN DISINTEGRATED CONDITION. AS THE COMPLAINT SAMPLE RECEIVED IN OPEN CONDITION FURTHER INVESTIGATION ON COMPLAINT SAMPLE CANNOT BE PERFORMED EXCEPT VISUAL INSPECTION. THE FOIL PACK WAS INSPECTED UNDER A 10 X MAGNIFICATION FOR ANY DAMAGE AND SHOWED A MULTITUDE OF WRINKLES OVER THE WHOLE FOIL ALONG WITH SEVERAL PINHOLES AT TOP LABEL SIDE AS WELL AS BOTTOM CAVITY SIDE OF THE PACKS WERE OBSERVED. FIVE RETAIN SAMPLES OF INCIDENT CODES AND LOT NUMBER WAS RETRIEVED FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS LIKE PSEUDO SEAL, PIN HOLE, PRODUCT IN SEAL, PRESS MARKS ON PACK BUT NO SUCH DEFECTS WERE OBSERVED. THE PRIMARY PACKS WERE OPENED, AND SUTURES AND NEEDLES WERE FOUND INTACT. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE KINKS, WEAK SPOTS, BROKEN PIECE, FRAY, BRITTLENESS, DETACHED NEEDLES BUT NO SUCH DEFECTS WERE OBSERVED. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE BEND, CRACKED BARREL, FINS BUT NO SUCH DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WERE TESTED FOR KNOT PULL TENSILE TEST AND FOUND TO MEET THE SPECIFICATION. AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED FOR ANY PROCESS DEVIATION, BUT NO DEVIATION WAS OBSERVED. FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION. FROM THE ABOVE ANALYSIS, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. AS THE COMPLAINT IS RELATED TO PERFORMANCE ¿ BREAKAGE SUTURE , KNOT PULL TENSILE STRENGTH TEST IS APPLICABLE & MEASURABLE PARAMETER. KNOT PULL TENSILE STRENGTH TEST WAS PERFORMED OVER FIVE RETAIN SAMPLES TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. THE AVERAGE KNOT PULL TENSILE STRENGTH VALUE OF THE RETAIN SAMPLE WAS FOUND TO MEET THE IN-HOUSE AVERAGE KNOT PULL TENSILE STRENGTH REQUIREMENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7894993&pc=GAM | 2210968-2018-75912 |  | Malfunction | ETHICON INC. | 2018-09-20 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OP THE PATIENT COMPLAINED OF SPITTING SUTURE PRIOR TO THE TWO WEEK POST-OP VISIT. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS MEDICAL OR SURGICAL INTERVENTION NEEDED TO ADDRESS THE ISSUE? WHAT IS THE PATIENT'S CURRENT CONDITION? WHAT IS THE LOT NUMBER OF THE PRODUCT INVOLVED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884812&pc=GAM | 2210968-2018-75867 |  | Malfunction | ETHICON INC. | 2018-09-18 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Appropriate Term/Code Not Available | No Code Available | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OPERATIVELY, THE ABSORBABLE SUTURES WERE NOT ABSORBING, AND THE SUTURES WERE "SPITTING THROUGH THE SKIN". IT IS UNKNOWN HOW OR IF THE PATIENT WAS TREATED FOR THIS CONDITION. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DATE OF INITIAL PROCEDURE; PROCEDURE NAME; WHICH TISSUE WAS THE SUTURE USED ON? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? WAS THE SUTURE PACKAGE REPROCESSED OR STERILIZED? HOW MANY DAYS POST -OP WAS THE SUTURE SPITTING/NON-ABSORPTION ISSUE NOTED? WAS WOUND DEHISCENCE NOTED? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED? PATIENT'S GENDER, AGE, WEIGHT, MEDICAL/ SURGICAL HISTORY & MEDICINES USED? PATIENT'S CURRENT CONDITION. LOT NUMBER OF THE VICRYL SUTURE. ANY REPRESENTATIVE SAMPLES OF THE SAME LOT AVAILABLE FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884819&pc=GAM | 2210968-2018-75863 |  | Malfunction | ETHICON INC. | 2018-09-18 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Appropriate Term/Code Not Available | No Code Available | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OPERATIVELY, THE ABSORBABLE SUTURES WERE NOT ABSORBING, AND THE SUTURES WERE "SPITTING THROUGH THE SKIN". IT IS UNKNOWN HOW OR IF THE PATIENT WAS TREATED FOR THIS CONDITION. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DATE OF INITIAL PROCEDURE. PROCEDURE NAME. WHICH TISSUE WAS THE SUTURE USED ON? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? WAS THE SUTURE PACKAGE REPROCESSED OR STERILIZED? HOW MANY DAYS POST -OP WAS THE SUTURE SPITTING/NON-ABSORPTION ISSUE NOTED? WAS WOUND DEHISCENCE NOTED? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED? PATIENT'S GENDER, AGE, WEIGHT, MEDICAL/ SURGICAL HISTORY & MEDICINES USED? PATIENT'S CURRENT CONDITION. LOT NUMBER OF THE VICRYL SUTURE. ANY REPRESENTATIVE SAMPLES OF THE SAME LOT AVAILABLE FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884832&pc=GAM | 2210968-2018-75865 |  | Malfunction | ETHICON INC. | 2018-09-18 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Appropriate Term/Code Not Available | No Code Available | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OPERATIVELY, THE ABSORBABLE SUTURE WAS NOT ABSORBING, AND THE SUTURE WAS "SPITTING THROUGH THE SKIN". IT IS UNKNOWN HOW OR IF THE PATIENT WAS TREATED FOR THIS CONDITION. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DATE OF INITIAL PROCEDURE. PROCEDURE NAME. WHICH TISSUE WAS THE SUTURE USED ON? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? WAS THE SUTURE PACKAGE REPROCESSED OR STERILIZED? HOW MANY DAYS POST -OP WAS THE SUTURE SPITTING/NON-ABSORPTION ISSUE NOTED? WAS WOUND DEHISCENCE NOTED? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED? PATIENT'S GENDER, AGE, WEIGHT, MEDICAL/ SURGICAL HISTORY & MEDICINES USED? PATIENT'S CURRENT CONDITION. LOT NUMBER OF THE VICRYL SUTURE. ANY REPRESENTATIVE SAMPLES OF THE SAME LOT AVAILABLE FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884889&pc=GAM | 2210968-2018-75864 |  | Malfunction | ETHICON INC. | 2018-09-18 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Appropriate Term/Code Not Available | No Code Available | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OPERATIVELY, THE ABSORBABLE SUTURE WAS NOT ABSORBING, AND THE SUTURE WAS "SPITTING THROUGH THE SKIN". IT IS UNKNOWN HOW OR IF THE PATIENT WAS TREATED FOR THIS CONDITION. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DATE OF INITIAL PROCEDURE . PROCEDURE NAME . WHICH TISSUE WAS THE SUTURE USED ON? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? WAS THE SUTURE PACKAGE REPROCESSED OR STERILIZED? HOW MANY DAYS POST -OP WAS THE SUTURE SPITTING/NON-ABSORPTION ISSUE NOTED? WAS WOUND DEHISCENCE NOTED? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED? PATIENT'S GENDER, AGE, WEIGHT, MEDICAL/ SURGICAL HISTORY & MEDICINES USED? PATIENT'S CURRENT CONDITION . LOT NUMBER OF THE VICRYL SUTURE . ANY REPRESENTATIVE SAMPLES OF THE SAME LOT AVAILABLE FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7881324&pc=GAM | 2210968-2018-75859 |  | Malfunction | ETHICON INC. | 2018-09-17 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 UND | Component Incompatible | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS FOUND THAT THE PACKAGES OF THEM WERE NOT FOR THE PRODUCT EXPECTED, BUT FOR A DIFFERENT PRODUCT IN THE BOX. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7822200&pc=GAM | 2210968-2018-75461 | 2018-07-26 | Malfunction | ETHICON INC. | 2018-08-28 | GAM | VICRYL SUTURE 27"(70CM) 3-0 VIO | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, BEFORE USE, IT AS NOTED THAT THERE WAS A DOUBLE-ARMED SUTURE IN THE PACKAGE INSTEAD OF A SINGLE-ARMED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824663&pc=GAM | 2210968-2018-75501 | 2017-12-01 | Malfunction | ETHICON INC. | 2018-08-28 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CESAREAN SECTION ON (B)(6) 2017 AND SUTURE WAS USED. WHILST THE SURGEON WAS CLOSING A PATIENT HE NOTICED A SMALL METAL POINT ATTACHED TO THE TOP OF A SUTURE. THIS SUTURE WAS NOT USED BUT COULD HAVE RESULTED IN THE METAL PIECE BEING LEFT INSIDE THE PATIENT HAD IT NOT HAVE BEEN NOTICED. SUTURE REMOVED FROM SET AND SET ASIDE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. INVESTIGATION SUMMARY: A FAILED SUTURE NEEDLE WAS SUBMITTED FOR EVALUATION. A FRACTURE WAS OBSERVED AT THE TIP OF THE NEEDLE. ONE SIDE OF THE NEEDLE WAS RECEIVED, THE MATING FRACTURE SURFACE WAS NOT PROVIDED FOR THIS EVALUATION. THE FRACTURE SURFACES WERE EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS PREDOMINANTLY COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. MECHANICAL DAMAGE OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDES ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE TIP OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. THE NEEDLE BREAKAGE WAS CONFIRMED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7769632&pc=GAM | 2210968-2018-75048 | 2018-07-18 | Malfunction | ETHICON INC. | 2018-08-09 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 VIO | Device Misassembled During Manufacturing /Shipping | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, THE STAFF FOUND THAT THE NEEDLE IN THE FOURTH SLOT SEEMED LIKE 3/8 CIRCLE AND WRONG SHAPE. THE SUTURE WAS NOT USED ON THE PATIENT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: PRODUCT COMPLAINT # PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. AN EMPTY OPENED FOIL, A PAPER LID, A WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS OF PRODUCT WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF NEEDLES, NO INCORRECT COMPONENT WAS OBSERVED. HOWEVER, THE NEEDLE OF SLOT # 4 WAS NOTED THAT THE CURVE OF THE NEEDLE IS NOT SMOOTH IN COMPARISON WITH THE NEEDLE SPECIFICATION AND IS OUTSIDE OF CURVE TOLERANCE. IN ADDITION, THE SUTURE WAS EXAMINED AND NO DEFECTS OR DAMAGED WERE FOUND. THE REMAINING NEEDLES WERE EXAMINED AND NO DEFECTS OR DAMAGED WERE FOUND. THE SUTURES WERE EXAMINED AND NO DEFECTS OR DAMAGED WERE FOUND. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, AN INCORRECT CURVATURE WAS OBSERVED ON THE NEEDLE THAT AT FIRST SEEMED TO BE AN INCORRECT COMPONENT WHICH CORRESPONDS TO THE REPORTED COMPLAINT. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7751752&pc=GAM | 2210968-2018-74915 | 2018-07-09 | Malfunction | ETHICON INC. | 2018-08-06 | GAM | VICRYL PLUS SUTURE 18"(45CM) 0 VIO | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS NOTED THAT THE SUTURE WAS DEGRADED WHEN OUT OF PACKAGE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. IT WAS REPORTED PACKAGING INTEGRITY. AN EMPTY OPENED FOIL AND A WINDING FORMER WITH EIGHT PARKED NEEDLE/SUTURE COMBINATIONS FROM SLOT # 1 TO 8 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, THE SWAGE AND ATTACHMENT AREA OF THE EIGHT NEEDLES WERE AS EXPECTED. THE SUTURES WERE EXAMINED AND NO DEFECTS OR SUTURE DEGRADE COULD BE OBSERVED. IN ADDITION, A FUNCTIONAL TEST WAS PERFORMED ON THE SAMPLES AND THE SAMPLES MET THE MINIMUM INDIVIDUAL REQUIREMENTS. ALSO, THE FOIL WAS EXAMINED FOR VISUAL INSPECTION AND SHOWED MULTIPLES WRINKLES AND HOLES IN CAVITY ON THE TOP AND BOTTOM FOIL PACKAGE DUE TO EXCESSIVE MANIPULATION. ALSO, THE FOIL PRESENT DELAMINATION ON THE OUTSIDE (TOP AND BOTTOM SIDE) OF THE FOIL. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF PACKAGING INTEGRITY, IS A HOLE IN CAVITY; DUE TO THE IMPROPER HANDLING OF PACKAGE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THERE ANY HOLE, TEAR, OR PUNCTURE IN THE PACKAGE THAT COMPROMISED STERILITY? WAS THERE ANY OPEN OR INCOMPLETE SEAL?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7751762&pc=GAM | 2210968-2018-74914 | 2018-07-09 | Malfunction | ETHICON INC. | 2018-08-06 | GAM | VICRYL PLUS SUTURE 18"(45CM) 0 VIO | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS NOTED THAT THE SUTURE WAS DEGRADED WHEN OUT OF PACKAGE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. IT WAS REPORTED PACKAGING INTEGRITY. AN EMPTY OPENED FOIL, A PAPER LID, A WINDING FORMER WITH SEVEN PARKED NEEDLES FROM SLOT # 1 TO 6 AND 8 AND SEVERAL SUTURE PIECES WERE RETURNED FOR ANALYSIS. THIS PRODUCT CODE REQUIRES 8 STRANDS PER PACKET. DURING THE VISUAL INSPECTION OF SAMPLE, THE SWAGE AND ATTACHMENT AREA OF THE SEVEN NEEDLES WERE AS EXPECTED. THE SUTURE COULD NOT BE DISPENSED SINCE IT HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED. ALSO, THE FOIL WAS EXAMINED FOR VISUAL INSPECTION AND SHOWED MULTIPLES WRINKLES AND HOLES IN CAVITY ON THE TOP AND BOTTOM FOIL PACKAGE DUE TO EXCESSIVE MANIPULATION. ALSO, THE FOIL PRESENT DELAMINATION ON THE OUTSIDE (TOP AND BOTTOM SIDE) OF THE FOIL. IN ADDITION, FOR ALL VICRYL PLUS SUTURES, THE PRODUCTS HAVE A CYCLE LIFE OF FIVE YEARS, AND DUE TO THE MANUFACTURE DATE OF THIS, THE EXPIRATION DATE WAS ON (B)(4) 2018. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF PACKAGING INTEGRITY, IS A HOLE IN CAVITY; DUE TO THE IMPROPER HANDLING OF PACKAGE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THERE ANY HOLE, TEAR, OR PUNCTURE IN THE PACKAGE THAT COMPROMISED STERILITY? WAS THERE ANY OPEN OR INCOMPLETE SEAL?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7734665&pc=GAM | 2210968-2018-74748 | 2018-01-17 | Malfunction | ETHICON, INC. | 2018-07-31 | GAM | VICRYL PLUS SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 338 | Event Description: THIS REPORT SUMMARIZES 338 INITIAL MALFUNCTION EVENTS, INCLUDING 119 SUTURE BREAKAGE, 71 NEEDLE BREAKAGE AND 148 NEEDLE PULL-OFF THAT OCCURRED INTRAOPERATIVELY. IT ALSO SUMMARIZES 40 FOLLOW UP MALFUNCTION EVENTS, INCLUDING 4 SUTURE BREAKAGE, 14 NEEDLE BREAKAGE AND 22 NEEDLE PULL-OFF THAT OCCURRED INTRAOPERATIVELY. Manufacturer Narrative: DATE SENT TO THE FDA: 7/31/2018 . INITIAL SUTURE BREAKAGE MALFUNCTION EVENTS: 119, INITIAL SUTURE BREAKAGE DEVICES RETURNED: 121· 48 - USE ERROR CAUSED OR CONTRIBUTED TO EVENT · 67 - NO DEVICE FAILURE · 1- MANUFACTURING DEFICIENCY · 5 - UNABLE TO CONFIRM COMPLAINT , INITIAL NEEDLE BREAKAGE MALFUNCTION EVENTS: 71 , INITIAL NEEDLE BREAKAGE DEVICES RETURNED: 90 · 12 - USE ERROR CAUSED OR CONTRIBUTED TO EVENT · 66 - NO DEVICE FAILURE · 3 - MANUFACTURING DEFICIENCY · 9 - UNABLE TO CONFIRM COMPLAINT , INITIAL NEEDLE PULL OFF MALFUNCTION EVENTS: 148, INITIAL NEEDLE PULL OFF DEVICES RETURNED: 304 · 20 - USE ERROR CAUSED OR CONTRIBUTED TO EVENT · 241 - NO DEVICE FAILURE · 34 - MANUFACTURING DEFICIENCY · 9 - UNABLE TO CONFIRM COMPLAINT . FOLLOW UP SUTURE BREAKAGE MALFUNCTION EVENTS: 4 , FOLLOW UP SUTURE BREAKAGE DEVICES RETURNED: 22 · 1 - USE ERROR CAUSED OR CONTRIBUTED TO EVENT · 21 - NO DEVICE FAILURE · 0 - MANUFACTURING DEFICIENCY · 0 - UNABLE TO CONFIRM COMPLAINT , FOLLOW UP NEEDLE BREAKAGE MALFUNCTION EVENTS: 14 , FOLLOW UP NEEDLE BREAKAGE DEVICES RETURNED: 37 · 4 - USE ERROR CAUSED OR CONTRIBUTED TO EVENT · 27 - NO DEVICE FAILURE · 3 - MANUFACTURING DEFICIENCY · 3 - UNABLE TO CONFIRM COMPLAINT FOLLOW UP NEEDLE PULL OFF MALFUNCTION EVENTS: 22 FOLLOW UP NEEDLE PULL OFF DEVICES RETURNED: 65 · 1 - USE ERROR CAUSED OR CONTRIBUTED TO EVENT · 57 - NO DEVICE FAILURE · 4 - MANUFACTURING DEFICIENCY · 3 - UNABLE TO CONFIRM COMPLAINT . ALL DEVICES LABELED FOR SINGLE USE; 0 DEVICES REPROCESSED AND RE-USED 0 REMEDIAL ACTION COMPLAINT INFORMATION WILL BE INCLUDED IN COMPLAINT TRENDING THAT IS REVIEWED ON A REGULAR BASIS TO DETERMINE IF FURTHER ACTION IS NECESSARY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7736005&pc=GAM | 2210968-2018-74771 | 2018-07-01 | Malfunction | ETHICON INC. | 2018-07-31 | GAM | VICRYL SUTURE 18"(45CM) 6-0 UND | Component Misassembled | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT CUSTOMER ORDERED THE SUTURE PRODUCT WITH THE UNDERSTANDING THAT IT WOULD COME WITH A CERTAIN SIZE AND TYPE OF NEEDLE (S24). THE BOX THAT THEY RECEIVED DISPLAYS IT HAS A DIFFERENT, LABELED NEEDLE (S14). THE SUTURES INSIDE ALSO DISPLAY S14. THERE WAS NO PATIENT INVOLVEMENT. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 09/18/2018 EVALUATION: A SEALED BOX, AN UNOPENED BOX WITH AN EMPTY OPENED FOIL, AND UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLES, THE INFORMATION OF PRODUCT CODE, LOT NUMBER AND GRAPHIC ON THE LABEL OVERWRAP AND PAPER LID PRINTED WAS COMPARED WITH SYSTEM DATABASE. IN ADDITION, A CHARACTERIZATION WAS PERFORMED. THE CHARACTERIZATION OF THE SAMPLE WAS CONDUCTED, AND THE FOLLOWING WAS OBSERVED: THE SAMPLE WAS MEASURED IN NEEDLE DIAMETER, POINT TYPE, NEEDLE TYPE, NEEDLE RADIUS/CURVATURE AND ANGLE BELONG TO PRODUCT CODE D5264. ALSO, THE SUTURE WAS MEASURED IN DIAMETER SUTURE AND LENGTH AND THE STRUCTURE OF THE STRAND BELONG TO VICRYL SUTURE DIAMETER 6-0¿AND IT IS CORRECT TO THE PRODUCT CODE D5264. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. HOWEVER, WHEN THE CODE WAS REVIEWED ON THE ETHICON WEB PRODUCT CATALOG, IT WAS FOUND TO HAVE A DIFFERENT PRODUCT ON THE PHOTO AND THE CHARACTERISTICS DON¿T CORRESPOND TO THE PHYSICAL SAMPLES RECEIVED. THE PRODUCT SHOW AT THE WEB PAGE BELONGS TO A DIFFERENT PRODUCT. PER THE CONDITION OF THE SAMPLE RECEIVED AND THE INFORMATION OF THE CUSTOMER, THE DIFFERENCE ON THE HOMEPAGE CORRESPONDS TO THE REPORTED COMPLAINT. HOWEVER, THE SAMPLES RECEIVED MEET THE FINISH GOOD REQUIREMENTS. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7725923&pc=GAM | 7725923 | 2018-05-06 | Malfunction | ETHICON, INC. | 2018-07-27 | GAM | VICRYL PLUS | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: SURGEON WAS CLOSING FASCIA WITH 0 VICRYL CT-1 SUTURE, AND THE NEEDLE BROKE. NEEDLE WAS RECOVERED. SUTURE: 0 CT-1, VICRYL PLUS: VCP946, LOT# LM2678, EXP: 10/31/2022.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7715594&pc=GAM | 2210968-2018-74585 |  | Malfunction | ETHICON INC. | 2018-07-24 | GAM | VICRYL SUTURE 27"(70CM) 4-0 UND | Device Misassembled During Manufacturing /Shipping | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. PRIOR TO THE PROCEDURE IT WAS REPORTED THAT THE BOX CONTAINED THE INCORRECT CODE AND LOT NUMBER OF SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PC-000225543 DATE SENT TO THE FDA: 08/01/2018 ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: 1. DID 2 BOXES OF J214H LOT MAM223 CONTAIN VCP311 LOT LPM513 SUTURES? YES Manufacturer Narrative: PRODUCT COMPLAINT # PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. AN OPENED BOX OF PRODUCT CODE (B)(4), LOT MAM223 CONTAINING THIRTY-SIX UNOPENED SAMPLES OF PRODUCT CODE (B)(4), LOT LPM513 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF BOX, A MISMATCHED BETWEEN BOX AND THE INDIVIDUAL FOILS INSIDE WERE OBSERVED. FOR THIS REASON, A CHARACTERIZATION OF SAMPLES WAS CONDUCTED, AND THE FOLLOWING WAS OBSERVED: THIRTY-SIX FOILS WERE EXAMINED AND NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA OF THE NEEDLES WERE AS EXPECTED. ALSO, THE SAMPLES WERE EXAMINED BY NEEDLE CHARACTERISTIC ON POINT TYPE, NEEDLE TYPE AND NEEDLE CURVATURE AND SUTURE CHARACTERISTIC AND THE PRODUCT CONTAINED IN EACH PACKET BELONG TO PRODUCT CODE (B)(4), LOT LPM513. IN ADDITION, THE MANUFACTURING DATE IS ON (B)(4), 2017. THE BOX WITH THE PRODUCT CODE (B)(4), LOT MAM223 WERE EXAMINED AND EXCESSIVE MANIPULATION COULD BE OBSERVED. ALSO, THE MANUFACTURING DATE IS ON (B)(4), 2018. ALSO, THE MANUFACTURED DATES WERE REVIEWED THERE ARE A DIFFERENCE OF A MONTH OF SEPARATION BETWEEN THE MANUFACTURE DATE OF THE REPORTED LOTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO MANUFACTURED DATES AND THAT THE BOX WAS RECEIVED OPENED A MIX OF PRODUCT COULD NOT BE HAPPENED IN OUR MANUFACTURING SITES AND THE ASSIGNABLE CAUSE OF THE ASSEMBLY INCORRECT IMPLANTABLE IS AN IMPROPER HANDLING AND THE REPORTED COMPLAINT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WERE ONLY 2 DEVICES OF THE INCORRECT VICRYL PLUS VCP311 LOT LPM513 SUTURES FOUND IN THE BOX OF VICRYL SUTURE J214H LOT MAM223? NO, VCP311H*36EA (LOT: LPM513) WERE IN THE BOX OF J214H (LOT: MAM223). WAS THE ISSUE NOTED IN A NEWLY OPENED BOX? YES. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID 2 BOXES OF J214H LOT MAM223 CONTAIN VCP311 LOT LPM513 SUTURES?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7715737&pc=GAM | 2210968-2018-74588 |  | Malfunction | ETHICON INC. | 2018-07-24 | GAM | VICRYL SUTURE 27"(70CM) 4-0 UND | Device Misassembled During Manufacturing /Shipping | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. PRIOR TO THE PROCEDURE IT WAS REPORTED THAT THE BOX CONTAINED THE INCORRECT CODE AND LOT NUMBER OF SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PC-000225543 DATE SENT TO THE FDA: 08/01/2018 ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: 1. DID 2 BOXES OF J214H LOT MAM223 CONTAIN VCP311 LOT LPM513 SUTURES? YES Manufacturer Narrative: PRODUCT COMPLAINT # PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. AN OPENED BOX OF PRODUCT CODE (B)(4), LOT MAM223 CONTAINING THIRTY-FIVE UNOPENED SAMPLES OF A DIFFERENT PRODUCT CODE (B)(4) , LOT LPM513 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF BOX, A MISMATCHED BETWEEN BOX AND THE INDIVIDUAL FOILS INSIDE WERE OBSERVED. FOR THIS REASON, A CHARACTERIZATION OF SAMPLES WAS CONDUCTED, AND THE FOLLOWING WAS OBSERVED: THIRTY-FIVE FOILS WERE EXAMINED AND NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA OF THE NEEDLES WERE AS EXPECTED. ALSO, THE SAMPLES WERE EXAMINED BY NEEDLE CHARACTERISTIC ON POINT TYPE, NEEDLE TYPE AND NEEDLE CURVATURE AND SUTURE CHARACTERISTIC AND THE PRODUCT CONTAINED IN EACH PACKET BELONG TO PRODUCT CODE (B)(4), LOT LPM513. IN ADDITION, THE MANUFACTURING DATE IS ON (B)(4), 2017. THE BOX WITH THE PRODUCT CODE (B)(4), LOT MAM223 WERE EXAMINED AND EXCESSIVE MANIPULATION COULD BE OBSERVED. ALSO, THE MANUFACTURING DATE IS ON (B)(4), 2018. ALSO, THE MANUFACTURED DATES WERE REVIEWED THERE ARE A DIFFERENCE OF A MONTH OF SEPARATION BETWEEN THE MANUFACTURE DATE OF THE REPORTED LOTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO MANUFACTURED DATES AND THAT THE BOX WAS RECEIVED OPENED A MIX OF PRODUCT COULD NOT BE HAPPENED IN OUR MANUFACTURING SITES AND THE ASSIGNABLE CAUSE OF THE ASSEMBLY INCORRECT IMPLANTABLE IS AN IMPROPER HANDLING AND THE REPORTED COMPLAINT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WERE ONLY 2 DEVICES OF THE INCORRECT VICRYL PLUS VCP311 LOT LPM513 SUTURES FOUND IN THE BOX OF VICRYL SUTURE J214H LOT MAM223? NO, VCP311H*36EA (LOT: LPM513) WERE IN THE BOX OF J214H (LOT: MAM223) WAS THE ISSUE NOTED IN A NEWLY OPENED BOX? YES. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID 2 BOXES OF J214H LOT MAM223 CONTAIN VCP311 LOT LPM513 SUTURES?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7677763&pc=GAM | 2210968-2018-74236 | 2018-06-14 | Malfunction | ETHICON INC. | 2018-07-11 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Expiration Date Error | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2018 SUTURE WAS RECEIVED BY A DISTRIBUTOR AND TWO FOILS WITH THE SAME LOT NUMBER WERE OBSERVED HAVING TWO DIFFERENT MANUFACTURING AND EXPIRATION DATES. ONE FOIL PACK HAD MANUFACTURING DATE 10/2017 AND EXPIRATION DATE 09/2022 WHEREAS ANOTHER FOIL PACK HAD A DIFFERENT MANUFACTURING DATE 08/2017 AND EXPIRY DATE 07/2022. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ALL THE RETAIN SAMPLES(24EA) OF INCIDENT CODES AND LOT NUMBER WERE RETRIEVED FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR MANUFACTURING AND EXPIRY DATE 06EA OUT OF 24 EA RETRIEVED FOR ANALYSIS HAS A MANUFACTURING DATE - (B)(4) 2017 & EXPIRY DATE - (B)(4) 2022. THIS COMPLAINT IS A MISMATCH PACKAGING OF SAME PRODUCT CODE, LOT NUMBER WITH DIFFERENT MANUFACTURING DATE AND EXPIRY DATE. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHERE WAS THE SALES REP WHEN HE FOUND THE ISSUE? SALES REP WAS CALLED BY THE SUB-DEALER WHEN THE STOCK WAS SUPPLIED THIS IS WHEN THE ISSUE WAS HIGHLIGHTED BY THE SUB-DEALER. HOW HE FOUND THE ISSUE? WHEN HE TRIED TO FEED THE INVENTORY IN HIS STOCK, SAME BATCH WITH DIFFERENT EXPIRY THE SOFTWARE RAISED AN ALARM THIS IS HOW THE SUB-DEALER FOUND OUT REGARDING THE EXPIRY MISMATCH. WAS THE CUSTOMER INVOLVED? CUSTOMER YES SUB-DEALER WAS THE CUSTOMER IN THIS CASE. DID THE SALES REP IDENTIFY THE ISSUE AT A CUSTOMER'S LOCATION? YES THE ISSUE WAS IDENTIFIED AT CUSTOMER LOCATION. WAS THE AFFECTED FOIL PACKAGE WITHIN J&J CONTROL AT ALL TIMES? YES, IT WAS UNDER J&J TEAM CONTROL AT ALL TIMES, MATERIAL WAS SHIPPED FROM OUR AUTHORIZED DISTRIBUTOR TO A SUB-DEALER. DID BOTH FOILS COME FROM THE SAME BOX OF PRODUCT? YES THE FOILS CAME FROM THE SAME BOX. DID THE LABEL ON THE OUTER BOX REFLECT THE INFORMATION FROM THE BATCH MANUFACTURING RECORD OF THE LOT. YES THE LABEL ON THE OUTER BOX MENTIONED THE SAME BATCH MANUFACTURING RECORD AS PER ONE OF THE FOILS INSIDE THE BOX. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7669834&pc=GAM | 2210968-2018-74176 | 2018-05-28 | Malfunction | ETHICON INC. | 2018-07-09 | GAM | VICRYL SUTURE 27"(70CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT AN ANIMAL UNDERWENT A CANINE C-SECTION PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. ON (B)(6) 2018, THE ANIMAL EXPERIENCED POST-OPERATIVE SUTURE BREAKAGE. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7669836&pc=GAM | 2210968-2018-74177 | 2018-05-28 | Malfunction | ETHICON INC. | 2018-07-09 | GAM | VICRYL SUTURE 27"(70CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT AN ANIMAL UNDERWENT A CANINE C-SECTION PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. ON (B)(6) 2018, THE ANIMAL EXPERIENCED POST-OPERATIVE SUTURE BREAKAGE. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018 . THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7661550&pc=GAM | 2210968-2018-74034 |  | Malfunction | ETHICON INC. | 2018-07-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 UND | Component Incompatible | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE ACCOUNT ORDERED VCP259H WHICH IS A 2-0 COATED VICRYL PLUS ON A CT-1, AND INSTEAD THEY RECEIVED AN 0 SUTURE ON A CP NEEDLE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY THE ISSUE "INCORRECT PRODUCT WAS MAILED": - WHAT WAS THE PRODUCT CODE PRINTED IN THE PACKAGE RECEIVED? WAS IT VCP259H OR A DIFFERENT PRODUCT CODE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7597416&pc=GAM | 2210968-2018-73444 | 2018-01-01 | Malfunction | ETHICON INC. | 2018-06-13 | GAM | VICRYL SUTURE 36"(90CM) 0 UND | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ON AN UNKNOWN DATE, THAT THE DISTRIBUTOR RECEIVED A BOX OF SUTURE AND THE BOX WAS CRUSHED. THE DISTRIBUTOR STATED THAT THE PACKAGES INSIDE THE BOX (PRIMARY PACKAGING) WERE DAMAGED AND COULD POSSIBLE AFFECT STERILITY. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: A SEALED BOX OF PRODUCT CODE J958, LOT MB2685 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE BOX, NO DAMAGED ON THE FILM WAS NOTED; BUT, ONE SIDE OF THE BOX WAS SLIGHT CRUSHED AND WRINKLED. HOWEVER, THE INTEGRITY OF PRODUCT WAS NOT COMPROMISED. IN ADDITION, CANNOT BE DETERMINED THE CONDITION OF THE STORAGE, TRANSPORTATION AND HANDLING OF THE SAMPLE DUE TO THE SAMPLE CONDITION IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT OF: ¿THE BOX WAS CRUSHED¿. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7581150&pc=GAM | 2210968-2018-73370 |  | Malfunction | ETHICON INC. | 2018-06-08 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 UND | Appropriate Term/Code Not Available | Foreign Body Reaction; No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CUTANEOUS WOUND REPAIR FOLLOWING AN EXCISON ON THE RIGHT UPPER ARM IN (B)(6) 2016 AND DEEP SUTURES WERE PLACED. THE PATIENT COMPLAINED OF AN ITCHY HYPERTROPHIC SCAR WHICH WAS TREATED IN THE PAST YEAR WITH INTRALESIONAL KENALOG. THE SCAR IS NO LONGER HYPERTROPHIC. THE DOCTOR REPORTS EXAMINING THE PATIENT LAST WEEK AND THE PATIENT HAS 5 SUTURES THAT APPEAR TO HAVE NOT ABSORBED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY PATIENT CONDITION PRESENT THAT COULD CONTRIBUTE TO THE SUTURE COMING OUT OF THE PATIENT ? THE SUTURE DID NOT COME OUT OF THE PATIENT UNTIL THE PAST MONTH (20 MONTHS AFTER THE SURGERY). IT BECAME APPARENT AFTER THE HYPERTROPHIC AND ITCHY SCAR WAS INJECTED WITH KENALOG TO THIN IT OUT A FEW MONTHS AGO. DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT? NO. HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? BURIED VERTICAL MATTRESS. INTERRUPTED. HAS THERE BEEN ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED OR ANY PLANNED RELATING TO THE SUTURE NOT ABSORBED? DO YOU HAVE ANY SUGGESTIONS? NOTHING HAS BEEN DONE SINCE SHE PRESENTED TO THE OFFICE SURPRISED TO SEE THERE WAS SUTURE IN THE SCAR. WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? I DO NOT KNOW WHY VICRYL SUTURE IS STILL INTACT IN HER SKIN 20 MONTHS AFTER THE SURGERY. THAT IS WHY I AM CONTACTING ETHICON / J&J. PATIENT PRE-EXISTING MEDICAL CONDITIONS NONE. WHAT IS THE PATIENT CURRENT STATUS? UPSET ABOUT THE SUTURE MATERIAL SHOWING UP IN HER SCAR 20 MONTHS AFTER THE SURGERY. OTHERWISE SHE IS FINE FROM A MEDICAL STANDPOINT. WILL ANY PRODUCT BE RETURNED FOR ANALYSIS? NO. SPECIFIC PRODUCT CODE / LOT NUMBER OF VICRYL SUTURE USED. DO NOT KNOW AS THEY WERE ORDERED AND USED NEARLY 2 YEARS AGO. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THERE ANY PATIENT CONDITION PRESENT THAT COULD CONTRIBUTE TO THE SUTURE COMING OUT OF THE PATIENT ? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HAS THERE BEEN ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED OR ANY PLANNED RELATING TO THE SUTURE NOT ABSORBED? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? PATIENT PRE-EXISTING MEDICAL CONDITIONS WHAT IS THE PATIENT CURRENT STATUS? WILL ANY PRODUCT BE RETURNED FOR ANALYSIS? SPECIFIC PRODUCT CODE / LOT NUMBER OF VICRYL SUTURE USED
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7580217&pc=GAM | 2210968-2018-73363 |  | Malfunction | ETHICON INC. | 2018-06-07 | GAM | VICRYL SUTURE 36"(90CM) 1 UND | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY THE DISTRIBUTOR THAT PRODUCT CODE JB947, PRINTED ON THE OUTSIDE OF THE BOX, DOESN'T RELATE TO THE LOT NUMBER PRINTED ON THE BOX. THE DEVICE WAS NOT USED ON THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # ==> PC-000191689 DATE SENT TO THE FDA: 07/10/2018 THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. EVALUATION: A SEALED BOX WITH UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, THE FRONT LABEL OF THE BOX IS PRESENT WITH THE PRODUCT CODE AND MATCHES. HOWEVER, THE MFG (MANUFACTURING) BARCODE PRESENT IS CORRECT FOR THIS PRODUCT. IT IS A SPECIAL CODE, SINCE; IT WAS CREATED WHEN THE SYSTEM HAD A 5-DIGIT LIMIT AND ACCORDING TO REQUIREMENT. ACCORDING THE SAMPLE CONDITION, NO PACKAGING LABELING IDENTIFICATION WAS OBSERVED. SINCE, THE GRAPHICS ON THE BOX WERE ACCORDING ETHICON REQUIREMENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7536994&pc=GAM | 2210968-2018-72990 |  | Malfunction | ETHICON INC. | 2018-05-23 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT UPON OPENING OF THE PACKAGING, SMALL HOLES WERE NOTICED IN THE BLISTER, POSSIBLY CAUSED BY THE NEEDLES. THE ISSUE WAS NOTICED PRIOR TO ANY USE IN A PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: DURING THE VISUAL INSPECTION OF SAMPLE, A PIN HOLE WAS OBSERVED ON BACK FOIL PACKET. THE FOIL WAS OPEN AND THE NEEDLES WERE PLACE CORRECTLY, HOWEVER, MARKS INTO THE FOIL WERE NOTED CAUSED BY THE CORNER THE FOLDER, THIS CONDITION CAUSES DAMAGE FROM INSIDE TO OUTSIDE ON THE FOIL AND THE STERILE BARRIER WAS COMPROMISED. PER THE CONDITION OF THE SAMPLE RECEIVED THE ASSIGNABLE CAUSE THE PACKAGING INTEGRITY WAS CAUSED BY PIN HOLES. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE REPORTED ISSUE NOTICED UPON OPENING A BRAND NEW PACKAGE? AS DESCRIBED A NEW PACKAGE WAS OPENED. ARE THE ACTUAL SAMPLES BEING RETURNED ? YES. IF NOT, DO YOU HAVE ANY PHOTOS OF THE REPORTED ISSUE? NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7537174&pc=GAM | 2210968-2018-72994 |  | Malfunction | ETHICON INC. | 2018-05-23 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT UPON OPENING OF THE PACKAGING, SMALL HOLES WERE NOTICED IN THE BLISTER WHICH WAS POSSIBLY CAUSED BY THE NEEDLES. THE ISSUE WAS NOTICED PRIOR TO ANY USE IN A PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: DURING THE VISUAL INSPECTION OF SAMPLE, A PIN HOLE WAS OBSERVED ON BACK FOIL PACKET. THE FOIL WAS OPEN AND THE NEEDLES WERE PLACE CORRECTLY, HOWEVER, MARKS INTO THE FOIL WERE NOTED CAUSED BY THE CORNER THE FOLDER, THIS CONDITION CAUSES DAMAGE FROM INSIDE TO OUTSIDE ON THE FOIL AND THE STERILE BARRIER WAS COMPROMISED. PER THE CONDITION OF THE SAMPLE RECEIVED THE ASSIGNABLE CAUSE THE PACKAGING INTEGRITY WAS CAUSED BY PIN HOLES. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE REPORTED ISSUE NOTICED UPON OPENING A BRAND NEW PACKAGE? AS DESCRIBED A NEW PACKAGE WAS OPENED. ARE THE ACTUAL SAMPLES BEING RETURNED? YES. IF NOT, DO YOU HAVE ANY PHOTOS OF THE REPORTED ISSUE? NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7537277&pc=GAM | 2210968-2018-72998 |  | Malfunction | ETHICON INC. | 2018-05-23 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT UPON OPENING OF THE PACKAGING, SMALL HOLES WERE NOTICED IN THE BLISTER, POSSIBLY CAUSED BY THE NEEDLES. THE ISSUE WAS NOTICED PRIOR TO ANY USE IN A PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: DURING THE VISUAL INSPECTION OF SAMPLE, A PIN HOLE WAS OBSERVED ON BACK FOIL PACKET. THE FOIL WAS OPEN AND THE NEEDLES WERE PLACE CORRECTLY, HOWEVER, MARKS INTO THE FOIL WERE NOTED CAUSED BY THE CORNER THE FOLDER, THIS CONDITION CAUSES DAMAGE FROM INSIDE TO OUTSIDE ON THE FOIL AND THE STERILE BARRIER WAS COMPROMISED. PER THE CONDITION OF THE SAMPLE RECEIVED THE ASSIGNABLE CAUSE THE PACKAGING INTEGRITY WAS CAUSED BY PIN HOLES. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE REPORTED ISSUE NOTICED UPON OPENING A BRAND NEW PACKAGE? AS DESCRIBED A NEW PACKAGE WAS OPENED. ARE THE ACTUAL SAMPLES BEING RETURNED? YES. IF NOT, DO YOU HAVE ANY PHOTOS OF THE REPORTED ISSUE? NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7526391&pc=GAM | 2210968-2018-72910 |  | Malfunction | ETHICON INC. | 2018-05-18 | GAM | VICRYL SUTURE 27"(70CM) 0 VIO | Break | Foreign Body In Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT HYSTERECTOMY PROCEDURE ON UNKNOWN DATE IN 2018 AND SUTURE WAS USED. A VERY SMALL TIP OF THE NEEDLE BROKE OFF THE NEEDLE WHEN PASSED THROUGH TISSUE. THE TIP OF THE NEEDLE COULD NOT BE RETRIEVED AND X-RAY CONFIRMED THE TIP IS IN THE VAGINAL CUFF. THE SURGEON HAS ELECTED NOT TO REMOVE THE SMALL NEEDLE TIP AND PATIENT IS ASYMPTOMATIC. THE PATIENT WILL BE FOLLOWED ANNUALLY WITH ULTRASOUND. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW LARGE (IN CM) IS THE SIZE OF THE NEEDLE IN THE VAGINAL CUFF? THE SIZE OF THE NEEDLE TIP WAS ESTIMATED TO BE 1 OR 2MM. DOES THE SURGEON PLAN TO REMOVE THE NEEDLE PIECE IN THE FUTURE? NO. WHAT IS THE SURGEONS OPINION REGARDING THE NEEDLE LOCATION AND CONSEQUENCE TO THE PATIENT? THE SURGEON WILL MONITOR THE PATIENT ANNUALLY THROUGH IMAGING. DO YOU HAVE ANY PRODUCT FOR EVALUATION? NO PRODUCT IS AVAILABLE FOR EVALUATION. WHAT IS THE CURRENT CONDITION OF THE PATIENT? THE PATIENT IS DISCHARGED AND ASYMPTOMATIC. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. UPON ADDITIONAL INFORMATION REVIEW, THIS MEDWATCH REPORT 2210968-2018-72910 HAS BEEN UPDATED TO MALFUNCTION. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PROCEDURE NAME: HYSTERECTOMY. LOT NUMBER: UNKNOWN. LOCATION AND SIZE OF THE RETAINED NEEDLE TIP; X-RAY/ULTRA SOUND REPORT?: LOCATION WAS REPORTED AS VAGINAL CUFF. SIZE OF NEEDLE TIP IS UNKNOWN, AND WAS REPORTED AS A VERY SMALL PART OF THE NEEDLE TIP, NOT NOTICED UNTIL THE SCRUB NURSE TRIED TO PLACE THE NEEDLE IN THE DISCARDED NEEDLE COUNT TRAY. THE PATIENT WAS X-RAYED TO CONFIRM THE NEEDLE TIP WAS PRESENT. ANY ADDITIONAL TISSUE DISSECTION OR DAMAGE DUE TO SEARCHING FOR THE NEEDLE TIP: NO ADDITIONAL DISSECTION OR DAMAGE WAS REPORTED. IT WAS REPORTED THAT THE SURGEON ELECTED TO LEAVE THE NEEDLE TIP IN SITU AND PERFORM AN ANNUAL ULTRASOUND TO MONITOR THE PATIENT. ANY OTHER ADVERSE PATIENT EVENT? IF YES, EXPLAIN AND INFORM WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED: NO OTHER ADVERSE EFFECTS WERE REPORTED AT THIS TIME. PATIENT INITIALS, GENDER, AGE, DATE OF BIRTH, WEIGHT.: UNKNOWN. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW LARGE (IN CM) IS THE SIZE OF THE NEEDLE IN THE VAGINAL CUFF? DOES THE SURGEON PLAN TO REMOVE THE NEEDLE PIECE IN THE FUTURE? WHAT IS THE SURGEONS OPINION REGARDING THE NEEDLE LOCATION AND CONSEQUENCE TO THE PATIENT? DO YOU HAVE ANY PRODUCT FOR EVALUATION? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7490259&pc=GAM | 7490259 | 2018-04-02 | Malfunction | ETHICON, INC. | 2018-05-07 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break; Detachment Of Device Component | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: DURING SURGERY ON A PATIENT A 4-0 VICRYL WAS BEING USED WHEN THE NEEDLE BROKE IN TWO PIECES. ONE PIECE WAS STILL ATTACHED TO THE THREAD AND OTHER PIECE WAS REMOVED BY THE DOCTOR. MANUFACTURER RESPONSE FOR SYNTHETIC ABSORBABLE VICRYL SUTURE WITH NEEDLE, COATED VICRYL UNDYED BRAIDED 4-0 P-3 (PER SITE REPORTER). SENT EMAIL TO CUSTOMER SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488314&pc=GAM | 2210968-2018-72612 | 2018-04-08 | Malfunction | ETHICON INC. | 2018-05-04 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PRINTING ON THE PACKAGING WAS ILLEGIBLE. NO PATIENT INVOLVEMENT WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: AN UNOPENED SAMPLE WAS RETURNED FOR ANALYSIS. THERE WERE NO DEFECTS FOUND ON THE PACKAGE DURING THE VISUAL INSPECTION AND THE PRINTING INFORMATION IS CORRECT AND LEGIBLE. IN ADDITION, THE PRINT INFORMATION ON THE FOIL PACKET WAS COMPARISON WITH SYSTEM AND WAS CORRECT AND MET THE ETHICON REQUIREMENTS. A PICTURE WAS INCLUDED INTO THE COMPLAINT AND DURING THE VISUAL INSPECTION A SAMPLE OF PRODUCT CODE (B)(4) AND A SAMPLE OF PRODUCT CODE (B)(4) WERE COMPARISON AND ONLY WAS OBSERVED THAT THE INK OF THE CODE IS DARKER THEN THE CODE, BUT THIS CONDITION IS NOT CONSIDERED A DEFECT, SINCE THE INFORMATION IN BOTH PACKETS IS LEGIBLE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, NO ILLEGIBLE LABELING IDENTIFICATION WAS OBSERVED. Manufacturer Narrative: ADDITIONAL INFORMATION REQUESTED AND FOLLOWING WAS OBTAINED: WHAT EXACTLY IS ILLEGIBLE ON THE PACKAGING? THE PRINTING SEEMS NOT THE SAME AS USUAL AS CUSTOMER SAID. WAS THERE AN ISSUE WITH SCANNING THE BARCODE? NOT REPORTED. WAS THERE SOMETHING MISSING FROM THE LABEL? NO. Manufacturer Narrative: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT EXACTLY IS ILLEGIBLE ON THE PACKAGING? WAS THERE AN ISSUE WITH SCANNING THE BARCODE? WAS THERE SOMETHING MISSING FROM THE LABEL?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7477182&pc=GAM | 2210968-2018-72501 | 2018-04-02 | Malfunction | ETHICON INC. | 2018-05-01 | GAM | VICRYL SUTURE 18"(45CM) 4-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE ON (B)(6) 2018, SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE IN TWO PIECES. ONE PIECE WAS STILL ATTACHED TO THE THREAD AND THE OTHER PIECE WAS REMOVED BY THE SURGEON. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: A FAILED SUTURE NEEDLE WAS SUBMITTED FOR EVALUATION. THE COMPONENT WAS IDENTIFIED. IT WAS REQUESTED TO ASSESS THE FRACTURE MODE OF THE FAILURE. A FRACTURE WAS OBSERVED AT THE SUTURE ATTACHMENT OF THE NEEDLE. THE NEEDLE WAS RECEIVED IN TWO PIECES, ONLY ONE OF THE MATING FRACTURE SURFACES WAS EXAMINED FOR THIS EVALUATION. A MICROSCOPE WAS USED TO EXAMINE THE FRACTURE SURFACES AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACES WERE EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. MECHANICAL DAMAGE OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDES ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE ATTACHMENT AREA OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE NEEDLE PIECE FALL IN THE PATIENT? YES. WERE X-RAYS TAKEN? RESULTS- LOCATION AND SIZE OF NEEDLE PIECE NO. WAS THE NEEDLE PIECE REMOVED FROM THE PATIENT? YES. WAS NEEDLE PIECE REMOVED DURING SAME PROCEDURE? YES. WAS ADDITIONAL TISSUE DISSECTION OR DAMAGE DUE TO SEARCHING/REMOVAL OF THE NEEDLE PIECE FORM THE PATIENT? NO. HOW WAS CASE COMPLETED? YES. ANY OTHER ADVERSE PATIENT CONSEQUENCES? IF YES EXPLAIN WITH INFORMATION ON WHAT. MEDICAL/SURGICAL TREATMENT WAS PROVIDED? NO. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7463956&pc=GAM | 2210968-2018-72310 |  | Malfunction | ETHICON INC. | 2018-04-26 | GAM | VICRYL PLUS SUTURE 18"(45CM) 0 UND | Break | No Consequences Or Impact To Patient |  |  | 385 | Event Description: THIS REPORT SUMMARIZES 385 MALFUNCTION EVENTS, INCLUDING 142 SUTURE BREAKAGE, 82 NEEDLE BREAKAGE AND 164 NEEDLE PULL-OFF THAT OCCURRED INTRAOPERATIVELY. Manufacturer Narrative: SUTURE BREAKAGE MALFUNCTION EVENTS: 142. SUTURE BREAKAGE DEVICES RETURNED: 135, 18 - USE ERROR CAUSED OR CONTRIBUTED TO EVENT, 110 - NO DEVICE FAILURE, 1 - MANUFACTURING DEFICIENCY, 6 - UNABLE TO CONFIRM COMPLAINT. NEEDLE BREAKAGE MALFUNCTION EVENTS: 82. NEEDLE BREAKAGE DEVICES RETURNED: 121, 17 - USE ERROR CAUSED OR CONTRIBUTED TO EVENT, 93 - NO DEVICE FAILURE, 3 - MANUFACTURING DEFICIENCY, 8 - UNABLE TO CONFIRM COMPLAINT. NEEDLE PULL OFF MALFUNCTION EVENTS: 164. NEEDLE PULL OFF DEVICES RETURNED: 192, 10 - USE ERROR CAUSED OR CONTRIBUTED TO EVENT, 154 - NO DEVICE FAILURE, 20 - MANUFACTURING DEFICIENCY, 8 - UNABLE TO CONFIRM COMPLAINT. ALL DEVICES LABELED FOR SINGLE USE; 0 DEVICES REPROCESSED AND RE-USED 0 REMEDIAL ACTION COMPLAINT INFORMATION IS TRENDED ON A REGULAR BASIS TO DETERMINE IF FURTHER INVESTIGATION IS WARRANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7449809&pc=GAM | 2210968-2018-72239 | 2018-03-21 | Malfunction | ETHICON INC. | 2018-04-22 | GAM | VICRYL VIO 70CM M3 USP2/0 SGLE ARMED SH PLUS | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT A STERILIZATION PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. POST-OPERATIVELY, THE SUTURE BROKE. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. AN EMPTY OPENED BOX AND UNOPENED SAMPLES OF PRODUCT CODE WERE RETURNED FOR ANALYSIS. THE ISSUE SAMPLE WAS NOT RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA OF NEEDLES WERE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG THE STRAND AND NO DEFECTS, DAMAGED OR ANOMALIES WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE RECEIVED SAMPLES, NO BREAKAGE WAS FOUND ON THE SUTURES AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7445704&pc=GAM | 2210968-2018-72198 | 2018-03-21 | Malfunction | ETHICON INC. | 2018-04-19 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A BONE SURGERY PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE OFF. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: DID THE NEEDLE PULL OFF FROM SUTURE THREAD OR NEEDLE BROKE? :- NO BREAKAGE NEEDLE VERIFY IF THE ISSUE OCCURRED DURING USE ON PATIENT? IF NO, EXPLAIN. :- NO DID THE NEEDLE OR NEEDLE PIECE(S) FALL IN PATIENT'S BODY? NO WAS IT REMOVED SUCCESSFULLY? SAME PROCEDURE? :- YES ANY ADDITIONAL TISSUE DISSECTION/DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE/NEEDLE PIECE(S)? :- NO ANY OTHER PATIENT CONSEQUENCES? IF YES, WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED. :-NO HOW WAS CASE COMPLETED? :- USED ANOTHER SUTURE. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4). Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ONE OPENED COMPLAINT SAMPLE FOIL PACK ALONG WITH SBF FOLDER, NEEDLE AND SUTURE FOR PRODUCT WAS RETURNED FOR ANALYSIS. THE SUTURE PACKS WERE VISUALLY INSPECTED TO CHECK WITH MENTIONED VOC NEEDLE BREAKAGE, BUT NEEDLE BREAKAGE WAS NOT EVIDENT. AS THE COMPLAINT SAMPLE RECEIVED IN OPEN CONDITION FURTHER INVESTIGATION ON COMPLAINT SAMPLE CANNOT BE PERFORMED EXCEPT VISUAL INSPECTION. THE FOIL PACK WAS VISUALLY INSPECTED UNDER MAGNIFICATION FOR ANY DAMAGE AND SHOWED MULTIPLE WRINKLES OVER THE WHOLE FOIL ALONG WITH SEVERAL PINHOLES AT THE TOP LABEL SIDE AS WELL AS ON BOTTOM CAVITY SIDE OF THE PACKS. RETURNED NEEDLE WAS INSPECTED WITH MAGNIFICATION AND NO DEFECT WAS OBSERVED. FIVE RETAIN SAMPLES OF INCIDENT CODES AND LOT NUMBER WAS RETRIEVED FOR ANALYSIS. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTABLE DEFECTS, BUT NO SUCH DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WERE TESTED FOR VISUAL INSPECTION AS WELL AS DESCRIPTION TEST AND FOUND TO MEET THE SPECIFICATION REQUIREMENT. AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED FOR ANY PROCESS DEVIATION, BUT NO DEVIATION WAS OBSERVED. FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION. FROM THE ABOVE ANALYSIS, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. THE NEEDLES WERE SUPPLIED BY JOHNSON & JOHNSON MEDICAL (B)(4). HENCE JOHNSON & JOHNSON MEDICAL (B)(4) WAS ALSO INVOLVED INTO INVESTIGATION. THE DHR & COA REVIEW INDICATES THAT ALL RECORDS WERE WITHIN SPECIFICATIONS. ALL ACTIVITIES AND QUALITY INSPECTIONS WERE FOUND ACCORDING TO SPECIFICATIONS. NO ANOMALIES DURING THE PROCESSING OF THE BATCH RELATED TO THE EVENT THAT COULD CAUSE PRODUCT FAILURE WAS IDENTIFIED. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THE NEEDLE PULL OFF FROM SUTURE THREAD OR NEEDLE BROKE? VERIFY IF THE ISSUE OCCURRED DURING USE ON PATIENT? IF NO, EXPLAIN. DID THE NEEDLE OR NEEDLE PIECE(S) FALL IN PATIENT'S BODY? WAS IT REMOVED SUCCESSFULLY? SAME PROCEDURE? ANY ADDITIONAL TISSUE DISSECTION/DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE/NEEDLE PIECE(S)? ANY OTHER PATIENT CONSEQUENCES? IF YES, WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED. HOW WAS CASE COMPLETED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7428320&pc=GAM | 7428320 | 2018-03-28 | Malfunction | ETHICON, INC. | 2018-04-13 | GAM | VICRYL PLUS | Break; Detachment Of Device Component | No Consequences Or Impact To Patient |  |  | 1 | Event Description: DURING SURGERY, A VERY SMALL PIECE OF THE END OF AN ATRAUMATIC NEEDLE BROKE OFF INSIDE PATIENT. NO OTHER COMPLICATIONS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7402600&pc=GAM | 2210968-2018-71961 | 2018-01-01 | Malfunction | ETHICON INC. | 2018-04-06 | GAM | VICRYL RAP UND 1 90CM CT PLUS | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT SUTURE WAS USED DURING AN A PATIENT PROCEDURE. IT WAS REPORTED THAT THE SUTURE ABSORBED TOO QUICKLY. THERE WERE NO PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION AVAILABLE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7402685&pc=GEA | 2210968-2018-71965 | 2018-02-25 | Malfunction | ETHICON INC. | 2018-04-06 | GEA | VICRYL*VLT 0 45CM - ENDOLOOP | Break | No Consequences Or Impact To Patient | K925914 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT UNDERWENT AN APPENDICECTOMY PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE OPERATION, BEFORE USE ON PATIENT. ANOTHER SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # : (B)(4). A CANNULA AND THREE SUTURES PIECE OF PRODUCT CODE WERE RETURNED FOR ANALYSIS. THE FOIL PACKAGE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF SAMPLE, THE SUTURES PIECES WERE EXAMINED AND HAD BEGUN \ THE PROCESS OF DISINTEGRATION, SINCE THE SUTURE WAS BROKEN SEVERAL SECTIONS. THE PRODUCT CODE CONTAINS AN ABSORBABLE SUTURE AND AS THE SAMPLE WAS RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED. NO FUNCTIONAL TEST CAN BE PERFORMED DUE TO SAMPLE CONDITION. THE FOIL WAS NOT RETURNED FOR EVALUATION TO DETERMINE IF ANY DAMAGED WAS PRESENT, THAT COULD BE ATTRIBUTABLE TO DEGRADATION OF THE SUTURE. PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF APPEARANCE COSMETIC OR COLOR SUTURE AND BREAKAGE SUTURE, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE SURGICAL TECHNIQUE USED TO SUTURE? APPENDICECTOMY. WHAT WAS THE SIZE OF THE WOUND? NOT PROVIDED. WAS THERE ANY SURGICAL COMPLICATION DUE TO THE RUPTURE OF THE SUTURE? NO, THE PROBLEM WAS DETECTED BEFORE USING IT ON THE PATIENT. ACTION TAKEN TO CORRECT THE SITUATION? A NEW SUTURE WAS USED. WAS THERE ANY POST-SURGICAL COMPLICATION IN THE PATIENT? NO, THE PROBLEM WAS DETECTED BEFORE USING IT ON THE PATIENT, AND USED A NEW SUTURE. WHAT IS THE CURRENT STATUS OF THE PATIENT? RECOVERED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7368704&pc=GAM | 2210968-2018-71639 | 2017-08-22 | Malfunction | ETHICON INC. | 2018-03-24 | GAM | VICRYL PLUS SUTURE 18"(45CM) 4-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: AT THE END OF A LAP CHOLECYSTECTOMY PROCEDURE ON (B)(6) 2017, SUTURE WAS INTENDED FOR USE ON THE SKIN CLOSURE FOR THE SPECIMEN WOUND. HOWEVER, SUTURE WAS FOUND DETACHED FROM THEIR NEEDLE AS SOON AS THE PACKINGS WERE OPENED. THERE WERE NO PATIENT CONSEQUENCES. PRODUCT WAS RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). ACTUAL NEEDLE/SUTURE WAS RETURNED. THE NEEDLE WAS BROKEN IN THE SWAGE AREA. THERE WERE SEVERAL MARKS ON THE BROKEN AREA. A FRACTURE WAS OBSERVED AT THE SUTURE ATTACHMENT OF THE NEEDLE. THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE. THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7366459&pc=GAM | 2210968-2018-71634 |  | Malfunction | ETHICON INC. | 2018-03-23 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED IN A JOURNAL ARTICLE ENTITLED: RECURRENCE MECHANISMS AFTER INGUINAL HERNIA REPAIR BY THE ONSTEP TECHNIQUE: A CASE SERIES AUTHOR(S): S. OBERG, K. ANDRESEN, D. HAUGE2, J. ROSENBERG CITATION: HERNIA (2016) 20:681¿685; DOI 10.1007/S10029-016-1496-4. THE PURPOSE OF THIS STUDY WAS TO INVESTIGATE RECURRENCES AFTER ONSTEP OPERATION FOR INGUINAL HERNIAS, AND THUS SUGGEST METHODS TO IMPROVE THE TECHNIQUE. BETWEEN DEC 2011 AND APR 2015, 565 MALE PATIENTS, AGES 17 TO 19 YEARS, WITH 569 INGUINAL HERNIAS UNDERWENT ONSTEP REPAIR. THERE WERE A TOTAL OF 26 RECURRENCES, 18 OF THESE WAS NOTED ON THE FIRST HALF OF THE OBSERVATIONAL PERIOD, WHILE 8 WERE NOTED IN THE SECOND HALF. THE ONSTEP MESH (POLYSOFT MESH) IS MEDIALLY PLACED THROUGH AN OPENING IN FASCIA TRANSVERSALIS AND LATERALLY, IT IS SUTURED AROUND THE SPERMATIC CORD. AND IN ONE PATIENT, VICRYL SUTURES WERE UTILIZED AROUND THE SPERMATIC CORD. ONE OF THE PATIENTS HAD SUTURES BROKEN AROUND THE SPERMATIC CORD. THE MAIN REASONS FOR RECURRENCE MAY BE BROKEN SUTURES IN THE MESH LATERALLY OR A SUBOPTIMAL SIZE OF THE OPENING IN FASCIA TRANSVERSALIS MEDIALLY, ENABLING THE MESH TO DISPLACE. Manufacturer Narrative: (B)(4). THIS REPORT IS RELATED TO A JOURNAL ARTICLE, THEREFORE NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE INFORMATION IN THE ARTICLE INDICATES THAT TWO PATIENTS HAD SUTURES BROKEN AROUND THE SPERMATIC CORD. DID THE OTHER PATIENT (NOT THE ONE THAT WAS REOPERATED ON) ALSO HAVE VICRYL SUTURES PLACED OR SOME OTHER TYPE OF SUTURE? IF VICRYL SUTURE WAS USED, CAN YOU PROVIDE THE PRODUCT CODE AND LOT NUMBER INVOLVED? WAS THERE ANY SUTURE DEFICIENCY NOTED PRIOR TO / DURING/ POST OPERATIVELY? WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION REQUIRED FOR THE OTHER PATIENT? PLEASE PROVIDE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE, PROCEDURE DATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7368178&pc=GAM | 2210968-2018-71640 | 2017-07-29 | Malfunction | ETHICON INC. | 2018-03-23 | GAM | VICRYL RAP UND 120CM 2-0 V-34 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO AN UNKNOWN PROCEDURE, THE SUTURE WAS NOTICED TO HAVE A BLUNTED TIP. THERE WERE NO PATIENT CONSEQUENCES. PRODUCT WAS RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). ACTUAL SUTURE/NEEDLE WAS RETURNED FOR EVALUATION. THERE WERE SEVERAL MARKS ON THE SWAGE AREA AND BODY AND THE TIP OF THE NEEDLE APPEARS TO BE BY USE OF THE SURGICAL INSTRUMENTS. A FRACTURE WAS OBSERVED AT THE TIP OF THE NEEDLE. THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE, INDICATES THAT THE BREAKAGE OCCURRED AT THE TIP OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7348478&pc=GAM | 2210968-2018-71510 |  | Malfunction | ETHICON INC. | 2018-03-19 | GAM | VICRYL RAP VIO 45CM 6/0 P-1 PRIME | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS NOTED THAT THE UBD ON THE DEVICE WAS NOT MARKED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PC-(B)(4). THERE WAS NO SAMPLE RECEIVED FOR ANALYSIS. ONLY A PICTURE OF THE SAMPLE WAS SENT. UPON VISUAL INSPECTION OF THE PICTURE, A VICRYL SAMPLE WITHOUT THE LOT NUMBER AT THE TAP OF THE FOIL COULD BE OBSERVED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: 1. PLEAS CLARIFY THE ISSUE WITH THE USE BY DATE (UBD)? UBD WAS NOT MARKED ON THE PRODUCT. 2. WAS THE PRODUCT EXPIRED? NO. 3. WAS THERE CONFLICTING INFORMATION REGARDING THE EXPIRATION DATE? NO. Manufacturer Narrative: AN UNOPENED SAMPLE OF PRODUCT CODE V32H, WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, NO DEFECTS WERE FOUND ON THE PACKAGE. HOWEVER, IT WAS OBSERVED THAT THE EXPIRATION DATE AND LOT NUMBER WAS MISSING FROM LABELING ON FOIL. THE MANUFACTURING RECORDS COULDN'T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEAS CLARIFY THE ISSUE WITH THE USE BY DATE (UBD)? WAS THE PRODUCT EXPIRED? WAS THERE CONFLICTING INFORMATION REGARDING THE EXPIRATION DATE? WITHIN THE ATTACHMENT OF THE COMPLAINT WERE PICTURES OF ETHIBOND SUTURES FROM PRODUCT CODE W10B52. WAS THERE ANY QUALITY ISSUE ASSOCIATED WITH W10B52? IF SO, PLEASE PROVIDE DETAILS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7348484&pc=GAM | 2210968-2018-71509 |  | Malfunction | ETHICON INC. | 2018-03-19 | GAM | VICRYL RAP VIO 45CM 6/0 P-1 PRIME | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS NOTED THAT THE UBD ON THE DEVICE WAS NOT MARKED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PC-(B)(4). THERE WAS NO SAMPLE RECEIVED FOR ANALYSIS. ONLY A PICTURE OF THE SAMPLE WAS SENT. UPON VISUAL INSPECTION OF THE PICTURE, A VICRYL SAMPLE WITHOUT THE LOT NUMBER AT THE TAP OF THE FOIL COULD BE OBSERVED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: 1. PLEASE CLARIFY THE ISSUE WITH THE USE BY DATE (UBD)? UBD WAS NOT MARKED ON THE PRODUCT. 2. WAS THE PRODUCT EXPIRED? NO. 3. WAS THERE CONFLICTING INFORMATION REGARDING THE EXPIRATION DATE? NO. Manufacturer Narrative: AN UNOPENED SAMPLE OF PRODUCT CODE V32H, WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, NO DEFECTS WERE FOUND ON THE PACKAGE. HOWEVER, IT WAS OBSERVED THAT THE EXPIRATION DATE AND LOT NUMBER WAS MISSING FROM LABELING ON FOIL. THE MANUFACTURING RECORDS COULDN¿T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEAS CLARIFY THE ISSUE WITH THE USE BY DATE (UBD)? WAS THE PRODUCT EXPIRED? WAS THERE CONFLICTING INFORMATION REGARDING THE EXPIRATION DATE? WITHIN THE ATTACHMENT OF THE COMPLAINT WERE PICTURES OF ETHIBOND SUTURES FROM PRODUCT CODE W10B52. WAS THERE ANY QUALITY ISSUE ASSOCIATED WITH W10B52? IF SO, PLEASE PROVIDE DETAILS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335729&pc=GAM | 2210968-2018-71409 | 2017-09-13 | Malfunction | ETHICON INC. | 2018-03-13 | GAM | VICRYL SUTURE 27"(70CM) 3-0 VIO | Inadequacy of Device Shape and/or Size; Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC ADRENALECTOMY PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. AT THE MOMENT OF OPENING THE PACKAGE OF THE SUTURE THEY REALIZED THAT THE SIZE OF THE NEEDLE DID NOT MATCH THE DESCRIPTION THAT IS IN THE PACKAGE. PATIENT WAS NOT AFFECTED. THIS WAS NOT USED ON THE PATIENT. ANOTHER LIKE PRODUCT WAS USED TO COMPLETE THE PROCEDURE. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: A PARTIALLY DISPENSED NEEDLE-SUTURE COMBINATION INTO OPENED FOIL OF PRODUCT CODE J305, LOT AF2586 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE OPENED SAMPLE, THE NEEDLE DID NOT MATCH WITH THE PRINT INFORMATION ON THE FOIL TO PRODUCT CODE J305 AND NOR WITH THE NEEDLE DRAWING SPECIFICATION. THE PRODUCT CODE J305 MUST CONTAINS A CHANNEL NEEDLE, WIRE DIAMETER 22 (MILS), WITH A VICRYL SUTURE (VIOLET) DIAMETER 3-0, LENGTH 70 CM.. FOR THIS REASON, A CHARACTERIZATION WAS PERFORMED TO SAMPLE: AND THE RESULT SHOWED THAT THE NEEDLE SUTURE COMBINATION WAS RELATED TO A DRILL NEEDLE WITH WIRE DIAMETER 40 (MILS), POINT TYPE: TAPER POINT WITH VICRYL VIOLET SUTURE IN DIAMETER 2-0¿ AND LENGTH 68 CM RESULTING A PRODUCT MIX. ACCORDING TO THE SAMPLE CONDITION TO THE ASSIGNABLE CAUSE TO ASSEMBLY - INCORRECT COMPONENT IS MIXED PRODUCT. IN ADDITION, THE SAMPLE WILL BE FORWARDED FOR FURTHER EVALUATION. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320911&pc=GAM | 2210968-2018-71331 |  | Malfunction | ETHICON INC. | 2018-03-07 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED BY THE VET THAT AN ANIMAL UNDERWENT AN OOPHORECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THREE OR FOUR DAYS POST OP THE VET NOTED A RUPTURE OF THE ABDOMINAL WALL SUTURE. THE KNOTS WERE INTACT. THE ANIMAL HAD A NEW SURGERY FOR EVENTRATION. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). DATE SENT TO FDA: (B)(4) 2018 A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7321011&pc=GAM | 2210968-2018-71334 |  | Malfunction | ETHICON INC. | 2018-03-07 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED BY THE VET THAT AN ANIMAL UNDERWENT AN OOPHORECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THREE OR FOUR DAYS POST OP THE VET NOTED A RUPTURE OF THE ABDOMINAL WALL SUTURE. THE KNOTS WERE INTACT. THE ANIMAL HAD A NEW SURGERY FOR EVENTRATION. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300379&pc=GAM | 2210968-2018-71070 | 2018-02-01 | Malfunction | ETHICON INC. | 2018-02-27 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE NEEDLE BODY IS NOT AS EXPECTED. THE LABEL DEMONSTRATES THAT NEEDLE BODY IS ROUND. THE PHYSICAL NEEDLE BODY IS DIFFERENT. Manufacturer Narrative: PC-(B)(4). DATE SENT TO THE FDA: (B)(4) 2018. A WINDING FORMER WITH A NEEDLE-SUTURE COMBINATION AND ONE EMPTY OPENED FOIL OF PRODUCT CODE WAS RETURNED. THE SWAGE AND ATTACHMENT AREA WERE AS EXPECTED, NO DEFECTS OR DAMAGED WERE FOUND ON THE TIP AND BODY NEEDLE. THE NEEDLE MET THE FINISH GOOD REQUIREMENTS. ALSO, THE FOIL AND THE PRINT INFORMATION WAS REVIEWED TO BE CORRECT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7269761&pc=GAM | 2210968-2018-70856 | 2018-01-17 | Malfunction | ETHICON INC. | 2018-02-14 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. POST- OPERATIVELY APPROXIMATELY ONE WEEK, THE SUTURE BROKE IN THE MIDDLE. THE VETERINARIAN NOTED HERNIATION OF THE WOUND. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). DATE SENT TO THE FDA: (B)(4) 2018. Manufacturer Narrative: (B)(4).A USED NEEDLE-SUTURE IN SEVERAL PIECES WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE (NEEDLE/SUTURE COMBINATION IN SEVERAL PIECES), THE SWAGE AND ATTACHMENT AREA WERE AS EXPECTED AND MARKS WERE OBSERVED ON THE BODY NEEDLE, AND THE END OF THE SUTURE WAS OBSERVED CUT APPEARS TO BE BY USE OF A SURGICAL INSTRUMENT. THE OTHER SECTIONS OF THE SUTURE WERE EXAMINED, BODY FLUIDS WERE NOTED ON THE STRANDS, AND THE ENDS OF THE SUTURES PIECES WERE CUT. IN HANDLING THIS OR ANY OTHER SUTURE MATERIAL, CARE SHOULD BE TAKEN TO AVOID DAMAGE FROM HANDLING. AVOID CRUSHING OR CRIMPING DAMAGE DUE TO APPLICATION OF SURGICAL INSTRUMENTS SUCH AS FORCEPS OR NEEDLE HOLDERS. IN ADDITION, THE TWO SUTURE PIECES WERE TESTED BY KNOT TENSILE STRENGTH AND THE RESULT WAS ABOVE THE MINIMUM INDIVIDUAL STRENGTH REQUIREMENT FOR SIZE ¿2-0¿ VICRYL SUTURE: ACCORDING TO THE SAMPLE CONDITION IT WAS SUGGEST AN IMPROPER HANDLING OF THE SAMPLE. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE NAME OF THE INITIAL PROCEDURE? OVERIOHYSTERCTOMY WHAT WAS THE DATE OF THE INITIAL PROCEDURE? (B)(4) 2018 HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? SURGICAL KNOT AT THE END ON WHAT TISSUE WAS THE SUTURE USED? OVARIES, ABDOMINAL WALL (MUSCLE/FASCIA), SKIN. DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? NOTHING ABNORMAL WAS NOTICED DURING SURGERY. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE CAT (FELIS SILVESTRUS CATUS), FEMALE, (B)(6). WEIGHT IS UNKNOWN AND BMI IS N/A IN CATS. Manufacturer Narrative: PC-(B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE NAME OF THE INITIAL PROCEDURE? WHAT WAS THE DATE OF THE INITIAL PROCEDURE? HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? ON WHAT TISSUE WAS THE SUTURE USED? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? WHAT INTERVENTION WAS PERFORMED FOR THIS SUTURE EVENT? WAS AN ADDITIONAL PROCEDURE INDICATED? PLEASE DESCRIBE WHAT WAS DONE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7269981&pc=GAM | 2210968-2018-70855 | 2018-01-17 | Malfunction | ETHICON INC. | 2018-02-14 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. POST- OPERATIVELY APPROXIMATELY ONE WEEK, THE SUTURE BROKE IN THE MIDDLE. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). DATE SENT TO THE FDA: (B)(4) 2018. Manufacturer Narrative: PC-(B)(4). REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE REPRESENTATIVE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND THE SWAGE AND ATTACHMENT AREA OF THE NEEDLE WERE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE OBSERVED. THE NEEDLE/SUTURE COMBINATION WAS TESTED BY KNOT TENSILE STRENGTH AND THE RESULTS IS ABOVE THE MINIMUM INDIVIDUAL STRENGTH REQUIREMENT. PER THE CONDITION OF THE REPRESENTATIVE SAMPLES, NO PERFORMANCE - BREAKAGE SUTURE ¿ POST OP WAS FOUND ON THE STRAND AND THE TESTED SUTURE MET THE FINISHED GOODS REQUIREMENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE NAME OF THE INITIAL PROCEDURE? OVERIOHYSTERCTOMY. WHAT WAS THE DATE OF THE INITIAL PROCEDURE? (B)(4) 2018. HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? SURGICAL KNOT AT THE END ON WHAT TISSUE WAS THE SUTURE USED? OVARIES, ABDOMINAL WALL (MUSCLE/FASCIA), SKIN. DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? NOTHING ABNORMAL WAS NOTICED DURING SURGERY. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE CAT (FELIS SILVESTRUS CATUS), FEMALE, 6 MONTHS OLD. WEIGHT IS UNKNOWN AND BMI IS N/A IN CATS. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE NAME OF THE INITIAL PROCEDURE? WHAT WAS THE DATE OF THE INITIAL PROCEDURE? HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? ON WHAT TISSUE WAS THE SUTURE USED? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? WHAT INTERVENTION WAS PERFORMED FOR THIS SUTURE EVENT? WAS AN ADDITIONAL PROCEDURE INDICATED? PLEASE DESCRIBE WHAT WAS DONE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7266694&pc=GAM | 2210968-2018-70824 |  | Malfunction | ETHICON INC. | 2018-02-13 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED SUBCUTANEOUSLY. AFTER A WHILE, THE SUTURE MOVED THROUGH THE SKIN AND DID NOT REABSORB. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE. THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? WAS MEDICAL OR SURGICAL INTERVENTION NEEDED TO ADDRESS THE EVENT? PRODUCT CODE AND LOT # IF APPLICABLE, WILL PRODUCT BE RETURNED, RETURN DATE, TRACKING INFORMATION WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7248274&pc=GAM | 2210968-2018-70726 | 2018-01-19 | Malfunction | ETHICON INC. | 2018-02-07 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A REDON FIXATION PROCEDURE AN UNKNOWN DATE AND SUTURE WAS USED TO PLACE A DRAIN. POST OPERATIVELY, THE SUTURE BROKE AT MULTIPLE PLACES. THE DRAIN FELT OUT OF THE WOUND DUE TO THE FACT THAT THE SUTURE BROKE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: PRODUCT COMPLAINT # PC-(B)(4). AN INVALID LOT WAS INITIALLY REPORTED. A VALID LOT WAS NOT PROVIDED, THEREFORE, THE MANUFACTURING RECORDS COULD NOT BE REVIEWED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: LOT # ? LL88DWHR0. PROCEDURE NAME ? REDON FIXATION. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CONFIRM, WHEN WAS THE SUTURE NOTICED TO BE BROKEN? INTRA-OP- DURING USE ON PATIENT, WHILE SUTURING ? OR POST-OP- AFTER SUTURING, DURING RECOVERY PERIOD? WAS INITIAL PROCEDURE COMPLETED SUCCESSFULLY? AND HOW? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? HOW WAS SUTURE INITIALLY PLACED (INTERRUPTED OR CONTINUOUS)? ANY PATIENT CONSEQUENCES OR ADVERSE EVENT OUTCOME? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT ? ANY MEDICAL/SURGICAL INTERVENTION WAS PERFORMED FOR THIS SUTURE EVENT? WAS AN ADDITIONAL PROCEDURE INDICATED? PLEASE DESCRIBE WHAT WAS DONE? ANY CHANGE IN THE PATIENT¿S POST-OPERATIVE COURSE AS A RESULT? PATIENT CURRENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7218334&pc=GAM | 2210968-2018-70446 | 2017-12-17 | Malfunction | ETHICON INC. | 2018-01-25 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Misassembled | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO ANY PROCEDURE, IT WAS NOTICED THAT A SUTURE PRODUCT BOX CONTAINED WRONG SUTURE COMPONENT IN IT. THE SURGEON OPENED BOX AND FOUND SUTURE OF PRODUCT CODE V904, WITH LOT NUMBER GG8HQLR0. THERE WERE NO PATIENT CONSEQUENCES. ADDITIONAL INFORMATION IS NOT AVAILABLE. Manufacturer Narrative: (B)(4). THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: WHAT IS THE LOT NUMBER OF THE OPENED BOX OF VCP460H? ADDITIONAL INFORMATION: THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT, FOR THE V904H, LOT GG8HQLR0. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7219770&pc=GAM | 2210968-2018-70448 | 2017-12-01 | Malfunction | ETHICON INC. | 2018-01-25 | GAM | VICRYL VIO 75CM M3 USP2/0 SGLE ARMED FSL | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A VETERINARY PROCEDURE (OOPHORECTOMY) IN (B)(6) 2017, SUTURE WAS USED. POST OPERATIVELY, A "BALL" APPEARED AT THE SURGICAL WOUND. THE VET NOTED AN EVENTRATION WITH RUPTURE OF THE ABDOMINAL WALL SUTURE ON ITS LENGTH, THE KNOTS WERE INTACT. ANOTHER SURGERY FOR EVENTRATION WAS PERFORMED WITH THREAD OF ANOTHER LOT NUMBER. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7217126&pc=GAM | 2210968-2018-70441 |  | Malfunction | ETHICON INC. | 2018-01-24 | GAM | VICRYL VIO 75CM M3 USP2/0 SGLE ARMED FSL | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT OOPHORECTOMY PROCEDURE ON UNKNOWN DATE IN (B)(6) 2017 AND SUTURE WAS USED. POST OPERATIVELY, THE SUTURE BROKE. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7198891&pc=GAM | 2210968-2018-70316 | 2017-12-21 | Malfunction | ETHICON INC. | 2018-01-17 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 UND | Component Incompatible | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC GASTRIC BYPASS ON (B)(6) 2017 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN THE DOCTOR WAS TYING THE KNOTS WITH THE SUTURE, THE DOCTOR FELT THE SUTURE WAS LIGHTER THAN EXPECTED. THE SUTURE FELT LIKE A DIFFERENT SIZE THAN WHAT WAS STATED. THE PROCEDURE WAS COMPLETED USING THE SAME SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 02/07/2018 Manufacturer Narrative: PRODUCT COMPLAINT # PC-(B)(4). INVESTIGATION SUMMARYTWO UNOPENED SAMPLES OF PRODUCT CODE (B)(4), LOT # LE7065 WERE RETURNED FOR ANALYSIS. THE SAMPLE WITH THE ISSUE REPORTED WAS NOT RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLE WAS OPENED AND THE SWAGE AND ATTACHMENT AREA WERE AS EXPECTED. THE SUTURE WAS DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WERE OBSERVED. THE SAMPLE WAS MEASURED BY DIAMETER AND MEET THE REQUIREMENTS. THE SAMPLE WITH THE ISSUE REPORTED WAS NOT RECEIVED FOR EVALUATION. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7189352&pc=GAM | 2210968-2018-70276 | 2017-12-13 | Malfunction | ETHICON INC. | 2018-01-15 | GAM | VICRYL UND 45CM M1.5 4/0 FS-2 | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. IT WAS NOTED THAT THE "WRONG ITEM IN PACK". THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PC-(B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY THE EVENT "WRONG ITEM IN PACK": WHAT PRODUCT CODE AND LOT NUMBER WAS IN PLACE OF PRODUCT W9951 LOT LH8DDMPO? - W9106 SUTURES WE¿RE INSIDE THE BOX INSTEAD OF W9951. AND CORRECT LOT NUMBER IS LH8DDMP0. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 02/07/2018 Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY THE EVENT "WRONG ITEM IN PACK": WHAT PRODUCT CODE AND LOT NUMBER WAS IN PLACE OF PRODUCT (B)(4), LOT LH8DDMPO?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7178338&pc=GAM | 2210968-2018-70171 | 2017-12-12 | Malfunction | ETHICON INC. | 2018-01-10 | GAM | VICRYL SUTURE 27"(70CM) 2-0 VIO | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE ON (B)(6) 2017, SUTURE WAS USED. THE SUTURE PRODUCT WAS WITHOUT THE BAR CODE ON THE PACKAGING. NO PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. LOT#? WAS THE BARCODE MISSING ON THE OUTER SECONDARY PACKAGE/BOX? WAS THE LABELLING DAMAGED IN ANY WAY? IS THERE AN INDICATION OF HOW THE PRODUCT WAS DISTRIBUTED? IS THERE ANY INDICATION OF THE SOURCE? BASED ON THE PACKAGING, IS THERE ANY INDICATION OF WHICH MARKET THE ORIGINAL GENUINE PRODUCT MAY HAVE COME FROM? ARE THERE ANY SECURITY FEATURES ON THE PRODUCT OR THE PACKAGING? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7178947&pc=GAM | 2210968-2018-70178 | 2017-11-30 | Malfunction | ETHICON INC. | 2018-01-10 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A VETERINARY PROCEDURE (SPAYING & NEUTERING), SUTURE WAS USED. IT WAS REPORTED THAT SUTURE BROKE IN FASCIA. NO PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7171595&pc=GAM | 2210968-2018-70092 | 2017-12-11 | Malfunction | ETHICON INC. | 2018-01-08 | GAM | VICRYL VIO 70CM M3 USP2/0 SGLE ARMED CT-1 PLUS | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROSTATECTOMY ON (B)(6) 2017 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN THE BOX WAS OPENED, ONE OF THE SUTURES WAS OF A DIFFERENT TYPE. A CORRECT SUTURE FROM THE SAME BOX WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PC-(B)(4). AN OPENED BOX WITH THIRTY-SIX UNOPENED SAMPLES OF PRODUCT CODE (B)(4), LOT LL8BXPP0 (VICRYL DIAMETER 2-0) AND TWO UNOPENED SAMPLES OF PRODUCT CODE (B)(4), LOT JJ8GDLM0 (VICRYL DIAMETER 2) WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLE, PRODUCT RECEIVED UNDER PRODUCT CODE (B)(4), LOT JJ8GDLM0 (VICRYL DIAMETER 2) IN A BOX OF THE PRODUCT CODE (B)(4), LOT LL8BXPP0 (VICRYL DIAMETER 2-0). A CHARACTERIZATION WAS PERFORMED TO PRODUCT IDENTIFIED WITH THE CODE (B)(4); THE RESULTS WERE RELATED TO A NEEDLE CT-1 PLUS WITH VICRYL VIOLET SUTURE IN DIAMETER 2¿ AND LENGTH 90CM¿ AND ACCORDING THE CHARACTERISTIC, THE SAMPLES BELONG TO PRODUCT CODE (B)(4). THE DESCRIPTIONS OF THE BOX FOR THE PRODUCT CODE (B)(4), LOT # LL8BXPP0 ARE RELATED TO NEEDLE CT-1 WITH VICRYL VIOLET SUTURE IN DIAMETER 2-0¿ AND LENGTH 70CM¿. THE PRODUCT CODE (B)(4), LOT JJ8GDLM0 WITH EXPIRY DATE 2021, 01 AND THE PRODUCT CODE (B)(4), LOT # LL8BXPP0 (BOX) WITH EXPIRY DATE 2023, 03 THESE PRODUCT CODES WERE MANUFACTURED WITH 2 YEARS OF SEPARATION BETWEEN EACH BATCH. LOT LL8BXPP0: MANUFACTURER DATE: (B)(4) 2017; EXPIRATION DATE: (B)(4) 2023. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THESE LOTS (LL8BXPP0 AND JJ8GDLM0). Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7171596&pc=GAM | 2210968-2018-70093 | 2017-12-11 | Malfunction | ETHICON INC. | 2018-01-08 | GAM | VICRYL VIO 70CM M3 USP2/0 SGLE ARMED CT-1 PLUS | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROSTATECTOMY ON (B)(6) 2017 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN THE BOX WAS OPENED, ONE OF THE SUTURES WAS OF A DIFFERENT TYPE. A CORRECT SUTURE FROM THE SAME BOX WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PC-(B)(4). AN OPENED BOX WITH THIRTY-SIX UNOPENED SAMPLES OF PRODUCT CODE V339, LOT LL8BXPP0 (VICRYL DIAMETER 2-0) AND TWO UNOPENED SAMPLES OF PRODUCT CODE V348, LOT JJ8GDLM0 (VICRYL DIAMETER 2) WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLE, PRODUCT RECEIVED UNDER PRODUCT CODE (B)(4), LOT JJ8GDLM0 (VICRYL DIAMETER 2) IN A BOX OF THE PRODUCT CODE (B)(4), LOT LL8BXPP0 (VICRYL DIAMETER 2-0). A CHARACTERIZATION WAS PERFORMED TO PRODUCT IDENTIFIED WITH THE CODE (B)(4); THE RESULTS WERE RELATED TO A NEEDLE CT-1 PLUS WITH VICRYL VIOLET SUTURE IN DIAMETER 2¿ AND LENGTH 90CM¿ AND ACCORDING THE CHARACTERISTIC, THE SAMPLES BELONG TO PRODUCT CODE (B)(4). THE DESCRIPTIONS OF THE BOX FOR THE PRODUCT CODE (B)(4), LOT # LL8BXPP0 ARE RELATED TO NEEDLE CT-1 WITH VICRYL VIOLET SUTURE IN DIAMETER 2-0¿ AND LENGTH 70CM¿. THE PRODUCT CODE (B)(4) ,LOT JJ8GDLM0 WITH EXPIRY DATE 2021, 01 AND THE PRODUCT CODE (B)(4), LOT # LL8BXPP0 (BOX) WITH EXPIRY DATE 2023, 03 THESE PRODUCT CODES WERE MANUFACTURED WITH 2 YEARS OF SEPARATION BETWEEN EACH BATCH. LOT LL8BXPP0: MANUFACTURER DATE: (B)(4) 2017; EXPIRATION DATE: (B)(4) 2023. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THESE LOTS (LL8BXPP0 AND JJ8GDLM0). Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7171608&pc=GAM | 2210968-2018-70094 | 2017-12-11 | Malfunction | ETHICON INC. | 2018-01-08 | GAM | VICRYL VIO 70CM M3 USP2/0 SGLE ARMED CT-1 PLUS | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROSTATECTOMY ON (B)(6) 2017 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN THE BOX WAS OPENED, ONE OF THE SUTURES WAS OF A DIFFERENT TYPE. A CORRECT SUTURE FROM THE SAME BOX WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: AN OPENED BOX WITH THIRTY-SIX UNOPENED SAMPLES OF PRODUCT CODE V339, LOT LL8BXPP0 (VICRYL DIAMETER 2-0) AND TWO UNOPENED SAMPLES OF PRODUCT CODE V348, LOT JJ8GDLM0 (VICRYL DIAMETER 2) WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLE, PRODUCT RECEIVED UNDER PRODUCT CODE V348, LOT JJ8GDLM0 (VICRYL DIAMETER 2) IN A BOX OF THE PRODUCT CODE V339, LOT LL8BXPP0 (VICRYL DIAMETER 2-0). A CHARACTERIZATION WAS PERFORMED TO PRODUCT IDENTIFIED WITH THE CODE V348; THE RESULTS WERE RELATED TO A NEEDLE CT-1 PLUS WITH VICRYL VIOLET SUTURE IN DIAMETER 2¿ AND LENGTH 90CM¿ AND ACCORDING THE CHARACTERISTIC, THE SAMPLES BELONG TO PRODUCT CODE V348. THE DESCRIPTIONS OF THE BOX FOR THE PRODUCT CODE V339, LOT # LL8BXPP0 ARE RELATED TO NEEDLE CT-1 WITH VICRYL VIOLET SUTURE IN DIAMETER 2-0¿ AND LENGTH 70CM¿. THE PRODUCT CODE V348, LOT JJ8GDLM0 WITH EXPIRY DATE 2021, 01 AND THE PRODUCT CODE V339, LOT # LL8BXPP0 (BOX) WITH EXPIRY DATE 2023, 03 THESE PRODUCT CODES WERE MANUFACTURED WITH 2 YEARS OF SEPARATION BETWEEN EACH BATCH. LOT LL8BXPP0: MANUFACTURER DATE: 10/01/2017; EXPIRATION DATE: 01/31/2023 THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THESE LOTS (LL8BXPP0 AND JJ8GDLM0). Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7171648&pc=GAM | 2210968-2018-70098 | 2017-12-12 | Malfunction | ETHICON INC. | 2018-01-08 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. DURING THE PROCEDURE IN THE EMERGENCY AREA, WHEN THE DOCTOR OPENED THE SUTURE PACKAGING, A DIFFERENT TYPE SUTURE WAS IN THE PACKAGE. THE SUTURE WAS NOT USED ON THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE ATTACH PHOTOS/PICTURES TO THIS FILE FOR THE QUALITY ISSUE WITH VICRYL SUTURE 3-0 AS REPORTED. PRODUCT CODE FOR THE VICRYL SUTURE 3-0? LOT # ? WAS THE SUTURE REMOVED FROM A BRAND-NEW BOX OR WAS THIS ISSUE DISCOVERED AFTER THE BOX HAD BEEN PREVIOUSLY OPENED AND SITTING IN THE STOREROOM OR OTHER LOCATION PRIOR TO USE? WAS THE PACKAGE NOT PROPERLY SEALED? WAS THE INTEGRITY OF THE SEAL COMPROMISED IN ANY WAY? WAS THE BOX NOTICED TO HAVE BEEN TAMPERED WITH? ANY SAMPLE AVAILABLE TO BE RETURNED FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7125466&pc=GAM | 2210968-2017-71834 |  | Malfunction | ETHICON INC. | 2017-12-18 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Component Incompatible | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN OB-GYN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. PRIOR TO USE ON THE PATIENT , IT WAS FOUND THAT THE SHAPE OF THE NEEDLE WAS NOT FOR THAT SPECIFIC PRODUCT CODE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH LH9CSME0 MFG DATE: 07/01/2017, EXP DATE: 12/31/2022, BATCH KM9GMBEO, MFG DATE: 11/01/2016, EXP DATE: 04/30/2022. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH LH9CSME0, MFG DATE: 07/01/2017, EXP DATE: 12/31/2022. BATCH KM9GMBEO, MFG DATE: UNK, EXP DATE: UNK. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: LOT # ? THE LOT NUMBER ¿ LH9CSME0 KM9GMBEO ¿ UNKNOWN IF THESE ARE THE CORRECT LOT NUMBERS. PLEASE CLARIFY WHAT DO YOU MEAN BY "" SHAPE OF THE NEEDLE WAS NOT FOR VR945""? DID THE PACKAGE FOR VR945 CONTAIN INCORRECT NEEDLE INSIDE AND NOTICED BEFORE USE ON PATIENT? YES. OR WAS THERE ISSUE NOTICED WITH NEEDLE APPEARANCE ONLY AND WAS THE CORRECT NEEDLE FOR VR945? NO. DID THE NEEDLE HAVE ANY SURFACE RELATED DEFECT ? NO INFORMATION. IF YES, WAS IT NOTICED PRIOR TO USE ON PATIENT ? OR DURING USE ON PATIENT? ANY SAMPLE AVAILABLE TO RETURN FOR EVAL? NO SAMPLE WILL BE RETURNED. NO FURTHER INFORMATION WILL BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7096275&pc=GAM | 2210968-2017-71660 | 2017-11-01 | Malfunction | ETHICON INC. | 2017-12-07 | GAM | VICRYL SUTURE 27(70CM) 3-0 VIO | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SURGICAL PROCEDURE IN (B)(6) 2017 AND SUTURE WAS USED. IT WAS NOTICED THAT THE NEEDLE SHAPE LOOKED LIKE A J SHAPE TAPER POINT RATHER THAN ACTUAL SPECIFICATION OF A CIRCLE. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # PC-(B)(4). THERE WAS NO SAMPLE RECEIVED FOR ANALYSIS. ONLY A PICTURE OF THE SAMPLE WAS SENT. UPON VISUAL INSPECTION OF THE PICTURE, AN OPENED SAMPLE WITH BODY FLUIDS AT THE PAPER FOLDER AND THE NEEDLE/SUTURE OUTSIDE OF THE WINDER FORMER COULD BE OBSERVED, IN ADDITION, THE NEEDLE HAS AN INCORRECT CURVATURE AND NO AN INCORRECT COMPONENT. Manufacturer Narrative: PC-(B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT TISSUE WAS DAMAGE? ORGAN. WAS THERE ANY PATIENT CONSEQUENCE? NO, BUT TISSUE DAMAGE TO WIDE BECAUSE SWAGE IS TOO LARGE (TISSUE DAMAGE IS NOT FIT WITH THE NEEDLE SIZE) -WAS MEDICAL OR SURGICAL INTERVENTION REQUIRED TO RESOLVE THE TISSUE DAMAGE? STITCH BACK-WHAT IS THE LOT NUMBER? LAZ361. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 12/07/2017. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TISSUE WAS DAMAGED? WAS THERE ANY PATIENT CONSEQUENCE? WAS MEDICAL OR SURGICAL INTERVENTION REQUIRED TO RESOLVE THE TISSUE DAMAGE? WHAT IS THE LOT NUMBER?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7078496&pc=GAM | 7078496 | 2017-11-21 | Malfunction | ETHICON, LLC | 2017-12-04 | GAM | VICRYL RAPIDE | Detachment Of Device Component | Device Embedded In Tissue or Plaque |  |  | 1 | Event Description: AFTER A VAGINAL DELIVERY THE PHYSICIAN WAS REPAIRING A LACERATION WITH A 4-0 VICRYL RADPIDE, WHILE SETTING THE SUTURE IT WAS NOTED THAT THE TIP OF THE NEEDLE WAS RETAINED IN THE PATIENT'S LABIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7079148&pc=GAM | 2210968-2017-71560 | 2017-11-10 | Malfunction | ETHICON INC. | 2017-12-04 | GAM | VICRYL VIO 70CM M0.7 USP6/0 SGLE ARM V-18 | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HYPOSPADIAS PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN THE PACKAGE WAS OPENED, IT WAS NOTED THAT THE NEEDLE WAS THINNER THAN NORMAL IN THE ATTACHMENT REGION. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ADDITIONAL EVALUATION SUMMARY: A USED NEEDLE-SUTURE PIECE WAS RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE NEEDLE-SUTURE PIECE, THE SWAGE AND ATTACHMENT AREA OF NEEDLE WERE AS EXPECTED. THE SUTURE PIECE PRESENT BODY FLUIDS AND IT WAS OBSERVED WITH OPEN BRAID. NO OTHER DAMAGED OR DEFECTS WERE OBSERVED ON THE NEEDLE OR SUTURE PIECE. ADDITIONALLY, A FUNCTIONAL TEST WAS PERFORMED, THE PULL AND TENSILE FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. IN ADDITION, THE TOP FOIL WAS EXAMINED AND NO DAMAGE OR DEFECTS WERE FOUND. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF RECEIVED SAMPLE, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT . AN OPENED SAMPLE WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE DISPENSED NEEDLE-SUTURE COMBINATION , THE SWAGE AND ATTACHMENT AREA WERE AS EXPECTED. THE SUTURE WAS EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WAS OBSERVED. ADDITIONALLY, A FUNCTIONAL TEST WAS PERFORMED , AND THE PULL AND TENSILE FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. IN ADDITION, THE LABELED WINDING FORMER WAS EXAMINED AND NO DAMAGE OR DEFECTS WERE NOTED. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE RECEIVED SAMPLE, NO DEFECTS WERE FOUND ON THE SAMPLE. AN UNOPENED REPRESENTATIVE SAMPLE WAS RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLE, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLE WAS OPENED AND THE SWAGE AND ATTACHMENT AREA OF THE NEEDLES WERE AS EXPECTED. THE SUTURE WAS DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WERE OBSERVED. ADDITIONALLY, A FUNCTIONAL TEST WAS PERFORMED, THE PULL AND TENSILE FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE RECEIVED SAMPLE, NO DEFECTS WERE FOUND . Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7052305&pc=GAM | 2210968-2017-71395 | 2017-10-27 | Malfunction | ETHICON INC. | 2017-11-22 | GAM | VICRYL SUTURE 27"(70CM) 2-0 VIO | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY THE DISTRIBUTOR THAT WHEN THE CUSTOMER RECEIVED THEIR ORDER ON (B)(6) 2017, THE OUTER BOX LOOKED FINE BUT WHEN OPENED THE SUTURE WAS OUTSIDE OF ITS PACKAGING AND NO LONGER STERILE. NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW WAS THE INTEGRITY OF THE SEAL COMPROMISED? DID THE PACKAGING LOOK DAMAGED? WHERE THERE ANY QUALITY ISSUES NOTED PRIOR TO THE PACKAGING BEING HANDLED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7046267&pc=GAM | 2210968-2017-71342 | 2017-10-25 | Malfunction | ETHICON INC. | 2017-11-20 | GAM | VICRYL SUTURE 18"(45CM) 1 VIO | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2017 AND SUTURE WAS USED. IT WAS FOUND THAT THERE WAS A SUTURE WITH THE THICKNESS OF 2-0 IN THE PACKAGE BEFORE USE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IS THIS COMPLAINT RELATED TO SUTURE DIAMETER OR IS THE ISSUE RELATED TO AN INCORRECT COMPONENT IN THE PACKAGE? THIS COMPLAINT IS RELATED TO AN INCORRECT COMPONENT IN THE PACKAGE. DO YOU HAVE A PHOTO? NO PHOTO WAS PROVIDED FROM THE HP. NO FURTHER INFORMATION WILL BE PROVIDED. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IS THIS COMPLAINT RELATED TO SUTURE DIAMETER OR IS THE ISSUE RELATED TO AN INCORRECT COMPONENT IN THE PACKAGE? DO YOU HAVE A PHOTO?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7028406&pc=GAM | 2210968-2017-71233 | 2017-10-13 | Malfunction | ETHICON INC. | 2017-11-14 | GAM | VICRYL POLYGLACTIN SUTURE | Detachment Of Device Component; Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ELECTIVE CAESAREAN SECTION PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. WHILE SUTURING OF THE UTERUS, THE OR STAFF NOTED THAT THE VERY TIP OF THE SUTURE NEEDLE WAS MISSING (AROUND 1-2MM IN LENGTH). THE NEEDLE TIP AS DESCRIBED WAS FOUND ON A SWAB. IN ADDITION, THE OUTER PACKAGING AND THE LOT NUMBER INSIDE DIFFER. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. 1. CLARIFY THE PRODUCT CODE AND LOT NUMBER? 2. CLARIFY THAT THE NEEDLE WAS NOT IN THE PATIENT? ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. TWO POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH LH8DZBP0, MFG DATE: 07/2017, EXP DATE: 12/2022. BATCH LH8DZBP1, MFG DATE: 07/2017, EXP DATE: 12/2022. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7006010&pc=GAM | 2210968-2017-70819 | 2017-08-13 | Malfunction | ETHICON INC. | 2017-11-07 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CESAREAN SECTION PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. THE PATIENT WAS DISCHARGED ON (B)(6) 2017. IT WAS NOTED THAT ONE PIECE OF SUTURE CAME OUT FROM VAGINA ON (B)(6) 2017. COLOR DOPPLER ULTRASOUND WAS COMPLETED AND NO OTHER ANOMALIES WERE FOUND. NOW THE PATIENT'S CONDITION IS STABLE. THE SURGEON OPINES THAT IT MIGHT BE A SUTURE ABSORPTION ISSUE, THAT THE SUTURE ABSORBED TOO QUICKLY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). AN EMPTY OPENED FOIL AND TWO SUTURE PIECES WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE TWO SUTURE PIECES WAS OBSERVED BODY FLUIDS AND TISSUES ALONG OF THE SUTURE PIECES. THE SUTURE PIECES WERE NOTED TO HAVE SOME DEGRADATION AS THE SUTURE IS ABSORBABLE. NO CONCLUSION COULD BE REACH ON WHAT CAUSE THE REPORTED EVENT. Manufacturer Narrative: ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6979548&pc=GAM | 2210968-2017-70657 |  | Malfunction | ETHICON INC. | 2017-10-26 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE PATIENT CAME BACK TO THE HOSPITAL ONE MONTH POST OP FOLLOW UP TO REMOVE THE FIXATION. DURING THE VISIT THE DOCTOR NOTICED THAT THE SUTURE WHICH HAD BEEN USED TO CLOSE THE SUBCUTANEOUS TISSUE HAS ENCAPSULATED IN TISSUE AND LOOKS AS THOUGH NO ABSORPTION HAS STARTED. THE DOCTOR CLAIMS THAT THE SUTURE LOOKS EXACTLY THE SAME AS ON THE DAY OF IMPLANTATION. THERE IS TISSUE REACTION VISIBLE AROUND THE SUTURE, BUT IT IS NOT AN INFECTION AS BACTERIAL GROWTH HAS BEEN TAKEN AND THE TESTS HAVE COME BACK NEGATIVE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE ACTUAL PRODUCT CODE ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE PRODUCT CODES: V224H; VCP345H; V346H; VCP316H. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: EXACT SUTURE CODES AND LOT NUMBERS ARE NOT KNOWN, THE SURGEON CLAIMS THAT HE HAS BEEN USING EITHER 2/0; 3/0 OR = SIZE VICRYL, THEN FROM WHAT'S AVAILABLE TO THE HOSPITAL WE HAVE BEEN ABLE TO NARROW IT DOWN TO THE 4 FOLLOWING CODES:V224H 2/0 SH; VCP345H 2/0 CT-1; V346H 0 CT-1; VCP316H 3/0 SH; EVENT DATE: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS. PROCEDURE NAME: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS. DID PATIENT EXPERIENCE ANY SYMPTOMS?: UNABLE TO DISCLOSE MEDICAL RECORDS, ACCORDING TO GENERAL DESCRIPTION GIVEN BY THE SURGEON THERE WAS VISIBLE TISSUE REACTION THAT WAS NOT INFECTION. WAS ANY MEDICAL/SURGICAL INTERVENTION PROVIDED BECAUSE OF ENCAPSULATION/POSSIBLE NON-ABSORPTION OF THE SUTURE?: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS. PATIENT DEMOGRAPHICS: INITIALS / ID; AGE OR DATE OF BIRTH; BMI; GENDER: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS. OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS. ARE ANY REPRESENTATIVE SAMPLE AVAILABLE FOR EVALUATION?: NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6982572&pc=GAM | 2210968-2017-70669 |  | Malfunction | ETHICON INC. | 2017-10-26 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE PATIENT CAME BACK TO THE HOSPITAL ONE MONTH POST OP FOLLOW UP TO REMOVE THE FIXATION. DURING THE VISIT THE DOCTOR NOTICED THAT THE SUTURE WHICH HAD BEEN USED TO CLOSE THE SUBCUTANEOUS TISSUE HAS ENCAPSULATED IN TISSUE AND LOOKS AS THOUGH NO ABSORPTION HAS STARTED. THE DOCTOR CLAIMS THAT THE SUTURE LOOKS EXACTLY THE SAME AS ON THE DAY OF IMPLANTATION. THERE IS TISSUE REACTION VISIBLE AROUND THE SUTURE, BUT IT IS NOT AN INFECTION AS BACTERIAL GROWTH HAS BEEN TAKEN AND THE TESTS HAVE COME BACK NEGATIVE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE ACTUAL PRODUCT CODE ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE PRODUCT CODES: V224H; VCP345H; V346H; VCP316H. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: EXACT SUTURE CODES AND LOT NUMBERS ARE NOT KNOWN, THE SURGEON CLAIMS THAT HE HAS BEEN USING EITHER 2/0; 3/0 OR = SIZE VICRYL, THEN FROM WHAT'S AVAILABLE TO THE HOSPITAL WE HAVE BEEN ABLE TO NARROW IT DOWN TO THE 4 FOLLOWING CODES:V224H 2/0 SH; VCP345H 2/0 CT-1; V346H 0 CT-1; VCP316H 3/0 SH; EVENT DATE: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS PROCEDURE NAME: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS DID PATIENT EXPERIENCE ANY SYMPTOMS?: UNABLE TO DISCLOSE MEDICAL RECORDS, ACCORDING TO GENERAL DESCRIPTION GIVEN BY THE SURGEON THERE WAS VISIBLE TISSUE REACTION THAT WAS NOT INFECTION. WAS ANY MEDICAL/SURGICAL INTERVENTION PROVIDED BECAUSE OF ENCAPSULATION/POSSIBLE NON-ABSORPTION OF THE SUTURE?: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS PATIENT DEMOGRAPHICS: INITIALS / ID; AGE OR DATE OF BIRTH; BMI; GENDER: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS. OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS: UNABLE TO DISCLOSE DUE TO PRIVACY REGULATIONS. ARE ANY REPRESENTATIVE SAMPLE AVAILABLE FOR EVALUATION?: NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6937118&pc=FTL | 6937118 | 2017-08-21 | Malfunction | ETHICON, INC. | 2017-10-10 | FTL | VICRYL | Break; Detachment Of Device Component | No Code Available |  |  | 1 | Event Description: SURGICAL RESIDENT WAS INSERTING A "STAY STITCH" IN UMBILICUS WHEN THE SUTURE NEEDLE BROKE INTO TWO PIECES LEAVING A HALF OF THE NEEDLE IN THE PATIENT. THE SURGEON EXPLORED THE SURGICAL SITE AND ABDOMEN LOOKING FOR THE NEEDLE REMNANT AND WAS UNABLE TO LOCATE. X-RAY WAS CALLED TO THE OR AND THE NEEDLE REMNANT WAS LOCATED AND SUCCESSFULLY REMOVED FROM THE PATIENT. PACKAGING FROM BROKEN SUTURE SAVED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6862553&pc=GAM | 2210968-2017-70265 |  | Malfunction | ETHICON INC. | 2017-09-13 | GAM | VICRYL RAPIDE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SPONTANEOUS DELIVERY ON AN UNKNOWN DATE AND SUTURE WAS USED. APPROXIMATELY, SIX MONTHS AFTER CHILDBIRTH, A SMALL FOREIGN SUBSTANCE APPEARED FROM THE VAGINA. THE SUBSTANCE WAS SUSPECTED TO BE A PIECE OF THE SUTURE. THERE WAS NO MEDICAL INTERVENTION PERFORMED AND NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAD BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT OB-GYN PROCEDURE DID THE PATIENT HAVE? =>NATURAL DELIVERY. WHAT SYMPTOMS DID THE PATIENT RETURN 6 MONTHS POST DELIVERY? =>NO SYMPTOM RETURNED. A SUBSTANCE CAME OUT OF THE VAGINA. DOES THE DOCTOR OPINION THAT THE PATIENT SUFFERED CONSEQUENCES DUE TO POSSIBLE NON OR SLOW ABSORPTION OF THE SUTURE? =>NO INFORMATION WAS PROVIDED. WHAT MEDICAL / SURGICAL INTERVENTION WAS PERFORMED? => NO INTERVENTION WAS PERFORMED. NO FURTHER INFORMATION WILL BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6846708&pc=GAM | 2210968-2017-70204 |  | Malfunction | ETHICON INC. | 2017-09-06 | GAM | VICRYL SUTURE 12"(30CM) 8-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VETERINARIAN THAT AN ANIMAL UNDERWENT UNILATERAL CATARACT SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE INCISION WAS CLOSED WITH SIMPLE INTERRUPTED SUTURES. THE VET SURGEON USES SUTURE FORCEPS FOR TYING THE KNOTS AND ALWAYS HANDLES THE SUTURE CAREFULLY. POST OPERATIVELY THE SUTURE WAS BREAKING DOWN. THE PATIENT DID NOT NEED ADDITIONAL SURGICAL TREATMENT. Manufacturer Narrative: TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. Manufacturer Narrative: DEVICE EVALUATION: THE ACTUAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. ONLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, THERE APPEARS TO BE A RUPTURE OF THE SUTURE. DUE TO THE SAMPLE WAS NOT RETURNED, NO CONCLUSION COULD BE REACHED AS HOW THIS COULD HAVE HAPPENED. Manufacturer Narrative: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6838549&pc=GAM | 2210968-2017-70180 | 2017-08-01 | Malfunction | ETHICON INC. | 2017-09-01 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY PROCEDURE ON (B)(6)2017 AND SUTURE WAS USED. AFTER CLOSING THE SKIN THERE WAS SUTURE BREAKDOWN OF SURGICAL WOUND WITHIN WEEK OF SURGERY. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE. THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS. FOUR IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODES NW2493/NW2494, ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: - WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? - WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? - HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? - HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? - WHAT DATE DID THE SUTURE BREAK? - HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. - WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6838605&pc=GAM | 2210968-2017-70181 | 2017-08-01 | Malfunction | ETHICON INC. | 2017-09-01 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY PROCEDURE ON (B)(6)2017 AND SUTURE WAS USED. AFTER CLOSING THE SKIN THERE WAS SUTURE BREAKDOWN OF SURGICAL WOUND WITHIN WEEK OF SURGERY. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS. FOUR IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODES NW2493/NW2494, ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: - WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? - WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? - HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? - HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? - WHAT DATE DID THE SUTURE BREAK? - HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. - WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT. WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6838803&pc=GAM | 2210968-2017-70184 |  | Malfunction | ETHICON INC. | 2017-09-01 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PLASTIC SURGERY PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. AFTER CLOSING THE SKIN THERE WAS SUTURE BREAKDOWN OF SURGICAL WOUND WITHIN WEEK OF SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED INDICATES THAT THIS EVENT DOES NOT MEET REPORTING MALFUNCTION CRITERIA. THIS EVENT THEREFORE IS NOT REPORTABLE. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS 4 IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODE NW2493/NW2494 AND LOTS T7003, T7004 AND T7001, IS FORTH LOT NUMBER MISSING? ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? WHAT DATE DID THE SUTURE BREAK? HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6838846&pc=GAM | 2210968-2017-70185 |  | Malfunction | ETHICON INC. | 2017-09-01 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PLASTIC SURGERY PROCEDURE ON (B)(6)2017 AND SUTURE WAS USED. AFTER CLOSING THE SKIN THERE WAS SUTURE BREAKDOWN OF SURGICAL WOUND WITHIN WEEK OF SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED INDICATES THAT THIS EVENT DOES NOT MEET REPORTING MALFUNCTION CRITERIA. THIS EVENT THEREFORE IS NOT REPORTABLE. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS. FOUR IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODE NW2493/NW2494 AND LOTS T7003, T7004 AND T7001, IS FORTH LOT NUMBER MISSING? ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: - WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? - WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? - HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? - HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? - WHAT DATE DID THE SUTURE BREAK? - HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. - WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6838918&pc=GAM | 2210968-2017-70186 |  | Malfunction | ETHICON INC. | 2017-09-01 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | Wound Dehiscence; No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PLASTIC SURGERY PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. AFTER CLOSING THE SKIN THERE WAS SUTURE BREAKDOWN OF SURGICAL WOUND WITHIN WEEK OF SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED THAT INDICATES THIS EVENT DOES NOT MEET REPORTING MALFUNCTION CRITERIA. THIS EVENT THEREFORE IS NOT REPORTABLE. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS: 4 IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODE NW2493/NW2494 AND LOTS T7003, T7004 AND T7001, IS FORTH LOT NUMBER MISSING? ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? WHAT DATE DID THE SUTURE BREAK? HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6839228&pc=GAM | 2210968-2017-70187 | 2017-08-01 | Malfunction | ETHICON INC. | 2017-09-01 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. AFTER CLOSING THE SKIN THERE WAS SUTURE BREAKDOWN OF SURGICAL WOUND WITHIN WEEK OF SURGERY. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS: 4 IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODES NW2493/NW2494, ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: - WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? - WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? - HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? - HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? - WHAT DATE DID THE SUTURE BREAK? - HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. - WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6837160&pc=GAM | 2210968-2017-70171 |  | Malfunction | ETHICON INC. | 2017-08-31 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. AFTER CLOSING THE SKIN THERE WAS SUTURE BREAKDOWN OF SURGICAL WOUND WITHIN WEEK OF SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE, THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS. 4 IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODES NW2493/NW2494, ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? WHAT DATE DID THE SUTURE BREAK? HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6837173&pc=GAM | 2210968-2017-70172 |  | Malfunction | ETHICON INC. | 2017-08-31 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. AFTER CLOSING THE SKIN THERE WAS SUTURE BREAKDOWN OF SURGICAL WOUND WITHIN WEEK OF SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE. THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS: 4 IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODES NW2493/NW2494, ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? WHAT DATE DID THE SUTURE BREAK? HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6837196&pc=GAM | 2210968-2017-70173 |  | Malfunction | ETHICON INC. | 2017-08-31 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. AFTER CLOSING THE SKIN THERE WAS SUTURE BREAKDOWN OF SURGICAL WOUND WITHIN WEEK OF SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE, THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS: 4 IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODES NW2493/NW2494, ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? WHAT DATE DID THE SUTURE BREAK? HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6831668&pc=GAM | 2210968-2017-70149 | 2017-08-01 | Malfunction | ETHICON INC. | 2017-08-30 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY PROCEDURE AND SUTURE WAS USED TO CLOSE THE SKIN. WITHIN ONE WEEK POST OPERATIVELY, BREAKDOWN OF SURGICAL WOUND WAS REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE, THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS 4 IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODES NW2493/NW2494, ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: - WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? - WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? - HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? - HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? - WHAT DATE DID THE SUTURE BREAK? - HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. - WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6831867&pc=GAM | 2210968-2017-70150 | 2017-08-01 | Malfunction | ETHICON INC. | 2017-08-30 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY PROCEDURE AND SUTURE WAS USED TO CLOSE THE SKIN. WITHIN ONE WEEK POST OPERATIVELY, BREAKDOWN OF SURGICAL WOUND WAS REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS 4 IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODES NW2493/NW2494, ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? WHAT DATE DID THE SUTURE BREAK? HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6831948&pc=GAM | 2210968-2017-70151 | 2017-08-01 | Malfunction | ETHICON INC. | 2017-08-30 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY PROCEDURE AND SUTURE WAS USED TO CLOSE THE SKIN. WITHIN ONE WEEK POST OPERATIVELY, BREAKDOWN OF SURGICAL WOUND WAS REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE DATE AND NAME OF INDEX SURGICAL PROCEDURE THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WERE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS: FOUR IMPACT PRODUCTS WERE CREATED FOR PRODUCT CODES (B)(4), ARE THESE POSSIBLE LOTS USED DURING THE PROCEDURE, OR DID SUTURES FROM ALL THE LOTS WERE USED IN THE PATIENT? HOW MANY SUTURES FROM EACH LOT BREAK AFTER THE PROCEDURE? FOR EACH SUTURE, PLEASE INFORM: - WHAT TISSUE TYPE AND LOCATION OF THE SUTURE PLACEMENT? - WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? - HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? - HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? - WHAT DATE DID THE SUTURE BREAK? - HOW WAS THE ISSUE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. - WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6784211&pc=GAM | 2210968-2017-33421 |  | Malfunction | ETHICON INC. | 2017-08-10 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Break | Wound Dehiscence | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A VAGINAL DELIVERY ON (B)(6) 2017 OR (B)(6) 2017 AND THE SUTURE WAS USED FOR PERINEORRHAPHY DURING THE DELIVERY. ON POST-OPERATIVE DAY 5, THE WOUND BECAME OPEN AND THE SUTURE BREAKAGE WAS NATURAL GIVEN THE TENSILE STRENGTH. THE DOCTOR OPINED THAT IT MIGHT DEPEND ON SUTURING TECHNIQUE. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE ADVISE HOW THE WOUND DEHISCENCE WAS ADDRESSED? NO INFORMATION. WAS THERE ANY POST OP SUTURE BREAKAGE NOTED? YES, BUT THE SURGEON SAID THAT IT OCCURRED ABOUT POD5 AND SUTURE BREAKAGE IS NATURAL GIVEN THE TENSILE STRENGTH. INITIAL PROCEDURE WAS PERFORMED BY YOUNG SURGEON, SO ADVISING DR. SAID IT MIGHT DEPEND ON SUTURING TECHNIQUE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6776917&pc=GAM | 2210968-2017-33397 | 2017-07-14 | Malfunction | ETHICON INC. | 2017-08-09 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PRIOR TO UNKNOWN PROCEDURE ON THE PATIENT ON UNKNOWN DATE , THE CUSTOMER WAS NOT ABLE TO READ THE BARCODE LABEL EXPIRATION DATE. AS THE DEVICE WAS NEVER USED ON THE PATIENT, THERE WERE NO ADVERSE CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: (B)(4). THE REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE EXPLAIN WERE UNITS AFFECTED FROM THE SAME BATCH? PLEASE EXPLAIN WHAT DOES (B)(4) WAS OPENED TO ADDRESS THE ISSUE MEAN? WHAT DO YOU MEAN BY ¿CUSTOMER NOT BEING ABLE TO READ THE BARCODE LABEL EXPIRATION DATE¿? WAS THE INFORMATION NOT CLEARLY PRINTED? WAS THE INFORMATION SMUDGED ON THE PACKAGING? DO YOU HAVE PICTURES OF THE PRODUCT SHOWING THIS ISSUE FOR EVALUATION PURPOSE? WAS THE LABEL OF THE PRODUCT DAMAGED DURING TRANSIT? WERE THESE PRODUCTS RECEIVED IN DAMAGED CONDITION? PLEASE EXPLAIN HOW THE BARCODE LABEL EXPIRATION DATE BECAME UNCLEAR/UNREADABLE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6773342&pc=GAT | 6773342 | 2017-07-07 | Malfunction | ETHICON, LLC | 2017-08-08 | GAT | VICRYL | Manufacturing, Packaging or Shipping Problem; Packaging Problem | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: SUTURE PACKAGE HAD EXTRA NEEDLE AND SUTURE NOTED BY TEAM AS PULLED OUT OF PACKAGE TO USE. SURGICAL TECH NOTICED ONE SUTURE ON TOP OF ANOTHER. NO ADVERSE EVENT. MANUFACTURER ERROR IN PACKAGING. RESULTED IN MISCOUNT AT END OF PROCEDURE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6773401&pc=GAM | 6773401 | 2017-07-07 | Malfunction | ETHICON, LLC | 2017-08-08 | GAM | VICRYL | Packaging Problem | No Consequences Or Impact To Patient |  |  | 1 | Event Description: SURGICAL TECH WENT TO LOAD SUTURE ONTO THE NEEDLE HOLDER AND NOTICED ONE SUTURE ON TOP OF THE OTHER. NO ADVERSE EVENT. MANUFACTURER ERROR IN PACKAGING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6748679&pc=GAM | 2210968-2017-33167 |  | Malfunction | ETHICON INC. | 2017-07-27 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | Device Embedded In Tissue or Plaque | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC APPENDECTOMY PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE APPROXIMATELY A QUARTER FROM THE SWAGE WHILE THE SURGEON WAS SUTURING IN FASCIA. THE BROKEN NEEDLE REMAINED IN FASCIA. THE SURGEON OPINED THAT THE RETENTION OF THE NEEDLE IS PREFERRED BECAUSE ATTEMPTING TO REMOVE THE NEEDLE WOULD DO MORE HARM THAN LEAVING THE NEEDLE AS IS. THE SURGEON OPINED THAT ATTEMPTING TO REMOVE COULD INCREASE THE RISK OF RECURRENT HERNIA. Manufacturer Narrative: THE ACTUAL DEVICE WAS RETURNED FOR EVALUATION. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF PREDOMINATELY MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. THIS WAS A DUCTILE FRACTURE. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED FROM MECHANICAL DEFORMATION, INDENTS AND SCRATCHES, DUE TO GRIPPING, BENDING AND OVERSTRESSING OF THE NEEDLE. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW THAT WOULD CAUSE PREMATURE FAILURE. Manufacturer Narrative: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBER IS REPORTED AS FOLLOWS: BATCH # HGM614, EXP. DATE 06/30/2019, MFR. DATE 06/05/2014 IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBER WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBER IS REPORTED AS FOLLOWS: BATCH # HGM614 . THE PHOTO OF THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE EVALUATION OF THE PHOTO IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: LOT # ? IT COULD BE LOT# HGM614, BUT THIS IS NOT CONFIRMED. THIS IS BASED ON THE REMAINING INVENTORY ON THEIR SHELF. WHAT WAS THE SIZE OF THE NEEDLE USED? 2-0 COATED VICRYL PLUS UR-5 NEEDLE (36 MM TAPER NEEDLE). WAS MORE THAN HALF THE NEEDLE RETAINED IN THE PATIENT? YES, MORE THAN HALF. WHAT TISSUE STRUCTURE IS THE NEEDLE LOCATED AT? FASCIA. WHERE WAS THE NEEDLE GRASPED PRIOR TO THE NEEDLE BREAKAGE (IE. SWAGE, CENTRAL PORTION OF NEEDLE)? UNKNOWN. WERE ANY MEASURES WERE TAKEN TO RETRIEVE THE BROKEN PIECE? IF YES, PLEASE SPECIFY WHAT WAS DONE? NO. HAS THERE BEEN ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED? NO. WAS THE PATIENT BROUGHT BACK TO OR TO HAVE NEEDLE REMOVED? NO. ARE THERE ANY PLANS TO REMOVE THE NEEDLE? IF YES, PLEASE INDICATE DATE. NO. THE SURGEON HAS DETERMINED THAT THE RETENTION OF THE NEEDLE IS PREFERRED BECAUSE ATTEMPTING TO REMOVE THE NEEDLE WOULD DO MORE HARM THAN LEAVING THE NEEDLE AS IS. THE SURGEON FEELS ATTEMPTING TO REMOVE COULD INCREASE THE RISK OF RECURRENT HERNIA. IS THERE ANY ADVERSE PATIENT OUTCOME? NO. WILL ANY ACTUAL OR REPRESENTATIVE SAMPLE BE RETURNED FOR EVALUATION? YES, THE SALES REP HAS THE PORTION OF THE NEEDLE OF THE NEEDLE THAT WAS RETAINED. IF THE CUSTOMER IS NOT RETURNING THE BROKEN NEEDLE, CAN WE OBTAIN PHOTOS OF THE BROKEN NEEDLE FOR EVALUATION? THE SALES REP ALREADY SENT PHOTOS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6744844&pc=GAM | 2210968-2017-33124 |  | Malfunction | ETHICON INC. | 2017-07-26 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Component Incompatible | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS FOUND THAT THE SHAPE OF THE NEEDLE WAS TAPER POINT THOUGH IT SHOULD HAVE BEEN A REVERSE CUTTING NEEDLE ACCORDING TO THE SPECIFICATION. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: SAMPLES WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION NO DEFECTS WERE FOUND ON THE PACKAGING, THE SAMPLES WERE OPENED AND THE LABELED INFORMATION IN THE PAPER LID BELONGS TO THE FOIL INFORMATION. ALSO, THE PAPER LID WAS REMOVED AND THE NEEDLE/SUTURE COMBINATIONS WERE EXAMINED AND IT WAS NOTED THAT THREE OF TWELVE NEEDLES WERE DIFFERENT TO OTHERS. FOR THIS REASON, A CHARACTERIZATION WAS PERFORMED . NINE NEEDLE SUTURES COMBINATIONS WERE RELATED TO A DRILL NEEDLE (40 MIL) POINT TYPE REVERSE CUTTING EDGE WITH VICRYL UNDYED SUTURE IN DIAMETER 2-0¿ AND LENGTH 27¿ AND ACCORDING THE CHARACTERISTIC, THE SAMPLES BELONG TO PRODUCT CODE VCP869. HOWEVER, THREE NEEDLE SUTURE COMBINATIONS, THE NEEDLES WERE RELATED TO A DRILL NEEDLE, POINT TYPE TAPER POINT WITH VICRYL UNDYED SUTURE IN DIAMETER 2-0¿ AND LENGTH 27¿. ACCORDING TO THE SAMPLE CONDITION THE ASSIGNABLE CAUSE IS MIXED NEEDLES. Manufacturer Narrative: TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CONFIRM THAT 1 DEVICE WAS USED ON THE PATIENT AND THE OTHER 12 ARE REPRESENTATIVE SAMPLES THAT WERE OPENED AND NOT USED ON THE PATIENT. ALL OF THE 13 PACKAGES WERE NOT USED ON THE PATIENT. TWO OF THEM WERE OPENED DURING THE PROCEDURE, AND THE REST (11 PACKAGES) ARE THE REPRESENTATIVE SAMPLES. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CONFIRM THAT 1 DEVICE WAS USED ON THE PATIENT AND THE OTHER 12 ARE REPRESENTATIVE SAMPLES THAT WERE OPENED AND NOT USED ON THE PATIENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6744894&pc=GAM | 2210968-2017-33126 |  | Malfunction | ETHICON INC. | 2017-07-26 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Component Incompatible | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS FOUND THAT THE SHAPE OF THE NEEDLE WAS TAPER POINT THOUGH IT SHOULD HAVE BEEN A REVERSE CUTTING NEEDLE ACCORDING TO THE SPECIFICATION. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CONFIRM THAT 1 DEVICE WAS USED ON THE PATIENT AND THE OTHER 12 ARE REPRESENTATIVE SAMPLES THAT WERE OPENED AND NOT USED ON THE PATIENT. ALL OF THE 13 PACKAGES WERE NOT USED ON THE PATIENT. TWO OF THEM WERE OPENED DURING THE PROCEDURE, AND THE REST (11 PACKAGES) ARE THE REPRESENTATIVE SAMPLES. SAMPLES WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION NO DEFECTS WERE FOUND ON THE PACKAGING, THE SAMPLES WERE OPENED AND THE LABELED INFORMATION IN THE PAPER LID BELONGS TO THE FOIL INFORMATION. ALSO, THE PAPER LID WAS REMOVED AND THE NEEDLE/SUTURE COMBINATIONS WERE EXAMINED AND IT WAS NOTED THAT THREE OF TWELVE NEEDLES WERE DIFFERENT TO OTHERS. FOR THIS REASON, A CHARACTERIZATION WAS PERFORMED . NINE NEEDLE SUTURES COMBINATIONS WERE RELATED TO A DRILL NEEDLE (40 MIL) POINT TYPE REVERSE CUTTING EDGE WITH VICRYL UNDYED SUTURE IN DIAMETER 2-0¿ AND LENGTH 27¿ AND ACCORDING THE CHARACTERISTIC, THE SAMPLES BELONG TO PRODUCT CODE VCP869. HOWEVER, THREE NEEDLE SUTURE COMBINATIONS, THE NEEDLES WERE RELATED TO A DRILL NEEDLE, POINT TYPE TAPER POINT WITH VICRYL UNDYED SUTURE IN DIAMETER 2-0¿ AND LENGTH 27¿. ACCORDING TO THE SAMPLE CONDITION THE ASSIGNABLE CAUSE IS MIXED NEEDLES. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 07/26/2017. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CONFIRM THAT 1 DEVICE WAS USED ON THE PATIENT AND THE OTHER 12 ARE REPRESENTATIVE SAMPLES THAT WERE OPENED AND NOT USED ON THE PATIENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6745021&pc=GAM | 2210968-2017-33128 | 2017-04-25 | Malfunction | ETHICON INC. | 2017-07-26 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN OOPHORECTOMY PERFORMED ON (B)(6) 2017 AND SUTURE WAS USED. POST-OPERATIVELY, ON (B)(6) 2017 THE ABDOMINAL WALL SUTURE WAS BROKEN. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE REPRESENTATIVE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLES WERE OPENED AND THE SWAGE AND ATTACHMENT AREA OF THE SAMPLE WERE AS EXPECTED. ALL SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE OBSERVED. THE SAMPLES WERE TESTED FOR KNOT TENSILE STRENGTH AND THE RESULT IS ABOVE THE MINIMUM INDIVIDUAL STRENGTH REQUIREMENT . PER THE CONDITION OF THE REPRESENTATIVE SAMPLES, NO SUTURE BREAKAGE WAS FOUND ON THE SUTURES AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 07/26/2017. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6745073&pc=GAM | 2210968-2017-33131 | 2017-04-25 | Malfunction | ETHICON INC. | 2017-07-26 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT A PANCREATIC CYST INCISION ON (B)(6) 2017 AND SUTURE WAS USED. POST OPERATIVELY, THE SUTURE BROKE ON (B)(6) 2017. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: DURING THE VISUAL INSPECTION OF THE ACTUAL SAMPLE, IT WAS OBSERVED THAT THE SUTURE PIECES WAS ADHERED TO THE BAG, DUE TO THAT THE SAMPLES ARE AN ABSORBABLE SUTURE AND THE DEGRADATION PROCESS ALREADY BEGUN. DUE TO CONDITION OF THE SAMPLE RETURNED, THE EVALUATION CANNOT BE PERFORMED TO DETERMINE THE ASSIGNABLE CAUSE OF PERFORMANCE - BREAKAGE SUTURE POST-OP. DURING THE VISUAL INSPECTION OF UNOPENED REPRESENTATIVE SAMPLES NO DEFECTS WERE OBSERVED ON THE PACKETS. ACCORDING THE REPRESENTATIVE SAMPLES CONDITION, NO DEFECTS OR SUTURE BREAKAGE WAS FOUND AND THE TESTED SAMPLES MET THE FG REQUIREMENTS. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6728431&pc=GAM | 2210968-2017-33025 | 2017-02-24 | Malfunction | ETHICON INC. | 2017-07-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY THE AFFILIATE THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. THE ANASTOMOSIS WAS GAPING. CURRENTLY THE PATIENT IS RECOVERING. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THE SUTURE BREAK AFTER THE OPERATION WAS COMPLETE? DID THE SUTURE BREAK DURING THE PROCEDURE? WAS THE PROCEDURE ON (B)(6) 2017? DID THE SUTURE BREAK ON (B)(6) 2017? WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED FOR THIS PATIENT? THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. SIX POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721224&pc=GAM | 2210968-2017-32984 | 2017-06-22 | Malfunction | ETHICON INC. | 2017-07-18 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Incorrect Device Or Component Shipped | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CATARACT REPAIR PROCEDURE ON (B)(6)2017 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGEON NOTICED THAT THE SHAPE OF THE NEEDLE WHICH HE REGARDED AS V549G WAS UNUSUAL. IT WAS FOUND THAT THE INNER PACKAGE WAS NOT FOR V549G BUT FOR V449G. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH KL6931, EXP. DATE: 09/30/2021; DATE OF MFG: 10/26/2016. BATCH KKZ567, EXP. DATE: 08/31/2021; DATE OF MFG: 09/15/2016. BATCH KD6190, EXP. DATE: 03/31//2021; DATE OF MFG: 04/19/2016. THE DEVICE HISTORY RECORDS WERE REVIEWED FOR THE POSSIBLE BATCH NUMBERS KKZ567, KL6931, KD6190 AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6716917&pc=GAM | 2210968-2017-32958 | 2017-06-22 | Malfunction | ETHICON INC. | 2017-07-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Incorrect Device Or Component Shipped | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CATARACT REPAIR PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGEON NOTICED THAT THE SHAPE OF THE NEEDLE WHICH HE REGARDED AS V549G WAS UNUSUAL. IT WAS FOUND THAT THE INNER PACKAGE WAS NOT FOR V549G BUT FOR V449G. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 07/17/2017. ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH KL6931, EXP. DATE: 09/30/2021; DATE OF MFG: 10/26/2016 BATCH KKZ567, EXP. DATE: 08/31/2021; DATE OF MFG: 09/15/2016 BATCH KD6190. THE DEVICE HISTORY RECORDS WERE REVIEWED FOR THE POSSIBLE BATCH NUMBERS KKZ567 AND KL6931 AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6716991&pc=GAM | 2210968-2017-32959 | 2017-06-22 | Malfunction | ETHICON INC. | 2017-07-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Incorrect Device Or Component Shipped | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CATARACT REPAIR PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGEON NOTICED THAT THE SHAPE OF THE NEEDLE WHICH HE REGARDED AS V549G WAS UNUSUAL. IT WAS FOUND THAT THE INNER PACKAGE WAS NOT FOR V549G BUT FOR V449G. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 07/17/2017. ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH KL6931, EXP. DATE: 09/30/2021; DATE OF MFG: 10/26/2016 BATCH KKZ567, EXP. DATE: 08/31/2021; DATE OF MFG: 09/15/2016 BATCH KD6190. THE DEVICE HISTORY RECORDS WERE REVIEWED FOR THE POSSIBLE BATCH NUMBERS KKZ567 AND KL6931 AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6714921&pc=GAM | 2210968-2017-32952 | 2017-06-22 | Malfunction | ETHICON INC. | 2017-07-14 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | Device Embedded In Tissue or Plaque | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT HEMORRHOIDECTOMY ON (B)(6) 2017 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE AND IT WAS LEFT INSIDE OF THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE NEEDLE PIECE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? WAS THE NEEDLE PIECE RETAINED AND LEFT INSIDE THE PATIENT? IF THE NEEDLE PIECE WAS LEFT INSIDE THE PATIENT, PLEASE ADVISE: ARE THERE ANY PLANS TO REMOVE THE NEEDLE PIECE? WHAT WAS THE NEEDLE SIZE USED? WHAT IS THE SIZE OF THE NEEDLE PIECE THAT WAS RETAINED? WAS X-RAY PERFORMED TO CONFIRM NEEDLE PIECE WAS RETAINED IN THE TISSUE? IS THERE ANY CHANGE IN THE PATIENT¿S POST-OPERATIVE COURSE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6656388&pc=GAM | 2210968-2017-32706 | 2017-06-08 | Malfunction | ETHICON INC. | 2017-06-21 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC TOTAL HYSTERECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OP, IT WAS NOTED THAT THE SUTURE ABSORBED NOT WELL. THE SURGEON CLEANED THE WOUND AND THE PATIENT CONDITION TURNED BETTER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE PROVIDE SPECIFIC INFORMATION REGARDING WOUND CLEANING PERFORMED BY SURGEON. WAS THERE ANY INFECTION? HOW TREATED? THERE WAS NO INFECTION. I DON'T HAVE OPERATION DETAIL. Manufacturer Narrative: ONE REPRESENTATIVE SAMPLE WAS RETURNED FOR EVALUATION. THE SAMPLE WAS EXAMINED FOR VISUAL DEFECTS. THE PACKET WAS OPENED AND THE NEEDLE WAS ATTACHED TO THE SUTURE. THE SWAGE AREA OF THE NEEDLE/SUTURE WAS EXAMINED AND NO ATTACHMENT DEFECTS WERE NOTED. THE SUTURE WAS DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS WERE FOUND. THE SUTURE DID NOT PRESENT DEFECTS. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS WILL BE MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE PROVIDE SPECIFIC INFORMATION REGARDING WOUND CLEANING PERFORMED BY SURGEON. WAS THERE ANY INFECTION? HOW TREATED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6656583&pc=GAM | 2210968-2017-32707 | 2017-06-08 | Malfunction | ETHICON INC. | 2017-06-21 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TRANSVAGINAL CERVICAL COGNIZATION ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OP, IT WAS NOTED THAT SUTURE ABSORBED NOT WELL. THE SURGEON CLEANED THE WOUND AND THE PATIENT CONDITION TURNED BETTER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE PROVIDE SPECIFIC INFORMATION REGARDING WOUND CLEANING PERFORMED BY SURGEON. WAS THERE ANY INFECTION? HOW TREATED? I DON'T HAVE DETAIL. THERE IS NO INFECTION. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE PROVIDE SPECIFIC INFORMATION REGARDING WOUND CLEANING PERFORMED BY SURGEON. WAS THERE ANY INFECTION? HOW TREATED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6656900&pc=GAT | 6656900 | 2017-06-05 | Malfunction | ETHICON INC. | 2017-06-21 | GAT | VICRYL (POLYGLACTIN 910) SUTURE | Packaging Problem | No Information |  |  | 1 | Event Description: WHEN THE SUTURE WAS OPENED ON THE STERILE FIELD, IT HAD TWO NEEDLES WHEN IT WAS ONLY SUPPOSED TO HAVE ONE NEEDLE IN THE PACKAGE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6653635&pc=GAM | 2210968-2017-32681 |  | Malfunction | ETHICON INC. | 2017-06-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OOPHORECTOMY IN (B)(6) 2017 AND SUTURE WAS USED. THE SUTURE BROKE POST-OP. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE SUTURE BREAK POST OP? YES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6653687&pc=GAM | 2210968-2017-32683 |  | Malfunction | ETHICON INC. | 2017-06-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A RABBIT UNDERWENT AN OOPHORECTOMY IN (B)(6) 2017 AND SUTURE WAS USED FOR CUTANEOUS CLOSURE. THE DAY AFTER THE SURGERY, THE SUTURE WAS BROKEN. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6650298&pc=GAM | 2210968-2017-32639 | 2017-05-31 | Malfunction | ETHICON INC. | 2017-06-19 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCTECTOMY ON (B)(6)2017 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN A CIRCULATING NURSE PASSED THE NEEDLE-SUTURE TO A SCRUB NURSE, THEY NOTICED THAT IT WAS A DOUBLE-ARMED SUTURE THOUGH SHOULD HAVE BEEN A SINGLE ARMED SUTURE. IT WAS NOT USED, AND ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF THE ASSEMBLY - INCORRECT COMPONENT CANNOT BE DETERMINED SINCE THE NEEDLE/SUTURE COMBINATIONS BELONG TO THE PRODUCT CODE J311H. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6640387&pc=GAM | 2210968-2017-32525 | 2017-05-25 | Malfunction | ETHICON INC. | 2017-06-14 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. PRIOR TO THE PROCEDURE, IT WAS NOTICED THAT A NEEDLE SHAPE DID NOT FIT THE DESCRIPTION ON THE SUTURE PACKAGING. THERE WAS A MISMATCH BETWEEN THE NEEDLE DESCRIPTION ON THE PACKAGING AND THE ACTUAL NEEDLE OF THE SUTURE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). THE REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. THE PHOTO OF THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED ONLY, HOWEVER, EVALUATION OF THE PHOTO IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6640410&pc=GAM | 2210968-2017-32528 | 2017-05-25 | Malfunction | ETHICON INC. | 2017-06-14 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. PRIOR TO THE PROCEDURE, IT WAS NOTICED THAT A NEEDLE SHAPE DID NOT FIT THE DESCRIPTION ON THE SUTURE PACKAGING. THERE WAS A MISMATCH BETWEEN THE NEEDLE DESCRIPTION ON THE PACKAGING AND THE ACTUAL NEEDLE OF THE SUTURE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). THE REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. THE PHOTO OF THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED ONLY, HOWEVER, EVALUATION OF THE PHOTO IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6640414&pc=GAM | 2210968-2017-32526 | 2017-05-25 | Malfunction | ETHICON INC. | 2017-06-14 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. PRIOR TO THE PROCEDURE, IT WAS NOTICED THAT A NEEDLE SHAPE DID NOT FIT THE DESCRIPTION ON THE SUTURE PACKAGING. THERE WAS A MISMATCH BETWEEN THE NEEDLE DESCRIPTION ON THE PACKAGING AND THE ACTUAL NEEDLE OF THE SUTURE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). THE REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. THE PHOTO OF THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED ONLY, HOWEVER, EVALUATION OF THE PHOTO IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6640415&pc=GAM | 2210968-2017-32527 | 2017-05-25 | Malfunction | ETHICON INC. | 2017-06-14 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. PRIOR TO THE PROCEDURE, IT WAS NOTICED THAT A NEEDLE SHAPE DID NOT FIT THE DESCRIPTION ON THE SUTURE PACKAGING. THERE WAS A MISMATCH BETWEEN THE NEEDLE DESCRIPTION ON THE PACKAGING AND THE ACTUAL NEEDLE OF THE SUTURE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). THE REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. THE PHOTO OF THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED ONLY, HOWEVER, EVALUATION OF THE PHOTO IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6639077&pc=GAM | 2210968-2017-32516 | 2017-04-28 | Malfunction | ETHICON INC. | 2017-06-13 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT AN ANIMAL UNDERWENT A MASS REMOVAL AND SPADE PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. FIVE TO SEVEN DAYS POST OP, THE SUTURE BROKE IN THE MIDDLE OF THE SUTURE LINE. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 06/13/2017. THE REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587494&pc=GAM | 2210968-2017-32247 | 2017-04-24 | Malfunction | ETHICON INC. | 2017-05-24 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2017, PRIOR TO AN UNKNOWN SURGICAL PROCEDURE, A HOLE WAS FOUND IN THE POUCH. THE PRODUCT WAS NOT USED ON PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: AN EMPTY OPENED FOIL, ONE PAPER LID AND ONE NEEDLE/SUTURE INTO WINDING FORMER WAS RETURNED FOR ANALYSIS DURING THE VISUAL INSPECTION OF THE SAMPLE, IT WAS OBSERVED THAT THE TOP FOIL WAS DAMAGED (BROKEN) NEAR TO THE SIDE PEELABLE FOIL AND END OF THE FOIL. THE BREAK WAS FROM THE OUTSIDE TO INSIDE DUE TO WRINKLES AT THE AREA AND ALSO THE FOIL WAS BROKEN IN A PART OF THE SEAL. IN ADDITION, THE NEEDLE/SUTURE WAS DISPENSED AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WERE OBSERVED. AS SAMPLE WAS RECEIVED OPEN AND IS AN ABSORBABLE SUTURE, THE DEGRADATION PROCESS WAS ALREADY BEGUN AND THE TIME OF EXPOSURE THAT HAS THE SUTURE IN THE ENVIRONMENT COULD NOT BE DETERMINED. HOWEVER, THE SUTURE WAS TESTED BY TENSILE STRENGTH. DUE TO THE SAMPLE CONDITION, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT OF: ¿THAT PRIOR TO AN UNKNOWN PROCEDURE, THERE WAS A HOLE ON THE POUCH¿. THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT WE WERE UNABLE TO DUPLICATE DURING OUR LABORATORY ANALYSIS. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587570&pc=GAM | 2210968-2017-32251 |  | Malfunction | ETHICON INC. | 2017-05-24 | GAM | VICRYL RAPIDE | Appropriate Term/Code Not Available | Pain | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CARPAL TUNNEL PROCEDURE AND SUTURE WAS USED. THE PATIENT STATED THAT THE INCISION IS PAINFUL AT ONE END AND THE SUTURE MATERIAL HAS NOT STARTED TO DISSOLVE. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 05/18/2017. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT DATE DID YOU EXPERIENCE THE SYMPTOMS? WERE YOU EVALUATED BY YOUR SURGEON? WHAT IS YOUR SURGEON OPINION REGARDING YOUR SYMPTOMS? WAS ANY MEDICAL TREATMENT PROVIDED FOR YOUR SYMPTOMS? DOES ETHICON HAVE YOUR PERMISSION TO CONTACT YOUR SURGEON, IN THE EVENT ETHICON WOULD LIKE TO CONTACT YOUR SURGEON FOR MORE CLINICAL INFORMATION TO BE USED FOR A PRODUCT QUALITY COMPLAINT INVESTIGATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6572613&pc=GAM | 2210968-2017-32127 | 2017-04-26 | Malfunction | ETHICON INC. | 2017-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO USE ON A PATIENT ON (B)(6) 2017, IT WAS FOUND THAT THE BOX AND SINGLE UNIT OF SUTURE HAD AN ADDITIONAL LABEL (COVERING THE ORIGINAL ONE) WITH A WRONG (B)(4) DESCRIPTION. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: RECEIVED AN OPENED BOX WITH NINE UNOPENED SAMPLES FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE BOX, AN ADDITIONAL LABEL WAS STUCK IN ONE SIDE OF THE BOX; ALSO, THE FOILS HAS AN ADDITIONAL LABEL ON THE BOTTOM. THE INFORMATION ON THE ADDITIONAL LABEL DOES NOT MATCH WITH THE PRINT ON THE BOX AND FOIL DUE TO THE INFORMATION IN THE ADDITIONAL LABEL IS WRONG ( SUTURE SURGICAL, NOT ABSORBABLE, BLACK MONOFILAMENT). THE PRINT INFORMATION ON THE BOX AND FOILS IS VIOLET BRAIDED ABSORBABLE SUTURE. IN ADDITION, THE FOILS WERE OPENED AND THE NEEDLE/SUTURE COMBINATIONS INTO FOILS BELONG TO A VIOLET BRAIDED ABSORBABLE VICRYL SUTURE. ACCORDING TO THE SAMPLE CONDITION, THE ASSIGNABLE CAUSE OF PACKAGING - LABELING IDENTIFICATION IS INCORRECT LABEL COPY. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THIS DISCOVERED BY A (B)(4) ASSOCIATE PRIOR TO SALE? THE PRODUCT WAS DELIVERED TO THE DISTRIBUTOR. THE FINAL CUSTOMER CLAIMED THE ISSUE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6572647&pc=GAM | 2210968-2017-32128 | 2017-04-26 | Malfunction | ETHICON INC. | 2017-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO USE ON A PATIENT ON (B)(6) 2017, IT WAS FOUND THAT THE BOX AND SINGLE UNIT OF SUTURE HAD AN ADDITIONAL LABEL (COVERING THE ORIGINAL ONE) WITH A WRONG (B)(4) DESCRIPTION. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: RECEIVED THREE SEALED BOXES WITH (B)(4) UNOPENED SAMPLES FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE BOX, AN ADDITIONAL LABEL WAS STUCK IN ONE SIDE OF THE BOX; ALSO, THE FOILS HAS AN ADDITIONAL LABEL ON THE BOTTOM. THE INFORMATION ON THE ADDITIONAL LABEL DOES NOT MATCH WITH THE PRINT ON THE BOX AND FOIL DUE TO THE INFORMATION IN THE ADDITIONAL LABEL IS WRONG (SUTURE SURGICAL, NOT ABSORBABLE, BLACK MONOFILAMENT). THE PRINT INFORMATION ON THE BOX AND FOILS IS VIOLET BRAIDED ABSORBABLE SUTURE. IN ADDITION, THE FOILS WERE OPENED AND THE NEEDLE/SUTURE COMBINATIONS INTO FOILS BELONG TO A VIOLET BRAIDED ABSORBABLE VICRYL SUTURE. ACCORDING TO THE SAMPLE CONDITION, THE ASSIGNABLE CAUSE OF PACKAGING - LABELING IDENTIFICATION IS INCORRECT LABEL COPY. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THIS DISCOVERED BY A J&J ASSOCIATE PRIOR TO SALE? THE PRODUCT WAS DELIVERED TO THE DISTRIBUTOR. THE FINAL CUSTOMER CLAIMED THE ISSUE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6572649&pc=GAM | 2210968-2017-32129 | 2017-04-26 | Malfunction | ETHICON INC. | 2017-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO USE ON A PATIENT ON (B)(6) 2017, IT WAS FOUND THAT THE BOX AND SINGLE UNIT OF SUTURE HAD AN ADDITIONAL LABEL (COVERING THE ORIGINAL ONE) WITH A WRONG (B)(4) DESCRIPTION. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: RECEIVED THREE SEALED BOXES WITH (B)(4) UNOPENED SAMPLES EACH AND ONE OPENED BOX WITH (B)(4) UNOPENED SAMPLES FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE BOX, AN ADDITIONAL LABEL WAS STUCK IN ONE SIDE OF THE BOX; ALSO, THE FOILS HAS AN ADDITIONAL LABEL ON THE BOTTOM. THE INFORMATION ON THE ADDITIONAL LABEL DOES NOT MATCH WITH THE PRINT ON THE BOX AND FOIL DUE TO THE INFORMATION IN THE ADDITIONAL LABEL IS WRONG (SUTURE SURGICAL, NOT ABSORBABLE, BLACK MONOFILAMENT). THE PRINT INFORMATION ON THE BOX AND FOILS IS VIOLET BRAIDED ABSORBABLE SUTURE. IN ADDITION, THE FOILS WERE OPENED AND THE NEEDLE/SUTURE COMBINATIONS INTO FOILS BELONG TO A VIOLET BRAIDED ABSORBABLE VICRYL SUTURE. ACCORDING TO THE SAMPLE CONDITION, THE ASSIGNABLE CAUSE OF PACKAGING - LABELING IDENTIFICATION IS INCORRECT LABEL COPY. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THIS DISCOVERED BY A J&J ASSOCIATE PRIOR TO SALE? THE PRODUCT WAS DELIVERED TO THE DISTRIBUTOR. THE FINAL CUSTOMER CLAIMED THE ISSUE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6572658&pc=GAM | 2210968-2017-32130 | 2017-04-26 | Malfunction | ETHICON INC. | 2017-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO USE ON A PATIENT ON (B)(6) 2017, IT WAS FOUND THAT THE BOX AND SINGLE UNIT OF SUTURE HAD AN ADDITIONAL LABEL (COVERING THE ORIGINAL ONE) WITH A WRONG (B)(4) DESCRIPTION. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: RECEIVED FIVE SEALED BOXES WITH (B)() UNOPENED SAMPLES. DURING THE VISUAL INSPECTION OF THE BOX, AN ADDITIONAL LABEL WAS STUCK IN ONE SIDE OF THE BOX; ALSO, THE FOILS HAS AN ADDITIONAL LABEL ON THE BOTTOM. THE INFORMATION ON THE ADDITIONAL LABEL DOES NOT MATCH WITH THE PRINT ON THE BOX AND FOIL DUE TO THE INFORMATION IN THE ADDITIONAL LABEL IS WRONG (SUTURE SURGICAL, NOT ABSORBABLE, BLACK MONOFILAMENT). THE PRINT INFORMATION ON THE BOX AND FOILS IS VIOLET BRAIDED ABSORBABLE SUTURE. IN ADDITION, THE FOILS WERE OPENED AND THE NEEDLE/SUTURE COMBINATIONS INTO FOILS BELONG TO A VIOLET BRAIDED ABSORBABLE VICRYL SUTURE. ACCORDING TO THE SAMPLE CONDITION, THE ASSIGNABLE CAUSE OF PACKAGING - LABELING IDENTIFICATION IS INCORRECT LABEL COPY. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THIS DISCOVERED BY A (B)(4) ASSOCIATE PRIOR TO SALE? THE PRODUCT WAS DELIVERED TO THE DISTRIBUTOR. THE FINAL CUSTOMER CLAIMED THE ISSUE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6556699&pc=GAM | 2210968-2017-32007 | 2017-04-18 | Malfunction | ETHICON INC. | 2017-05-10 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TRABECULECTOMY PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED. DURING THE PROCEDURE, A DIFFERENCE IN THE NEEDLE, WHICH WAS OUTLINED ON THE PACKAGING, WAS NOTICED; THE NEEDLE WAS CHANGED FROM A PRECISION POINT SPATULATED NEEDLE TO A PRECISION POINT REVERSE CUTTING NEEDLE. THE SURGEON REPORTED THAT THE NAME OF THE NEEDLE STAYED THE SAME, WHICH CREATED CONFUSION. THE SURGEON WAS CONCERNED AS A SPATULATED NEEDLE WAS REQUIRED FOR THIS PROCEDURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. DURING THE VISUAL INSPECTION OF THE SAMPLE, THE DESCRIPTION OF THE NEEDLE AT THE FOIL CORRESPONDS TO A MICRO POINT REVERSE-CUTTING NEEDLE. AT THAT TIME, THE LABELING WAS ACCORDING TO THE REQUIREMENTS OF (B)(4). HOWEVER, THE NEEDLE ATTACHED TO THE SUTURE CORRESPONDS TO A MICRO POINT SPATULA. DUE TO THE DIFFERENCE, A CHANGE WAS PERFORMED ON NOVEMBER 8, 2016 TO ALIGN THE LABELING WITH THE PRODUCT INSIDE ON THE PACKAGE. THE MANUFACTURE OF SAMPLE RETURN IS OCTOBER 21, 2016, BEFORE THE CHANGE WAS MADE. ACCORDING TO THE SAMPLE CONDITION, THE ASSIGNABLE ACCUSE OF PACKAGING - LABELING IDENTIFICATION IS DUE AN ISSUE DURING THE PACKAGING PROCESS. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE ADDITIONAL INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE PROVIDE A PHOTO OF THE PACKAGING AND PRODUCT IF AVAILABLE. PLEASE RETURN THE PRODUCT AND PACKAGING FOR ANALYSIS. PLEASE PROVIDE THE LOT NUMBER INVOLVED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6544340&pc=GAM | 2210968-2017-31931 | 2017-03-15 | Malfunction | ETHICON INC. | 2017-05-04 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HYSTERECTOMY PROCEDURE AND SUTURE WAS USED. AFTER THE PROCEDURE, THE SURGEON SAW THAT THE SUTURE ABSORBED "TOO QUICKLY¿. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: UNOPENED REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION AND EXAMINED FOR VISUAL DEFECTS: ACCORDING TO THE SAMPLE CONDITION, THE SAMPLES MEET THE FINISH GOOD REQUIREMENTS AND NO PERFORMANCE - ABSORPTION OF SUTURE COULD BE OBSERVED. Manufacturer Narrative: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. THEY WERE VISUALLY AND FUNCTIONALLY EXAMINED FOR STRENGTH AND THEY MET REQUIREMENTS. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IN HOW MANY DAYS DID THE SUTURE ABSORB? WHAT MEDICAL OR SURGICAL INTERVENTION WAS REQUIRED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6527003&pc=GAM | 2210968-2017-31851 | 2017-03-15 | Malfunction | ETHICON INC. | 2017-04-27 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HYSTERECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. AFTER THE PROCEDURE, THE SURGEON SAW THAT THE SUTURE ABSORBED "TOO QUICKLY¿. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: THE REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. Manufacturer Narrative: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION AND EXAMINED FOR VISUAL DEFECTS: ACCORDING TO THE SAMPLE CONDITION, NO DEFECTS WERE OBSERVED AND NO PERFORMANCE - ABSORPTION OF SUTURE COULD BE OBSERVED. Manufacturer Narrative: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. THEY WERE VISUALLY AND FUNCTIONALLY EXAMINED FOR STRENGTH AND THEY MET REQUIREMENTS. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IN HOW MANY DAYS DID THE SUTURE ABSORB? WHAT MEDICAL OR SURGICAL INTERVENTION WAS REQUIRED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6527004&pc=GAM | 2210968-2017-31850 | 2017-03-15 | Malfunction | ETHICON INC. | 2017-04-27 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HYSTERECTOMY PROCEDURE AND SUTURE WAS USED. AFTER THE PROCEDURE, THE SURGEON SAW THAT THE SUTURE ABSORBED "TOO QUICKLY¿. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: THE REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: UNOPENED REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION AND EXAMINED FOR VISUAL DEFECTS: ACCORDING TO THE SAMPLE CONDITION, THE SAMPLES MEET THE FINISH GOOD REQUIREMENTS AND NO PERFORMANCE - ABSORPTION OF SUTURE COULD BE OBSERVED. Manufacturer Narrative: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. THEY WERE VISUALLY AND FUNCTIONALLY EXAMINED FOR STRENGTH AND THEY MET REQUIREMENTS. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IN HOW MANY DAYS DID THE SUTURE ABSORB? WHAT MEDICAL OR SURGICAL INTERVENTION WAS REQUIRED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6521862&pc=GAM | 6521862 | 2017-04-01 | Malfunction | ETHICON, INC. | 2017-04-26 | GAM | VICRYL SUTURE | Break; Detachment Of Device Component; Patient-Device Incompatibility | Foreign Body In Patient |  |  | 1 | Event Description: DURING LAPAROSCOPIC CHOLECYSTECTOMY, CLOSURE OF FASCIAL INCISIONS WAS DIFFICULT DUE TO PATIENT'S BODY HABITUS, AND THE SUBXIPHOID INCISION, VICRYL NEEDLE BROKE. BOTH PORTIONS WERE RETRIEVED. NO PATIENT HARM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6508734&pc=GAM | 2210968-2017-31705 |  | Malfunction | ETHICON INC. | 2017-04-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT OVARY REMOVAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST OPERATIVELY THE SUTURE BROKE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE SUTURE BREAK POST OP? YES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6508735&pc=GAM | 2210968-2017-31704 |  | Malfunction | ETHICON INC. | 2017-04-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT OVARY REMOVAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST OPERATIVELY THE SUTURE BROKE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE SUTURE BREAK POST OP? YES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6508741&pc=GAM | 2210968-2017-31700 |  | Malfunction | ETHICON INC. | 2017-04-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT OVARY REMOVAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST OPERATIVELY THE SUTURE BROKE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE SUTURE BREAK POST OP? YES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6493188&pc=GAM | 6493188 | 2017-03-22 | Malfunction | ETHICON, INC. | 2017-04-14 | GAM | VICRYL | Break | Device Embedded In Tissue or Plaque |  |  | 1 | Event Description: AT THE START OF THE PROCEDURE, ONE ATRAUMATIC 0 VICRYL UR-6 WAS USED AND BROKEN WHEN SUTURING THE DEVICE INTO ABDOMINAL TISSUE. BOTH PIECES OF THIS ATRAUMATIC WERE OBTAINED AND COLLECTED IMMEDIATELY. INTRA-OPERATIVE RADIOLOGY WAS PERFORMED PER HOSPITAL POLICY. THE DOCTOR AND RADIOLOGIST WERE CALLED AND SPOKE WITH THE SURGEON. THE X-RAY WAS NEGATIVE. MANUFACTURER RESPONSE FOR NEEDLE, 0 VICRYL SUTURE (PER SITE REPORTER): WE HAVE REQUESTED A MEETING WITH VENDOR THROUGH MATERIALS MANAGEMENT. NEW MANAGER IS WORKING WITH US. THE TEAM FOLLOWED POLICY. BROKEN NEEDLE PIECES RETRIEVED, X-RAY TAKEN TO ENSURE NO PIECES LEFT BEHIND. X-RAY CLEARED BY RADIOLOGY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6464198&pc=GAM | 2210968-2017-31561 | 2017-03-13 | Malfunction | ETHICON INC. | 2017-04-06 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PTERYGIUM TRANSPLANT SURGERY ON (B)(6) 2017 AND SUTURE WAS USED. THE SURGEON NOTICED THAT THE WIRE WASN'T A SINGLE STRAND BUT IT WAS BRAIDED. THE PROCEDURE WAS COMPLETED BY REPLACING THE SUTURE WITH A SINGLE WIRE OF A LITTLE DIAMETER (10/0). NO PATIENT ADVERSE EVENTS HAVE BEEN REPORTED. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: A REPRESENTATIVE SAMPLE WAS RECEIVED. VISUAL INSPECTION WAS PERFORMED AND THE FOLDER CONTAINED THE CORRECT SUTURE MATERIAL, CONFORMING WITH DESCRIPTION ON THE PACKAGE: V600, COATED VICRYL, USP 9-0, VIOLET, BRAIDED ABSORBABLE SUTURE, SINGLE ARMED WITH GS-9 NEEDLE. AS STATED ON THE PACKAGE, SUTURE OF THIS PRODUCT IS BRAIDED, THUS NOT A MONOFILE BUT A POLYFILE SUTURE. THEREFORE USER ERROR WAS DETERMINED TO BE THE CAUSE OF THE ISSUE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS REPORTED THAT THE SURGEON NOTICED ¿THAT THE WIRE WASN'T A SINGLE STRAND BUT IT WAS BRAIDED.¿ THE REPORTED PRODUCT CODE IS V600G WHICH IS VICRYL POLYGLACTIN 910 SUTURE. VICRYL SUTURE IS A BRAIDED SUTURE. COULD YOU PLEASE CLARIFY THIS COMPLAINT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6461336&pc=GAM | 2210968-2017-31536 | 2017-03-14 | Malfunction | ETHICON INC. | 2017-04-05 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient; Device Embedded In Tissue or Plaque | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GYNECOMASTIA BREAST PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED. DURING THE CLOSURE, THE NEEDLE BROKE IN TWO HALVES EXACTLY IN THE CENTER AND FELL INTO THE CAVITY. THE SURGEON WAS USING C-ARM TO FIND THE NEEDLE PIECE. IT WAS ALSO REPORTED THAT THE NEEDLE PIECE WAS REMOVED USING C-ARM AND MORE THAN HALF THE NEEDLE RETAINED IN THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE ADDITIONAL INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE DESCRIBE WHAT ADDITIONAL TISSUE DAMAGE OCCURRED? LOCATION AND SIZE OF ADDITIONAL TISSUE DAMAGE? THE ACTUAL SAMPLE WAS RECEIVED FOR EVALUATION. THE SAMPLE CONSISTED OF A NEEDLE BROKEN IN TWO PIECES. BOTH NEEDLE PIECES WERE VISUALLY INSPECTED. NEEDLE PIECE # 1 SHOWED SEVERAL USER INSTRUMENT MARKS AND BENT UP APPEARANCE RIGHT BEHIND THE ATTACHMENT AREA, INDICATING THAT THE NEEDLE HAD BEEN GRASPED THERE. NEEDLE PIECE # 2 SHOWED SEVERAL USER INSTRUMENT MARKS AT THE BREAK POINT. CAUSE IS EXTERNAL AND USAGE ERROR. IN ADDITION, RETAIN SAMPLES WERE VISUALLY AND FUNCTIONALLY TESTED. ALL INDIVIDUAL DUCTILITY TEST VALUES WERE FOUND TO BE ABOVE AVERAGE IN-HOUSE REQUIREMENT FOR DUCTILITY TEST. ALL THE INDIVIDUAL STIFFNESS VALUES & DUCTILITY TEST VALUES WERE FOUND TO BE ON OR ABOVE INDIVIDUAL IN-HOUSE REQUIREMENT. THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS. Manufacturer Narrative: THE ACTUAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. THEREFORE, RETAIN SAMPLES WERE VISUALLY AND FUNCTIONALLY TESTED. ALL INDIVIDUAL DUCTILITY TEST VALUES WERE FOUND TO BE ABOVE AVERAGE IN-HOUSE REQUIREMENT FOR DUCTILITY TEST. ALL THE INDIVIDUAL STIFFNESS VALUES & DUCTILITY TEST VALUES WERE FOUND TO BE ON OR ABOVE INDIVIDUAL IN-HOUSE REQUIREMENT. THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN MORE ADDITIONAL/ FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE FOLLOW UP INFORMATION PROVIDED INDICATES THAT THERE WAS ADDITIONAL TISSUE DAMAGE - PLEASE DESCRIBE WHAT ADDITIONAL TISSUE DAMAGE OCCURRED? LOCATION AND SIZE OF ADDITIONAL TISSUE DAMAGE? WAS THE PROCEDURE COMPLETED SUCCESSFULLY? PLEASE PROVIDE ANY UPDATE TO THE PATIENT CURRENT CONDITION. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE NEEDLE PIECE REMOVED, AND HOW? ¿ YES THE NEEDLE PIECE WAS REMOVED USING C-ARM. WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? ¿ THEY USED C-ARM. WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? ¿ YES, THERE WAS ADDITIONAL TISSUE DAMAGED AND IT WAS TIME CONSUMING, WHICH DELAYED THE PROCEDURE. WAS MORE THAN HALF THE NEEDLE RETAINED IN THE PATIENT? - YES MORE THAN HALF THE NEEDLE RETAINED IN THE PATIENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6462774&pc=GAM | 2210968-2017-31542 | 2017-03-13 | Malfunction | ETHICON INC. | 2017-04-05 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient; Device Embedded In Tissue or Plaque | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GASTRECTOMY PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED FOR CLOSURE. DURING THE PROCEDURE, THE NEEDLE BROKE IN TWO HALVES EXACTLY IN THE CENTER AND FELL INTO THE CAVITY. THERE WAS NO ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE. IT WAS ALSO REPORTED THAT MORE THAN HALF OF THE NEEDLE RETAINED IN THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ATTEMPTS ARE BEING MADE TO OBTAIN THE ADDITIONAL INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. TWO PIECES OF NEEDLE WERE RECEIVED FOR ANALYSIS. THEREFORE, IT IS UNCLEAR IF ANY NEEDLE PIECE WAS RETAINED IN THE PATIENT. IS A PIECE OF THE NEEDLE FROM PRODUCT CODE VP2317 RETAINED IN THE PATIENT? IS A PIECE OF THE NEEDLE FROM PRODUCT CODE VP2347 RETAINED IN THE PATIENT? WHICH PRODUCT CODE WAS THE 2 NEEDLE PIECES RETURNED RELATED TO ( VP2317 OR FOR VP2347)? Manufacturer Narrative: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GASTRECTOMY PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED FOR CLOSURE. DURING THE PROCEDURE, THE NEEDLE BROKE IN TWO HALVES EXACTLY IN THE CENTER AND FELL INTO THE CAVITY. THERE WAS NO ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE. IT WAS ALSO REPORTED THAT THE BROKEN NEEDLE WAS REMOVED USING C-ARM AND THERE WAS NO DELAY IN THE PROCEDURE TIME. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PRODUCT CODE: VP2317 WAS NOTED IN THE EVENT DESCRIPTION AS NEEDLE BREAKING. ¿ YES NEEDLE BROKE IN THE SURGERY. A 2ND PRODUCT CODE: VP2347 WAS NOTED ON THE SYNERGY REPORT. ¿ BROKEN SAMPLE WAS NOT RETAINED. WAS THIS NEEDLE BROKEN: VP2347 AS WELL ? - YES. AND WAS IT BELIEVED TO HAVE BEEN RETAINED? - IT WAS NOT RETAINED. IF RETAINED IN THE PATIENT, WHAT TISSUE STRUCTURE WAS IT RETAINED IN? ¿ NO IDEA. OR WAS IT REMOVED FROM PATIENT DURING INITIAL PROCEDURE? ¿ YES IT WAS REMOVED USING C -ARM. IS THERE A NEEDLE PIECE RETAINED IN THE PATIENT AT THE CURRENT TIME? ¿ NO IDEA. IF THERE IS A NEEDLE PIECE IS CURRENTLY RETAINED, WHAT STRUCTURE IS IT RETAINED IN? ¿ NO IDEA. ARE THERE ANY PLANS TO REMOVE THE RETAINED PIECE? - RETAINED PIECE IS REMOVED AND THE SAME HAS BEEN SENT TO YOU. ARE THERE ANY ADVERSE PATIENT CONSEQUENCES? ¿ NO DELAY IN THE PROCEDURE TIME. BOTH NEEDLE PIECES WERE VISUALLY INSPECTED. NEEDLE PIECE # 1 SHOWED SEVERAL USER INSTRUMENT MARKS AND BENT UP APPEARANCE RIGHT BEHIND THE ATTACHMENT AREA, INDICATING THAT THE NEEDLE HAD BEEN GRASPED THERE. A SUTURE PIECE WAS STILL ATTACHED. NEEDLE PIECE # 2 SHOWED SEVERAL USER INSTRUMENT MARKS AT THE BREAK POINT. RETAIN SAMPLES WAS RETRIEVED FOR ANALYSIS. NO DEFECTS WERE OBSERVED IN THE RETAIN SAMPLE. THESE RETAIN SAMPLES WERE TESTED AND FOUND TO MEET SPECIFICATION. ALL INDIVIDUAL DUCTILITY TEST VALUES WERE FOUND TO BE ABOVE AVERAGE IN-HOUSE REQUIREMENT Manufacturer Narrative: (B)(4). A FRACTURE WAS OBSERVED IN THE BODY OF THE NEEDLE. ONLY ONE SIDE OF THE FRACTURED NEEDLE WAS RECEIVED, THE MATING FRACTURE SURFACE WAS NOT PROVIDED FOR THIS EVALUATION. THE EVALUATION REVEALED THE FRACTURE WAS PREDOMINATELY INTERGRANULAR, WHICH IS EVIDENCE OF A BRITTLE FAILURE. MICROVOID COALESCENCE WAS OBSERVED ON SOME OF THE INTERGRANULAR FACETS. THIS WAS A NON-DUCTILE FRACTURE. IT IS NOT COMMON FOR A NEEDLE, WHICH TYPICALLY FAILS PRIMARILY IN A DUCTILE MANNER. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IS THERE A NEEDLE PIECE RETAINED IN THE PATIENT AT THE CURRENT TIME? IF THERE IS A NEEDLE PIECE IS CURRENTLY RETAINED, WHAT STRUCTURE IS IT RETAINED IN? ARE THERE ANY PLANS TO REMOVE THE RETAINED PIECE? ARE THERE ANY ADVERSE PATIENT CONSEQUENCES? ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE NEEDLE PIECE REMOVED, AND HOW? ¿ NO NEEDLE PIECE WAS NOT REMOVED. WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? - NO. WAS THE PATIENT BROUGHT BACK TO OR TO HAVE NEEDLE REMOVED OR WAS THE NEEDLE REMOVED DURING THE INITIAL PROCEDURE? ¿ NO, NEEDLE POPPED OUT AS IT BROKE DURING THE PROCEDURE. IT FELL ON THE OT THEATRE AND CIRCULATING NURSE PICKED UP THE NEEDLE. WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? ¿ NO. WAS MORE THAN HALF THE NEEDLE RETAINED IN THE PATIENT? - YES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6462795&pc=GAM | 2210968-2017-31547 | 2017-03-13 | Malfunction | ETHICON INC. | 2017-04-05 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient; Device Embedded In Tissue or Plaque | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GASTRECTOMY PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED FOR CLOSURE. DURING THE PROCEDURE, THE NEEDLE BROKE IN TWO HALVES EXACTLY IN THE CENTER AND FELL INTO THE CAVITY. THERE WAS NO ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE. IT WAS ALSO REPORTED THAT MORE THAN HALF OF THE NEEDLE RETAINED IN THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. IT WAS REPORTED THAT THE PATIENT UNDERWENT A GASTRECTOMY PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED FOR CLOSURE. DURING THE PROCEDURE, THE NEEDLE BROKE IN TWO HALVES EXACTLY IN THE CENTER AND FELL INTO THE CAVITY. THERE WAS NO ADDITIONAL TISSUE DAMAGE DUE TO SEARCHING FOR THE NEEDLE PIECE. IT WAS ALSO REPORTED THAT THE BROKEN NEEDLE WAS REMOVED USING C-ARM AND THERE WAS NO DELAY IN THE PROCEDURE TIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED. ATTEMPTS ARE BEING MADE TO OBTAIN MORE ADDITIONAL INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IS A PIECE OF THE NEEDLE FROM PRODUCT CODE VP2347 RETAINED IN THE PATIENT? ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: A PRODUCT CODE: VP2347 WAS NOTED ON THE SYNERGY REPORT. ¿ BROKEN SAMPLE WAS NOT RETAINED. WAS THIS NEEDLE BROKEN: VP2347 AS WELL ? - YES. AND WAS IT BELIEVED TO HAVE BEEN RETAINED? - IT WAS NOT RETAINED. IF RETAINED IN THE PATIENT, WHAT TISSUE STRUCTURE WAS IT RETAINED IN? ¿ NO IDEA. OR WAS IT REMOVED FROM PATIENT DURING INITIAL PROCEDURE? ¿ YES IT WAS REMOVED USING C -ARM. IS THERE A NEEDLE PIECE RETAINED IN THE PATIENT AT THE CURRENT TIME? ¿ NO IDEA. IF THERE IS A NEEDLE PIECE IS CURRENTLY RETAINED, WHAT STRUCTURE IS IT RETAINED IN? ¿ NO IDEA. ARE THERE ANY PLANS TO REMOVE THE RETAINED PIECE? - RETAINED PIECE IS REMOVED AND THE SAME HAS BEEN SENT TO YOU. ARE THERE ANY ADVERSE PATIENT CONSEQUENCES? ¿ NO DELAY IN THE PROCEDURE TIME. Manufacturer Narrative: ONE REPRESENTATIVE SAMPLE WAS RECEIVED FOR EVALUATION. THE INTACT REPRESENTATIVE SAMPLE WAS INSPECTED FOR APPEARANCE TEST AND IT WAS FOUND TO MEET THE SPECIFICATION REQUIREMENT. FIVE RETAIN SAMPLES WERE RETRIEVED FOR ANALYSIS. THE NEEDLE RETAIN SAMPLES WERE VISUALLY INSPECTED FOR PHYSICAL APPEARANCE LIKE CURVATURE, TIP, ETC. ATTRIBUTE DEFECTS SUCH AS BEND, CRACKED BARREL AND WERE FOUND SATISFACTORY. NO DEFECTS WERE OBSERVED IN THE RETAIN SAMPLE. THESE RETAIN SAMPLES WERE TESTED FOR STIFFNESS & DUCTILITY AND FOUND TO MEET SPECIFICATION. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN MORE ADDITIONAL FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PRODUCT CODE: VP2317 WAS NOTED IN THE EVENT DESCRIPTION AS NEEDLE BREAKING. A 2ND PRODUCT CODE: VP2347 WAS NOTED ON THE SYNERGY REPORT. WAS THIS NEEDLE BROKEN ON VP2347 AS WELL ? AND WAS IT BELIEVED TO HAVE BEEN RETAINED? OR REMOVED FROM PATIENT? ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE NEEDLE PIECE REMOVED, AND HOW? ¿ NO NEEDLE PIECE WAS NOT REMOVED. WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? -NO. WAS THE PATIENT BROUGHT BACK TO OPERATING ROOM TO HAVE NEEDLE REMOVED OR WAS THE NEEDLE REMOVED DURING THE INITIAL PROCEDURE? ¿ NO, NEEDLE POPPED OUT AS IT BROKE DURING THE PROCEDURE. IT FELL ON THE OT THEATRE AND CIRCULATING NURSE PICKED UP THE NEEDLE. WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? ¿ NO. WAS MORE THAN HALF THE NEEDLE RETAINED IN THE PATIENT? - YES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6457235&pc=GAM | 2210968-2017-31502 | 2017-03-10 | Malfunction | ETHICON INC. | 2017-04-04 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. PRIOR TO THE PROCEDURE, THE DISCREPANCY OF THE LOT NUMBER WAS DETECTED. THE LOT MENTIONED ON THE INDIVIDUAL STERILITY PROTECTOR WAS DIFFERENT THAN ON THE PROTECTIVE PACKAGING OF THE SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: LOT KK8DHPP0 WAS REWORKED. REWORKED PEELABLE FOILS HAVE ANOTHER LOT NUMBER AT DIGIT 8 THAN THOSE THAT WERE USED ON TOP OF THE FOLDER. IN THIS DEVICE IT IS LOT KK8DHPP1. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. CAN YOU PLEASE PROVIDE PHOTOS OF THE PACKAGING THAT WERE MISSING THE LOT NUMBER? CAN YOU CLARIFY THE QUANTITY THAT DISPLAYED THE REPORTED ISSUE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6452798&pc=GAM | 6452798 | 2017-03-07 | Malfunction | ETHICON, INC. | 2017-04-03 | GAM | VICRYL | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: DURING CLOSURE OF LEFT HIP SURGICAL INCISION, AN ATRAUMATIC NEEDLE BROKE IN TWO PIECES. INTRA-OP X-RAY OBTAINED, NO FOREIGN BODY IDENTIFIED. THE SURGERY THEN CONCLUDED UNEVENTFULLY. THERE WAS NO HARM TO THE PATIENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6436123&pc=GAM | 2210968-2017-31358 | 2017-03-06 | Malfunction | ETHICON INC. | 2017-03-27 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A FRENECTOMY PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED AS SIMPLE INTERRUPTED IN THE SUPERFICIAL ORAL MUCOSA LAYER. FOLLOWING THE PROCEDURE, THE SITE WAS WELL HEALED BUT THE SUTURE HAD NOT DISSOLVED AS EXPECTED AND WAS REMOVED IN THE OFFICE AT FOLLOW-UP APPOINTMENT (B)(6) 2017. NO FURTHER TREATMENT WAS NECESSARY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE FRENUM LOCATED UNDER THE TONGUE OR UNDER THE UPPER LIP? WHAT WAS THE DATE OF THE POST OP APPOINTMENT? CAN YOU EXPLAIN ¿SUTURE HAD NOT DISSOLVE AS EXPECTED¿? WHAT WAS THE EXPECTATION OF SUTURE ABSORPTION FROM THE SURGEON? CAN YOU DESCRIBE THE APPEARANCE OF THE SUTURE AT THE POST OP APPOINTMENT? WHAT WAS THE INDICATION FOR REMOVING THE SUTURE? WHAT ARE THE PATIENT GENDER, AGE AND OTHER MEDICAL CONDITIONS? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6436161&pc=GAM | 2210968-2017-31359 | 2017-02-21 | Malfunction | ETHICON INC. | 2017-03-27 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A FRENECTOMY PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED IN THE SUPERFICIAL ORAL MUCOSA IN SIMPLE INTERRUPTED CLOSURE. FOLLOWING THE PROCEDURE, THE PATIENT RETURNED ON (B)(6) 2017 TO HAVE THE SUTURE REMOVED AS THE SUTURE HAD NOT DISSOLVED AS EXPECTED AT FOUR WEEKS POST-OP. THE TISSUE WAS FOUND TO BE WELL HEALED AND THE SUTURE WAS REMOVED IN THE OFFICE. THE PATIENT IS CURRENTLY DOING WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6431673&pc=GAM | 6431673 | 2016-09-06 | Malfunction | ETHICON, INC. | 2017-03-24 | GAM | VICRYL CP-1 | Knotted | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: THE SCRUB ASSISTANT HANDED OFF A 2.0 VICRYL CP-1 TO THE MD UPON SKIN CLOSURE. THE MD NOTICED A KNOT IN THE MIDDLE OF THE SUTURE. A SECOND 2.0 VICRYL CP-1 WAS HANDED OFF TO MD, AND WAS NOTED TO HAVE A KNOT IN IT AS WE WELL. THIRD 2.0 VICRYL CP-1 WAS HANDED OFF TO MD WITH NO FURTHER ISSUES. THERE WAS NO HARM TO PATIENT. SUTURES WERE GIVEN TO THE VENDOR.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429610&pc=GAM | 2210968-2017-31341 | 2017-03-03 | Malfunction | ETHICON INC. | 2017-03-23 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A PEDIATRIC UROLOGY PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED. DURING THE PROCEDURE, WHEN THEY OPENED A PACKAGE, THE SUTURE ITSELF WAS COMPLETELY DISINTEGRATED AND NOT USABLE. ANOTHER LIKE SUTURE WAS OPENED AND USED FOR THIS PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL AND FUNCTIONAL EXAMINATIONS WERE PERFORMED AND TESTED SAMPLES MET REQUIREMENTS. NO DEFECTS WERE FOUND ON THE PACKAGING. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS OR BREAKAGE SUTURE WAS OBSERVED. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6430240&pc=GAM | 2210968-2017-31335 | 2017-02-28 | Malfunction | ETHICON INC. | 2017-03-23 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LUMBAR LAMINECTOMY PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. THE SUTURE BROKE DURING USE AND A NEW ONE WAS OPENED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE. THE DOCTOR OPINED THAT THE SUTURE FEELS THINNER THAN BEFORE. Manufacturer Narrative: THE SAMPLES WERE ALSO MEASURED IN DIAMETER AND THE SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: (B)(4). A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.. REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. THEY WERE VISUALLY AND FUNCTIONALLY EXAMINED FOR TENSILE STRENGTH AND THEY MET THE REQUIREMENTS. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR SUTURE BREAKAGE WAS FOUND.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6399116&pc=GAM | 2210968-2017-31126 | 2017-02-19 | Malfunction | ETHICON INC. | 2017-03-13 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | Device Embedded In Tissue or Plaque | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED TO CLOSE THE UTERUS WITH TWO-LAYER CLOSURE TECHNIQUE. DURING THE PROCEDURE, TOWARDS THE END OF THE CLOSURE FOR THE SECOND LAYER OF THE UTERUS, A 2 MM PART OF THE NEEDLE BROKE INSIDE THE UTERUS. THE SURGEON DECIDED NOT TO TRY TO LOCATE THE LOST PIECE BECAUSE THE PATIENT WAS NOT BLEEDING AND THE BROKEN PIECE OF THE NEEDLE CAUSED NO DANGER TO THE PATIENT. THE SURGEON OPINED THAT THE NEEDLE WENT UNDER BURIED SUTURE WHILE SHE TRIED TO EXIT THE TISSUE. THE BURIED SUTURE HAD PUT A PRESSURE ON THE NEEDLE¿S TIP AND THE NEEDLE¿S TIP GOT BROKEN. THE SURGEON OPINED ALSO THAT THERE SHOULD NOT BE ANY SEVERE CONSEQUENCES. IT WAS ALSO REPORTED THAT THE SURGEON DID NOT HOLD THE NEEDLE¿S TIP GENTLY WHILE SUTURING. THE PATIENT IS WELL AND HAD NO HARM DONE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS ANY ANOMALY OR DEFICIENCY OF THE NEEDLE NOTED PRIOR TO USE: NO; WHAT IS THE SURGEON¿S OPINION OF CONTRIBUTING FACTORS TO NEEDLE BREAKING: THE SURGEON ASSUMES THAT THE NEEDLE WENT UNDER A BURIED SUTURE WHILE SHE TRIED TO EXIT THE TISSUE. THE BURIED SUTURE HAD PUT PRESSURE ON THE NEEDLE¿S TIP AND THE NEEDLE¿S TIP GOT BROKEN; WERE ANY X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE: NO; DOES THE SURGEON PLAN TO REMOVE THE NEEDLE FROM THE PATIENT'S TISSUE: NO; WHAT IS THE SURGEON'S OPINION OF CONSEQUENCES TO THE PATIENT: THE SURGEON THINK¿S THERE SHOULDN¿T BE ANY SEVERE CONSEQUENCES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6397867&pc=GAM | 2210968-2017-31106 | 2017-02-16 | Malfunction | ETHICON INC. | 2017-03-11 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | Device Embedded In Tissue or Plaque | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ESOPHAGUS SURGICAL PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED. DURING THE PROCEDURE, WHEN SUTURING THE GASTER WITH THE NEEDLE-SUTURE, IT WAS NOTICED THE TIP OF THE NEEDLE HAD BEEN MISSING. THE NEEDLE HAD BEEN HELD AT ITS TIP DURING USE. THE BROKEN TIP WAS SEARCHED FOR BY X-RAY PHOTOS DURING THE PROCEDURE, HOWEVER, IT WAS NOT FOUND. THERE IS A POSSIBILITY IT REMAINS IN THE BODY. AS OF (B)(6) 2017, THE PATIENT WAS IN THE HOSPITAL, AND WAS GETTING BETTER. NO ADVERSE CONSEQUENCES CAUSED BY THE BROKEN PIECE HAD BEEN OBSERVED UNTIL THEN. Manufacturer Narrative: (B)(4). THE ACTUAL SAMPLE WAS RETURNED FOR ANALYSIS. A FRACTURE WAS OBSERVED AT THE TIP OF THE NEEDLE. THE NEEDLE WAS RECEIVED IN ONE PIECE, THE MATING FRACTURE SURFACE WAS NOT PROVIDED FOR THIS EVALUATION. THE FRACTURE SURFACES AND SURROUNDING AREA OF THE NEEDLE WERE EXAMINED TO DETERMINE FRACTURE MODE. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE TIP OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE THE PACKAGE INSERT (IFU) CAUTIONS: GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH# KMM047, EXP. DATE: 10/31/2021, MFR. DATE: 11/01/2016. BATCH# JK5435, EXP. DATE: 08/31/2020, MFR. DATE: 09/29/2015. BATCH# KM6738, EXP. DATE: 10/31/2021, MFR. DATE: 11/24/2016. BATCH# JLM385, EXP. DATE: 09/30/2020, MFR. DATE: 10/03/2015. BATCH# KMM036, EXP. DATE: 10/31/2021, MFR. DATE: 11/01/2016. BATCH# KMM073, EXP. DATE: 10/31/2021, MFR. DATE: 11/01/2016. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE ADVISE OF ANY MEDICAL / SURGICAL INTERVENTION AND STATUS OF PATIENT. ¿THE BROKEN PIECE WAS SEARCHED FOR BY X-RAYS DURING THE INDEX OPERATION, HOWEVER, IT WAS NOT FOUND. AS OF (B)(6), THE PATIENT WAS IN THE HOSPITAL, AND WAS GETTING BETTER. NO ADVERSE CONSEQUENCES CAUSED BY THE BROKEN PIECE HAD BEEN OBSERVED UNTIL THEN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6394127&pc=GAM | 2210968-2017-31088 |  | Malfunction | ETHICON INC. | 2017-03-09 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER RECEIVED A BOX OF SUTURE AND THERE WAS A DIFFERENT ITEM INSIDE THE BOX THAN WHAT WAS INDICATED ON THE BOX LABEL. NO PATIENT INVOLVEMENT WAS REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE DESCRIBE WHAT WAS THE PRODUCT CODE ORDERED? IS THE ISSUE RELATED TO THE SHIPMENT BOX? WAS THE INCORRECT PRODUCT INSIDE OF THE SHIPMENT BOX? WHAT WAS THE PRODUCT CODE ON THE BOX OF SUTURE? WHAT WAS THE PRODUCT CODE OF THE FOILS INSIDE THE SUTURE BOX ? WERE THE 2 LOTS KKM178 AND KK5111 IN THE ONE SHIPPING BOX OR IN SEPARATE SHIPPING BOX? ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. TWO POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH KKM178 EXPIRATION DATE: 08/31/2021. DATE OF MFG.: 09/02/2016. BATCH KK5111 EXPIRATION DATE: 08/31/2021. DATE OF MFG.: 09/29/2016. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6394128&pc=GAM | 2210968-2017-31089 |  | Malfunction | ETHICON INC. | 2017-03-09 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER RECEIVED A BOX OF SUTURE AND THERE WAS A DIFFERENT ITEM INSIDE THE BOX THAN WHAT WAS INDICATED ON THE BOX LABEL. NO PATIENT INVOLVEMENT WAS REPORTED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 03/09/2017. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE DESCRIBE WHAT WAS THE PRODUCT CODE ORDERED? IS THE ISSUE RELATED TO THE SHIPMENT BOX? WAS THE INCORRECT PRODUCT INSIDE OF THE SHIPMENT BOX? WHAT WAS THE PRODUCT CODE ON THE BOX OF SUTURE? WHAT WAS THE PRODUCT CODE OF THE FOILS INSIDE THE SUTURE BOX ? WERE THE 2 LOTS KKM178 AND KK5111 IN THE ONE SHIPPING BOX OR IN SEPARATE SHIPPING BOX? THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. TWO POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH KKM178 EXPIRATION DATE: 08/31/2021. DATE OF MFG.: 09/02/2016. BATCH KK5111 EXPIRATION DATE: 08/31/2021. DATE OF MFG.: 09/29/2016. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6348811&pc=GAM | 6348811 | 2017-01-27 | Malfunction | ETHICON, INC. | 2017-02-22 | GAM | VICRYL | Break; Detachment Of Device Component | Foreign Body In Patient |  |  | 1 | Event Description: TIP OF NEEDLE OF A O VICRYL UR-6 SUTURE BROKE OFF AFTER ONE PASS THROUGH ABDOMEN. BOTH PIECES WERE RECOVERED AND ACCOUNTED FOR. AN X-RAY WAS OBTAINED AT END OF THE PROCEDURE, THE RESULTS WERE NEGATIVE PER RADIOLOGY. THERE WAS NO HARM TO PATIENT. MANUFACTURER RESPONSE FOR NEEDLE, O UR-6 (PER SITE REPORTER): HOSPITAL HAS REQUESTED TO MEET WITH THE ETHICON SALES REP TO DISCUSS THE MULTIPLE NEEDLES BREAKING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6307213&pc=GAM | 6307213 | 2016-12-15 | Malfunction | ETHICON, INC. | 2017-02-07 | GAM | VICRYL | Break; Detachment Of Device Component | Foreign Body In Patient |  |  | 1 | Event Description: THE NEEDLE BROKE WHILE PLACING CLOSING SUTURES IN. THE NEEDLE WAS A CUTTING NEEDLE CP-1 0 VICRYL WHICH IS AN ABSORBABLE SUTURE. THE NEEDLE WAS DISCARDED IN A SHARPS CONTAINER. BOTH PARTS OF NEEDLE APPEARED TO BE RECOVERED. X-RAY OF KNEE TAKEN TO CONFIRM. THERE WAS NO HARM TO THE PATIENT.Event Description: THE NEEDLE BROKE WHILE PLACING CLOSING SUTURES IN. THE NEEDLE WAS A CUTTING NEEDLE CP-1 0 VICRYL WHICH IS AN ABSORBABLE SUTURE. THE NEEDLE WAS DISCARDED IN A SHARPS CONTAINER. BOTH PARTS OF NEEDLE APPEARED TO BE RECOVERED. X-RAY OF KNEE TAKEN TO CONFIRM. THERE WAS NO HARM TO THE PATIENT. Manufacturer Narrative: UNIQUE DEVICE IDENTIFIER (UDI): FOR TYPE OF DEVICE: SUTURE, ABSORBABLE, SYNTHETIC, POLYGLYCOLIC ACID.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6309364&pc=GAM | 2210968-2017-30669 | 2017-01-01 | Malfunction | ETHICON INC. | 2017-02-07 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN ANIMAL PROCEDURE AND THE SUTURE WAS USED. AFTER THE PROCEDURE, A DOG HAD RUPTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: REPRESENTATIVE SAMPLES WERE RECEIVED FOR EVALUATION. VISUAL AND FUNCTIONAL EXAMINATIONS WERE PERFORMED AND SAMPLES MET THE SPECIFIED QUALITY REQUIREMENTS. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 02/07/2017. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THE SUTURE BREAK POST OPERATIVELY? ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE SUTURE BREAK POST OPERATIVELY? THREE DOGS WERE OPERATED WITH THIS SUTURE AND HAD TO BE REOPERATED. WHAT IS THE PRODUCT CODE AND LOT NUMBER? VICRYL VCP587H LOT: HM8DPJP0. IS ANY PRODUCT COMING BACK FOR EVALUATION?- NO MATERIAL LEFT DIRECTLY FROM THE OPERATED DOGS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6309367&pc=GAM | 2210968-2017-30670 | 2017-01-01 | Malfunction | ETHICON INC. | 2017-02-07 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN ANIMAL PROCEDURE AND THE SUTURE WAS USED. AFTER THE PROCEDURE, A DOG HAD RUPTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THE SUTURE BREAK POST OPERATIVELY? ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE SUTURE BREAK POST OPERATIVELY?- 3 DOGS WERE OPERATED WITH THIS SUTURE AND HAD TO BE REOPERATED. WHAT IS THE PRODUCT CODE AND LOT NUMBER? - VICRYL VCP587H, LOT: HM8DPJP0. IS ANY PRODUCT COMING BACK FOR EVALUATION?- NO MATERIAL LEFT DIRECTLY FROM THE OPERATED DOGS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6309369&pc=GAM | 2210968-2017-30671 | 2017-01-01 | Malfunction | ETHICON INC. | 2017-02-07 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN UNKNOWN ANIMAL PROCEDURE AND THE SUTURE WAS USED. AFTER THE PROCEDURE, A DOG HAD RUPTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THE SUTURE BREAK POST OPERATIVELY? ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE SUTURE BREAK POST OPERATIVELY?- THREE DOGS WERE OPERATED WITH THIS SUTURE AND HAD TO BE REOPERATED. WHAT IS THE PRODUCT CODE AND LOT NUMBER? - VICRYL VCP587H LOT: HM8DPJP0. IS ANY PRODUCT COMING BACK FOR EVALUATION?- NO MATERIAL LEFT DIRECTLY FROM THE OPERATED DOGS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6296465&pc=GAM | 2210968-2017-30556 | 2016-08-18 | Malfunction | ETHICON INC. | 2017-02-02 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CESAREAN SECTION ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, THE POINT OF THE NEEDLE BROKE WHEN BUSY SUTURING THE UTERUS. THERE WERE NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID A PIECE OF THE NEEDLE FALL INTO THE PATIENT? IF YES, WAS THE NEEDLE PIECE REMOVED, AND HOW? WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? WAS THE PATIENT BROUGHT BACK TO OR TO HAVE NEEDLE REMOVED OR WAS THE NEEDLE REMOVED DURING THE INITIAL PROCEDURE? WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6297299&pc=GAM | 2210968-2017-30567 |  | Malfunction | ETHICON INC. | 2017-02-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A PHACO/GONIOSYNECHIALYSIS/BLEB REVISION WITH MITOMYCIN-C /MMC PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS PLACED CONTINUOUS ON CONJUNCTIVA. AFTER THE PROCEDURE, DURING THE ROUTINE FOLLOW-UP APPOINTMENT, IT WAS DISCOVERED THAT THE SUTURE HAD DEHISCED/ IT WAS BROKEN. THE RE-SUTURE WAS NOT REQUIRED AND NO OTHER SKIN CLOSURE WAS USED WITH THE SUTURE. IT WAS ALSO REPORTED THAT EXTRA POSTOPERATIVE VISITS WERE REQUIRED AND THE PATIENT HAS DONE WELL. THE SURGEON OPINED THAT IT IS A SUTURE STRENGTH PROBLEM. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS MEDWATCH CAPTURES EVENT RELATED TO CASE 1. MEDWATCH FOR ADDITIONAL PATIENT EVENTS WILL BE SUBMITTED SEPARATELY. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CAN YOU PROVIDE THE BELOW INFORMATION FOR EACH PATIENT EVENTS? WHAT DATE WAS EACH INITIAL PROCEDURE? (B)(6) 2016. WHAT WAS THE NAME OF EACH PROCEDURE? PHACO/GONIOSYNECHIALYSIS/BLEB REVISION WITH MITOMYCIN-C (MMC). WHAT TISSUE WAS VICRYL SUTURE USED ON? CONJUNCTIVA. WHAT WAS THE CONDITION OF THE TISSUE (HEALTHY, DISEASED)? PREVIOUSLY OPERATED TISSUE. HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? CONTINUOUS. WAS THERE ANY PATIENT EVENT PRIOR TO SYMPTOMS (COUGHING, FALLING, MOVING)? NO EVENT THAT WAS REPORTED TO THE PHYSICIAN. WHAT WAS USED TO RE-SUTURE THE INCISION? RE-SUTURE NOT REQUIRED. WHAT IS THE SURGEON OPINION AS TO THE POTENTIAL CONTRIBUTING FACTORS OF THE WOUND OPENING AND SUTURE BREAKAGE POST OP? ¿BECAUSE I HAVE USED THIS SUTURE MANY TIMES IN THE PAST FOR THE SAME OPERATION, I FEEL IT IS A SUTURE STRENGTH PROBLEM.¿ ¿ PER MD. WHAT IS THE PATIENT AGE, WEIGHT, GENDER AND OTHER MEDICAL CONDITIONS? (B)(6) YEAR OLD MALE WITH H/O GLAUCOMA AND CATARACTS WITH A BMI OF (B)(6). WHAT IS THE CURRENT CONDITION OF THE PATIENT? IN EACH CASE THE PATIENT HAS DONE WELL. IN ALL 3 CASES MANY EXTRA POSTOPERATIVE VISITS WERE REQUIRED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6289941&pc=GAM | 2210968-2017-30465 | 2017-01-16 | Malfunction | ETHICON INC. | 2017-01-31 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | Device Embedded In Tissue or Plaque | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED. DURING SUTURING THE UTERUS, THE TIP OF THE NEEDLE BROKE OFF AND IT WAS NOT FOUND. IT WAS ALSO REPORTED THAT X-RAY WAS NOT TAKEN. THERE WAS NO ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING. THE DOCTOR OPINES THAT THE TIP IS TOO SMALL TO BE A CONCERN. Manufacturer Narrative: THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: UNOPENED/STERILE SAMPLES WERE RECEIVED FOR EVALUATION. NEEDLE SAMPLES HAVE BEEN SUBJECTED TO A VISUAL INSPECTION, NEEDLE DIAMETER TEST AND DUCTILITY TEST, AND ALL RESULTS MET THE SPECIFIED QUALITY REQUIREMENTS. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID A PIECE OF THE NEEDLE FALL INTO THE PATIENT? IF YES, WAS THE NEEDLE PIECE REMOVED, AND HOW? TIP OF NEEDLE BROKE OFF BUT WAS FOUND AND RETRIEVED. FOR THE SECOND PC, TIP BROKE OFF BUT WAS NOT FOUND. DOCTOR SAID THAT TIP IS TOO SMALL TO BE A CONCERN WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? NO XRAYS TAKEN. WAS THE PATIENT BROUGHT BACK TO OR TO HAVE NEEDLE REMOVED OR WAS THE NEEDLE REMOVED DURING THE INITIAL PROCEDURE? SECOND PC OF BROKEN NEEDLE WAS NOT FOUND. WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? NO ADDITIONAL TISSUE DAMAGE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6290186&pc=GAM | 2210968-2017-30469 | 2016-11-22 | Malfunction | ETHICON INC. | 2017-01-31 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | Device Embedded In Tissue or Plaque | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE WHEN THE DOCTOR WAS CLOSING THE WOUND IT WAS NOTED THE TIP OF THE NEEDLE WAS MISSING AND WAS NOT REMOVED. THE NEEDLE WAS RETAINED IN THE UTERINE TISSUE AT SUTURE LINE. THE PATIENT'S CURRENT CONDITION IS STABLE WITH NO PLANS TO READMIT PATIENT FOR REMOVAL OF NEEDLE. Manufacturer Narrative: (B)(4). THIS MEDWATCH REPORT IS IN RESPONSE TO RECEIPT OF MEDSUN MANDATORY AND VOLUNTARY REPORT FORM 0504440000-2016-8013. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID A PIECE OF THE NEEDLE FALL INTO THE PATIENT? IF YES, WAS THE NEEDLE PIECE REMOVED, AND HOW? YES. NEEDLE TIP WAS NOT REMOVED WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? YES. WAS THE PATIENT BROUGHT BACK TO OR TO HAVE NEEDLE REMOVED OR WAS THE NEEDLE REMOVED DURING THE INITIAL PROCEDURE? NEEDLE WAS NOT REMOVED DUE TO HIGHER RISK OF ATTEMPTING TO REMOVE IT VERSUS LEAVING NEEDLE IN CURRENT LOCATION. WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? NO. WHAT IS THE NAME OF THE PROCEDURE? CESAREAN SECTION. WHAT LOCATION AND WHAT TISSUE IS THE NEEDLE RETAINED IN? UTERINE TISSUE AT SUTURE LINE. UPDATE TO PATIENT CURRENT CONDITION - STABLE ON DISCHARGE WITH NO PLANS TO READMIT PATIENT FOR REMOVAL. ARE THERE ANY PLANS TO REMOVE THE PIECE OF NEEDLE? NOT THAT I AM AWARE OF AT THIS TIME. WE HAVE SENT YOU A PRODUCT RETURN MAILER. PLEASE PROVIDE PRODUCT RETURN DATE AND TRACKING NUMBER WHEN AVAILABLE - PRODUCT WILL NOT BE RETURNED, CAN SEND ANOTHER NEEDLE FROM THE SAME LOT # IF YOU WOULD LIKE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6287627&pc=GAM | 2210968-2017-30452 | 2017-01-18 | Malfunction | ETHICON INC. | 2017-01-30 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient; Device Embedded In Tissue or Plaque | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC GASTRECTOMY PROCEDURE ON (B)(6) 2017 AND THE SUTURE WAS USED. DURING THE PROCEDURE, WHEN CLOSING A SURGICAL WOUND AFTER REMOVING A PORT, IT WAS FOUND THAT THE TIP OF THE NEEDLE HAD BEEN MISSING. X-RAYS WERE TAKEN, HOWEVER, THE MISSING TIP WAS NOT FOUND. IT IS UNKNOWN WHETHER THE BROKEN TIP REMAINED IN THE PATIENT¿S TISSUE OR THERE HAD BEEN NO TIP ORIGINALLY. THE PATIENT¿S CURRENT STATUS IS GOOD. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). BASED ON EVALUATION, THE NEEDLE WAS NOT BROKEN SO THE NEEDLE TIP COULD NOT HAVE BEEN RETAINED IN THE PATIENT. THIS EVENT IS THEREFORE MALFUNCTION REPORTABLE. THE 1ST EVALUATION REVEALED MECHANICAL DAMAGE ON THE FRACTURE SURFACE, BUT NO CONCLUSIVE FRACTURE EVIDENCE. FURTHER EVALUATION WAS PERFORMED AND IT WAS OBSERVED THAT THE TIP ON THE NEEDLE IS INCOMPLETE (DEFORMED). Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE NEEDLE WAS RETURNED AND IT SEEMS NOT TO BE SHARPENED. NAME OF INDEX SURGICAL PROCEDURE: LAP OR ROBOTIC GASTRECTOMY; DOES THE PIECE/DEVICE REMAIN RETAINED IN THE PATIENT'S TISSUE: IT IS UNKNOWN WHETHER THE BROKEN TIP REMAINED IN THE PATIENT¿S TISSUE OR THERE HAD BEEN NO TIP ORIGINALLY (MEANING, THE TIP HAD NOT BEEN POLISHED.); IF APPLICABLE, WILL PRODUCT BE RETURNED, RETURN DATE, TRACKING INFORMATION: WE WILL RETURNED IT TO YOU. WE WILL INFORM RETURN DATE AND TRACKING INFORMATION WHEN AVAILABLE; WHAT IS THE PATIENT CURRENT STATUS: GOOD.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6263480&pc=GAM | 6263480 |  | Malfunction | ETHICON, INC. | 2017-01-19 | GAM | VICRYL | Bent; Device Damaged Prior to Use | No Patient Involvement |  |  | 1 | Event Description: THE TIP OF SUTURE NEEDLE IS BENT. ALL BOXES OF SAME LOT # PULLED . PULLED RANDOM SAMPLES, AND ALL HAD BENT TIP. MANUFACTURER RESPONSE FOR 0 VICRYL SUTURE, VICRYL (PER SITE REPORTER): TOOK DESCRIPTION OF PROBLEM; ISSUED RGA#. THEY WILL SEND RETURN KIT, AND REPLACEMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6265307&pc=GAM | 2210968-2017-30299 | 2017-01-06 | Malfunction | ETHICON INC. | 2017-01-19 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT GYNECOLOGICAL PROCEDURE ON (B)(6) 2017 AND SUTURE WAS USED. WHEN THE PACKAGE WAS OPENED, IT WAS FOUND THAT THERE WAS A NEEDLE WITH BV-1 SIZE IN THE PACKAGE. NO ADDITIONAL INFORMATION IS AVAILABLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: IT WAS RECEIVED ONLY ONE LOOSE NEEDLE. THE LOOSE NEEDLE WAS EXAMINED FOR VISUAL INSPECTION AND IT WAS OBSERVED THAT THE SHAPE OF THE SWAGE AREA OF THE NEEDLE IS TOTALLY CIRCULAR. THE NEEDLE FOR PRODUCT CODE J739D IS CT-1 RAW WIRE DIAMETER 40 MIL AND THE NEEDLE RECEIVED COULD NOT BE ASSEMBLED IN A 2-0 DIAMETER SUTURE. DUE TO THAT THE TRAY AND OTHER NEEDLE/SUTURE WERE NOT RETURNED FOR EVALUATION CANNOT BE DETERMINATE WHETHER IT WAS FOUND THAT THERE HAD BEEN A NEEDLE WITH BV-1 SIZE IN THE PACKAGE WHEN THE PACKAGE WAS OPENED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6240630&pc=GAM | 6240630 | 2016-11-11 | Malfunction | ETHICON, INC. | 2017-01-11 | GAM | VICRYL | Break; Detachment Of Device Component; Entrapment of Device | Foreign Body In Patient |  |  | 1 | Event Description: DURING START OF SURGERY THE NEEDLE OF A 0-VICRYL UR-6 BROKE IN THE TISSUE OF THE PATIENT'S UMBILICUS. BOTH PIECES WERE RETRIEVED RIGHT AWAY. AN X-RAY WAS PERFORMED TO VERIFY NO OTHER PIECES REMAINED. NOTHING WAS NOTED ON X-RAY UPON REVIEW FROM MD AND RADIOLOGIST.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227662&pc=GAM | 2210968-2017-30104 | 2016-12-14 | Malfunction | ETHICON INC. | 2017-01-05 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. THE SUTURE WAS FOUND STILL HARD AFTER USED ON THE PATIENT OVER TEN DAYS. THE ABSORPTION IS NOT AS GOOD AS BEFORE. THE SUTURE WAS SLOW TO ABSORB IN THE MUCOSA LAYER. THE NURSE OPINED THAT THE SUTURE IS SOFTER THAN BEFORE WHICH MAY LEAD TO THE WOUND NOT HEALING WELL. IN ADDITION, IT WAS REPORTED THAT THERE IS SOME BLANK SECTIONS ON THE PRODUCT IFU. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE ISSUE OCCUR WITH 1 OR 2 SUTURES? TWO. ARE THE 2 SAMPLES BEING RETURNED UNOPENED REPRESENTATIVE SAMPLES? YES. DID THE PATIENTS WOUND NOT HEAL WELL? YES. WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED? UNKNOWN. UPDATE TO PATIENT CURRENT CONDITION- DISCHARGED. Manufacturer Narrative: REPRESENTATIVE SAMPLES RECEIVED FOR EVALUATION. VISUAL INSPECTION WAS PERFORMED AND NO DEVIATIONS COULD BE DETECTED. FUNCTIONAL INSPECTIONS WERE PERFORMED AND THE SAMPLES MET THE REQUIREMENTS. REMAINING TENSILE STRENGTH CORRESPONDED TO THE ABSORPTION PROFILE OF THIS SUTURE MATERIAL. PICTURES WERE PROVIDED AND REVIEWED AND SHOWED THE PRINTING ON THE BOX LOOKS MARRED AND SMEARED. DATAMATRIX INFORMATION IS NOT READABLE AND DETAILS OF THE PRODUCT INFORMATION APPEAR PARTLY PALE AND SMEARED. ALSO THE BARCODE PRINTING APPEARS NOT CLEAR BUT DEFACED. PICTURES OF THE IFU WERE REVIEWED AND COMPARED WITH REQUIRED IFU AS PER PRODUCT SPECIFICATION LIST: NO DEVIATIONS COULD BE DETECTED. IFU IS INTENDED TO BE ONLY IN (B)(6) LANGUAGE AND OTHER PANELS WERE LEFT INTENTIONALLY BLANK. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PROVIDE PROCEDURE NAME: C-SECTION; IN WHAT TISSUE SUTURE USED AND SLOW TO ABSORB: IN MUCOSA LAYER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6210681&pc=GAM | 2210968-2016-15651 | 2016-12-06 | Malfunction | ETHICON INC. | 2016-12-28 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT A CANINE CATARACT EXTRACTION VETERINARY PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED INTERRUPTED IN THE PERIPHERAL CORNEA. THERE WERE NO QUALITY ISSUES NOTICED DURING THE PROCEDURE. ABOUT TEN HOURS LATER AFTER THE PROCEDURE, THE SUTURE BROKE, BUT A KNOT WAS INTACT. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: ACCORDING TO THE SAMPLE CONDITION NO SUTURE BREAKAGE WAS OBSERVED AND THE SAMPLES MET THE REQUIREMENTS. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH#: KBK011, EXP DATE: 01/31/2021, MFR DATE: 02/19/2016. BATCH#: KC6408, EXP DATE: 02/28/2021, MFR DATE: 03/17/2016. BATCH#: HE5145, EXP DATE: 01/31/2019, MFR DATE: 05/30/2014. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6210946&pc=GAM | 2210968-2016-15655 | 2016-12-07 | Malfunction | ETHICON INC. | 2016-12-28 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT A CANINE CATARACT EXTRACTION VETERINARY PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED INTERRUPTED IN THE PERIPHERAL CORNEA. THERE WERE NO QUALITY ISSUES NOTICED DURING THE PROCEDURE. ABOUT TEN HOURS LATER AFTER THE PROCEDURE, THE SUTURE BROKE, BUT A KNOT WAS INTACT. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 12/28/2016. ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH# (B)(4) EXP DATE: 01/31/2021, MFR DATE: 02/19/2016. BATCH# (B)(4) EXP DATE: 02/28/2021, MFR DATE: 03/17/2016 BATCH# (B)(4) EXP DATE 01/31/2019, MFR DATE: 05/30/2014 IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6207450&pc=GAM | 2210968-2016-15642 | 2016-12-06 | Malfunction | ETHICON INC. | 2016-12-27 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Unsealed Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. DURING THE PROCEDURE, IT WAS FOUND THAT THE SUTURE HAD BEEN ALREADY HYDROLYZED WHEN THE PACKAGE WAS OPENED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: SAMPLE AND OPENED FOIL WERE RETURNED FOR EVALUATION. THE SAMPLE WAS VISUALLY INSPECTED AND SUTURE HAS BEGUN PROCESS OF DISINTEGRATION. MULTIPLES HOLES IN CAVITY WERE FOUND ON THE PACKAGE AND THIS CONDITION CONTRIBUTED TO DISINTEGRATION OF THE SUTURE. THE CONDITIONS OF THE HOLES COULD NOT BE DETERMINED. UNOPENED SAMPLES WERE RETURNED FOR EVALUATION. THE SAMPLES WERE VISUALLY INSPECTED AND FUNCTIONALLY TESTED AND SAMPLES MET THE REQUIREMENTS. NO PINHOLES WERE OBSERVED ON THE FOILS. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6189021&pc=GAM | 6189021 | 2016-09-16 | Malfunction | ETHICON, INC. | 2016-12-20 | GAM | VICRYL SUTURE | Microbial Contamination of Device | Bacterial Infection; Post Operative Wound Infection |  |  | 1 | Event Description: VICRYL SUTURES CULTURED AFTER (3) SUPERFICIAL SURGICAL SITE INFECTIONS (SSIS) IDENTIFIED A FEW WEEKS AFTER TOTAL JOINT CASES DONE ALL ON THE SAME DAY AND IN THE SAME OPERATING ROOM. SUTURES USED DURING 3 TOTAL JOINT PROCEDURES WERE KEPT IN OPERATING ROOM, AND REMOVED AFTER THREE (3) PATIENT'S SUPERFICIAL SSIS WERE IDENTIFIED. ALL THREE (3) TYPES OF SUTURES WERE SUBMITTED FOR CULTURING (X3 EACH- 9 TOTAL) SEVEN CAME BACK WITH NO GROWTH. TWO CAME BACK WITH POSITIVE (+) CULTURE RESULTS THAT MAY BE ATTRIBUTED TO CONTAMINATION DURING TESTING AND ONE HAD COAGULASE NEGATIVE (B)(6) SPECIES AND THE OTHER HAD COAGULASE (B)(6)S. LOT #S - JMK776 AND JL7734.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6169062&pc=GAM | 6169062 | 2016-10-28 | Malfunction | ETHICON, INC. | 2016-12-13 | GAM | VICRYL | Break; Detachment Of Device Component | Foreign Body In Patient |  |  | 1 | Event Description: NEEDLE BROKE OFF DURING A C-SECTION. UNABLE TO LOCATE TIP OF NEEDLE. X-RAY CLEAR PRIOR TO BEING TRANSFERRED OUT OF OR TO PACU.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6156275&pc=GAM | 2210968-2016-15371 | 2016-08-15 | Malfunction | ETHICON INC. | 2016-12-08 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT LOWER CAESAREAN SECTION ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE DURING SUTURING. THIS OCCURRED AFTER 4 TO 5 PASSES THROUGH THE UTERUS. THE PROCEDURE WAS COMPLETED WITH LIKE DEVICE WITH 10 MINUTES DELAY. ALL THE NEEDLE PIECES WERE RETRIEVED FROM THE PATIENT AND NO ADDITIONAL DISSECTION WAS REQUIRED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6124338&pc=GAM | 2210968-2016-15090 | 2016-10-31 | Malfunction | ETHICON INC. | 2016-11-23 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A HERNIA REPAIR PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. WHEN THE PACKAGE WAS OPENED, A DIFFERENT TYPE OF THE SUTURE MONOFILAMENT WAS IN THE PACKAGE INSTEAD. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE ACTUAL DEVICE WAS RETURNED FOR EVALUATION. IT WAS DETERMINED THAT A PRODUCT MIX OCCURRED. ALSO, REPRESENTATIVE SAMPLES WERE RECEIVED. THEY WERE VISUALLY EXAMINED, NO DEFECTS WERE FOUND AND THE SUTURES BELONG TO THE PRODUCT FAMILY VICRYL. NO PRODUCT MIX WAS FOUND ON THE REPRESENTATIVE SAMPLES. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6124651&pc=GAM | 6124651 | 2016-08-09 | Malfunction | ETHICON, INC. | 2016-11-23 | GAM | VICRYL SUTURE | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: DURING CLOSURE OF WOUND 3-0 VICRYL SUTURE NEEDLE TIP BROKE (LESS THAN 1MM).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6124652&pc=GAM | 6124652 | 2016-08-30 | Malfunction | ETHICON, INC. | 2016-11-23 | GAM | VICRYL SUTURE | Break; Detachment Of Device Component | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: DURING PROCEDURE THE TIP OF THE SUTURE NEEDLE, 0 VICRYL ON CT-1 NEEDLE TIP BROKE. THE BROKEN NEEDLE AND TIP WERE IMMEDIATELY LOCATED AND VERIFIED BOTH PIECES ACCOUNTED FOR.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6119996&pc=GAM | 2210968-2016-15063 |  | Malfunction | ETHICON INC. | 2016-11-22 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY THE DISTRIBUTOR THAT THE PRODUCT WAS MISLABELED WITH PRODUCT CODE J214H. THE LOT NUMBER BELONGS TO PRODUCT CODE J214W. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. THE PHOTO OF PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. Manufacturer Narrative: A PHOTOGRAPHIC EVALUATION WAS PERFORMED. THIS IS AN INTERNATIONAL CODE. THE BOX SHOULD SHOW (B)(4). THEREFORE THE TB GRAPHIC IS THE CORRECT ONE FOR THIS BATCH. Manufacturer Narrative: THE SEALED BOXES WERE EXAMINED FOR VISUAL INSPECTION AND NO DEFECTS WERE OBSERVED ON THE BOXES. THE INFORMATION OF THE PRODUCT CODE, LOT NUMBER AND GRAPHIC ON THE LABEL BOXES PRINTED WAS COMPARED AND THE INFORMATION WAS CORRECT. THIS IS AN INTERNATIONAL CODE AND THE BOX SHOULD SHOW J214H. ACCORDING THE EVALUATION, NO DEFECTS WERE FOUND AND THE INFORMATION ON THE BOXES IS CORRECT ACCORDING TO OUR SYSTEM. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6121983&pc=GAM | 2210968-2016-15077 |  | Malfunction | ETHICON INC. | 2016-11-22 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CAESARIAN SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING AN UNKNOWN OB-GYN PROCEDURE AT A LATER DATE AT ANOTHER HOSPITAL, IT WAS FOUND THAT THE SUTURE FROM THE PREVIOUS CAESARIAN SECTION HAD BEEN LEFT IN THE SURGICAL INCISION SITE. THE DOCTOR OPINED THAT THE SUTURE MIGHT POSSIBLY BE EITHER AN UNKNOWN SUTURE PRODUCT OR ANOTHER COMPETITIVE PRODUCT (SURGISORB) AS IT WAS AN ABSORBABLE SUTURE TYPE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: DATE SENT TO THE FDA: 1/4/2017 ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE DATE OF THE PREVIOUS C-SECTION?...NO INFORMATION. WHAT WAS THE DATE OF THE OB-GYN PROCEDURE?...NO INFORMATION. WHAT TISSUE WAS THE UNKNOWN SUTURE FOUND IN? ¿THE UTERUS AND THE URINARY BLADDER. HOW MANY DAYS POST OP FROM THE PREVIOUS C-SECTION? ¿NO INFORMATION. WAS THE UNKNOWN SUTURE REMOVED FROM THE PATIENT?...YES. HOW WAS THE UNKNOWN SUTURE REMOVED FROM THE PATIENT?...NO INFORMATION. DO YOU HAVE THE SUTURE PIECE FOR EVALUATION?...YES. Manufacturer Narrative: THE RETURNED EXPLANTED SUTURE WAS DETERMINED NOT TO BE AN ETHICON SUTURE PRODUCT. THE RETURNED EXPLANTED SUTURE DID NOT CORRESPOND TO ETHICON COATED VICRYL, BASED ON CONSTRUCTION. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 11/22/2016. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6109276&pc=GAM | 2210968-2016-15029 | 2016-09-24 | Malfunction | ETHICON INC. | 2016-11-17 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE SURGICAL THREAD INSIDE THE PACKAGE WAS NOT SAME AS THE LABEL ON THE PACKAGE. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS RECEIVED ONE OPEN FOIL PACKAGE AND PAPER LID FOR PRODUCT VCP353H (1 VICRYL SUTURE 70 CM CT). IT WAS ALSO RECEIVED ONE SUTURE THAT LOOKS LIKE PRODUCT Y353H (1 MONOCRYL VIOLET SUTURE 70 CM CT). IT WAS IDENTIFIED THAT THE NEEDLE WAS MANUALLY POSITIONED IN THE PLASTIC TRAY SINCE IT WAS PARKED WITH THE TIP POINTED TO THE OUTER SIDE. PLEASE PROVIDE THE INFORMATION BELOW: THE CUSTOMER INFORMED THAT A PROLENE SUTURE WAS IN THE PACKAGE, COULD YOU PLEASE CONFIRM WITH THE CUSTOMER IF THE SAMPLE RETURNED IS FOR THIS COMPLAINT? CAN YOU CONFIRM THE LOT NUMBER, SINCE THAT INFORMATION WAS REMOVED FROM THE FOIL PACKAGE? ACCORDING TO THE CHARACTERISTICS OF THE SAMPLE, IT WAS NOT POSSIBLE TO ANALYZE BECAUSE DOESN'T MATCH THE PRODUCT REPORTED BY CUSTOMER. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE CUSTOMER INFORMED THAT A PROLENE SUTURE WAS IN THE PACKAGE, COULD YOU PLEASE CONFIRM WITH THE CUSTOMER IF THE SAMPLE RETURNED IS FOR THIS COMPLAINT? YES, THE SENT SAMPLE IS RELATED TO THIS COMPLAINT. CAN YOU CONFIRM THE LOT NUMBER, SINCE THAT INFORMATION WAS REMOVED FROM THE FOIL PACKAGE? THE CUSTOMER COULD NOT INFORM THE CORRECT LOT NUMBER, ONCE THE PROFESSIONAL THAT USED THE MATERIAL (BY MISTAKE) AND DISCHARGED THE PACKAGE. THE SALES REP REQUESTED THE PRODUCT LOT THAT THE CUSTOMER HAD IN HIS WAREHOUSE BUT THE SALES REP. DID NOT RECEIVE ANY ANSWER. Manufacturer Narrative: (B)(4) DATE SENT TO THE FDA: 11/17/2016. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. CAN YOU IDENTIFY THE LOT NUMBER?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6109314&pc=GAM | 2210968-2016-15030 | 2016-10-31 | Malfunction | ETHICON INC. | 2016-11-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A HERNIA REPAIR PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. WHEN THE PACKAGE WAS OPENED, A DIFFERENT TYPE OF THE SUTURE MONOFILAMENT WAS IN THE PACKAGE INSTEAD. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE ACTUAL DEVICE WAS RETURNED FOR EVALUATION. IT WAS DETERMINED THAT A PRODUCT MIX OCCURRED. ALSO, REPRESENTATIVE SAMPLES WERE RECEIVED. THEY WERE VISUALLY EXAMINED, NO DEFECTS WERE FOUND AND THE SUTURES BELONG TO THE PRODUCT FAMILY VICRYL. NO PRODUCT MIX WAS FOUND ON THE REPRESENTATIVE SAMPLES. Manufacturer Narrative: (B)(4) DATE SENT TO THE FDA: 11/17/2016. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6105549&pc=GAM | 2210968-2016-14976 | 2016-10-25 | Malfunction | ETHICON INC. | 2016-11-16 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016. PRIOR TO THE PROCEDURE, WHEN THE FOIL WAS OPENED THEY FOUND THE SUTURE HAD TURNED INTO POWDER. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6103648&pc=GAM | 2210968-2016-14955 |  | Malfunction | ETHICON INC. | 2016-11-15 | GAM | VICRYL RAPIDE | Detachment Of Device Component | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A VETERINARIAN PERFORMED A PROCEDURE ON AN ANIMAL ON AN UNKNOWN DATE. DURING THE PROCEDURE, THE NEEDLE SEPARATED FROM THE SUTURE AFTER USING ONLY ABOUT A THIRD OF IT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6103656&pc=GAM | 2210968-2016-14954 |  | Malfunction | ETHICON INC. | 2016-11-15 | GAM | VICRYL RAPIDE | Detachment Of Device Component | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT AN ANIMAL UNDERWENT A PROCEDURE ON AN UNKNOWN DATE. DURING THE PROCEDURE, THE NEEDLE SEPARATED FROM THE SUTURE AFTER USING ONLY ABOUT A THIRD OF IT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6103826&pc=GAM | 2210968-2016-14956 | 2016-10-26 | Malfunction | ETHICON INC. | 2016-11-15 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. THE STERILE PACKAGE WAS FOUND WITH A HOLE ON IT WHEN THE NURSE REMOVED IT FROM THE BOX. ANOTHER SAME LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE FOIL WAS EXAMINED FOR VISUAL INSPECTION AND A HOLE IN CAVITY WAS OBSERVED IN A CORNER ON THE PACKAGE. THE HOLE IS OBSERVED TO BE MADE FROM THE OUTSIDE TO INSIDE AND MARKS ARE OBSERVED OF SOME POINTED OBJECT NEAR OF THE HOLE ON THE FOIL. THE PACKAGING PROCESS HAS BEEN REVIEWED AND THERE IS NOT AREA WHERE THIS HOLE COULD HAVE BEEN CAUSED. IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6100394&pc=GAM | 2210968-2016-14911 | 2016-10-25 | Malfunction | ETHICON INC. | 2016-11-14 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. BEFORE THE PROCEDURE, WHEN THE FOIL WAS OPENED THEY FOUND THE SUTURE HAD TURNED INTO POWDER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES FOR THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6096926&pc=GAM | 2210968-2016-14889 |  | Malfunction | ETHICON INC. | 2016-11-11 | GAM | VICRYL RAPIDE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CIRCUMCISION PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. THE SUTURE TORE/BROKE ON POST-OP DAY 1-3. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6092117&pc=GAM | 2210968-2016-14852 | 2016-10-19 | Malfunction | ETHICON INC. | 2016-11-10 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. THE PACKAGE WAS FOUND WITH PINHOLES IN THE PACKAGE. STERILE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH JJM945, EXPIRATION DATE: 07/31/2020, DATE OF MFG.: 08/07/2015. BATCH HL5234, EXPIRATION DATE: 07/31/2019. DATE OF MFG.: 10/29/2014. BATCH HBZ772, EXPIRATION DATE: 01/31/2019. DATE OF MFG.: 02/12/2014. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6092119&pc=GAM | 2210968-2016-14853 | 2016-10-19 | Malfunction | ETHICON INC. | 2016-11-10 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. THE PACKAGE WAS FOUND WITH PINHOLES IN THE PACKAGE. STERILE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH JJM945, EXPIRATION DATE: 07/31/2020, DATE OF MFG.: 08/07/2015. BATCH HL5234, EXPIRATION DATE: 07/31/2019. DATE OF MFG.: 10/29/2014. BATCH HBZ772, EXPIRATION DATE: 01/31/2019. DATE OF MFG.: 02/12/2014. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6092120&pc=GAM | 2210968-2016-14851 | 2016-10-19 | Malfunction | ETHICON INC. | 2016-11-10 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. THE PACKAGE WAS FOUND WITH PINHOLES IN THE PACKAGE. STERILE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH JJM945, EXPIRATION DATE: 07/31/2020, DATE OF MFG.: 08/07/2015. BATCH HL5234, EXPIRATION DATE: 07/31/2019. DATE OF MFG.: 10/29/2014. BATCH HBZ772, EXPIRATION DATE: 01/31/2019. DATE OF MFG.: 02/12/2014. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6092126&pc=GAM | 2210968-2016-14854 | 2016-10-19 | Malfunction | ETHICON INC. | 2016-11-10 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. THE PACKAGE WAS FOUND WITH PINHOLES IN THE PACKAGE. STERILE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH JJM945, EXPIRATION DATE: 07/31/2020, DATE OF MFG.: 08/07/2015. BATCH HL5234, EXPIRATION DATE: 07/31/2019. DATE OF MFG.: 10/29/2014. BATCH HBZ772, EXPIRATION DATE: 01/31/2019. DATE OF MFG.: 02/12/2014. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6092144&pc=GAM | 2210968-2016-14857 | 2016-10-19 | Malfunction | ETHICON INC. | 2016-11-10 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. THE PACKAGE WAS FOUND WITH PINHOLES IN THE PACKAGE. STERILE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH JJM945, EXPIRATION DATE: 07/31/2020, DATE OF MFG.: 08/07/2015. BATCH HL5234, EXPIRATION DATE: 07/31/2019. DATE OF MFG.: 10/29/2014. BATCH HBZ772, EXPIRATION DATE: 01/31/2019. DATE OF MFG.: 02/12/2014. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6092146&pc=GAM | 2210968-2016-14856 | 2016-10-19 | Malfunction | ETHICON INC. | 2016-11-10 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. THE PACKAGE WAS FOUND WITH PINHOLES IN THE PACKAGE. STERILE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH JJM945, EXPIRATION DATE: 07/31/2020, DATE OF MFG.: 08/07/2015. BATCH HL5234, EXPIRATION DATE: 07/31/2019. DATE OF MFG.: 10/29/2014. BATCH HBZ772, EXPIRATION DATE: 01/31/2019. DATE OF MFG.: 02/12/2014. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6092148&pc=GAM | 2210968-2016-14855 | 2016-10-19 | Malfunction | ETHICON INC. | 2016-11-10 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. THE PACKAGE WAS FOUND WITH PINHOLES IN THE PACKAGE. STERILE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH JJM945, EXPIRATION DATE: 07/31/2020, DATE OF MFG.: 08/07/2015. BATCH HL5234, EXPIRATION DATE: 07/31/2019. DATE OF MFG.: 10/29/2014. BATCH HBZ772, EXPIRATION DATE: 01/31/2019. DATE OF MFG.: 02/12/2014. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6092514&pc=GAM | 6092514 | 2016-08-31 | Malfunction | ETHICON, INC. | 2016-11-10 | GAM | VICRYL | Break; Detachment Of Device Component | Foreign Body In Patient |  |  | 1 | Event Description: WHILE CLOSING THE SURGICAL WOUND, THE SURGEON DISCOVERED THE SUTURE NEEDLE HAD BROKEN IN HALF. BOTH SEGMENTS WERE RETRIEVED, A FOLLOW UP XRAY OBTAINED CONFIRMING THE PATIENT WAS NEGATIVE FOR RETAINED FOREIGN BODY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6090887&pc=GAM | 2210968-2016-14839 | 2016-10-20 | Malfunction | ETHICON INC. | 2016-11-09 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. PRIOR TO THE PROCEDURE, A SUTURE PACKAGE WAS PULLED AND PIN HOLES WERE NOTICED IN THE PACKAGE. ANOTHER LIKE PACKAGE WAS OPENED AND USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: TWO LABELED ZIPPER TRAYS WITH SUTURES NOT DISPENSED WITH PAPER LIDS SECURED CORRECTLY WERE RECEIVED FOR EVALUATION AND NO PIN HOLES WERE OBSERVED ON THE PAPER LIDS. ACCORDING TO THE EVALUATION RESULT, THE FOILS WERE NOT RETURNED FOR EVALUATION AND THE SAMPLES RECEIVED HAD NO PIN HOLES OBSERVED ON THE PAPER LIDS. THE SAMPLES MET THE FG REQUIREMENTS FOR NEEDLE PULL AND TENSILE STRENGTH. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 11/09/2016. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6087917&pc=GAM | 2210968-2016-14814 | 2016-10-20 | Malfunction | ETHICON INC. | 2016-11-08 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. THE STAFF PULLED SUTURE FROM THE BOX AND NOTICED THE PACKAGE HAD PIN HOLES; SO THEY DID NOT USE IT. A NEW PACKAGE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). OPEN FOIL WAS RECEIVED FOR EVALUATION. THE OPEN FOIL WAS EXAMINED FOR VISUAL INSPECTION AND PRESENT EXCESSIVE MANIPULATION, AND MULTIPLES HOLES IN CAVITY. THE CONDITIONS OF THE HOLES COULD NOT BE DETERMINED DUE TO THE CONDITION OF THE FOIL RECEIVED. IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 11/08/2016. (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6070078&pc=GAM | 2210968-2016-14626 |  | Malfunction | ETHICON INC. | 2016-11-01 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT DURING INSPECTION OF THE SURGICAL DISPOSABLE STORE OF THE PHARMACY ON (B)(6) 2016, THE INSPECTOR REMOVED THE SUTURE FOR TESTING. THE SUBSEQUENT ANALYSIS ON 09/24/2016 STATED THAT THE SUTURE WAS CONSIDERED TO BE SUB-STANDARD ON THE BASIS OF STERILITY TEST PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 11/01/2016. (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE MAKE EXTRA EFFORT TO GET PRODUCT BACK FROM THE HOSPITAL FOR THE SAME PRODUCT CODE VCP316H, LOT JK8DHQR0 FOR ETHICON TO PERFORM TESTING. IS IT POSSIBLE TO OBTAIN SAMPLES OF THE SAME PRODUCT CODE AND LOT NUMBER FOR ANALYSIS? WAS THERE ANY ANOMALIES (HOLE, TEAR, PUNCTURE, OPEN OR INCOMPLETE SEAL) NOTICED ON THE SUTURE PACKAGES THAT COMPROMISES STERILITY?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6066233&pc=GAM | 2210968-2016-14579 | 2016-09-30 | Malfunction | ETHICON INC. | 2016-10-31 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | Device Embedded In Tissue or Plaque | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE NEEDLE BROKE AND HAS NOT BEEN RETRIEVED. THE X-RAY WAS PERFORMED BUT THE TIP OF THE NEEDLE WAS NOT VISUALIZED. THERE WERE NO PATIENT CONSEQUENCE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHERE WAS THE NEEDLE HELD DURING USE (END, SWAGE, TIP, MIDDLE)? UNK. WHAT TISSUE WAS THE NEEDLE TIP RETAINED IN? THE NEEDLE TIP HAS NOT BEEN FOUND (NEITHER IN THE PATIENT NOR IN THE OPERATIVE FIELD). WHAT SIZE OF NEEDLE PIECE (IN MM) REMAINS IN THE PATIENT TISSUE? THE NEEDLE TIP HAS NOT BEEN FOUND (NEITHER IN THE PATIENT NOR IN THE OPERATIVE FIELD). WERE ALL THE NEEDLE PIECES REMOVED FROM THE PATIENT? THE NEEDLE TIP HAS NOT BEEN FOUND (NEITHER IN THE PATIENT NOR IN THE OPERATIVE FIELD). DOES THE SURGEON PLAN TO REMOVE THE NEEDLE PIECE FROM THE PATIENT? THE NEEDLE TIP HAS NOT BEEN FOUND (NEITHER IN THE PATIENT NOR IN THE OPERATIVE FIELD). WHAT IS THE SURGEON OPINION AS TO THE CAUSATIVE FACTOR FOR THE NEEDLE BREAKAGE? UNK. DO YOU HAVE THE OTHER HALF OF THE NEEDLE FOR EVALUATION? THE NEEDLE TIP HAS NOT BEEN FOUND (NEITHER IN THE PATIENT NOR IN THE OPERATIVE FIELD). WHAT ARE THE PATIENT AGE, WEIGHT AND MEDICAL HISTORY? UNK. WHAT IS THE CURRENT CONDITION OF THE PATIENT? FINE. Manufacturer Narrative: THIS MEDWATCH IS BEING VOIDED BECAUSE IT IS A DUPLICATE. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHERE WAS THE NEEDLE HELD DURING USE (END, SWAGE, TIP, MIDDLE)? WHAT TISSUE WAS THE NEEDLE TIP RETAINED IN? WHAT SIZE OF NEEDLE PIECE (IN MM) REMAINS IN THE PATIENT TISSUE? WERE ALL THE NEEDLE PIECES REMOVED FROM THE PATIENT? DOES THE SURGEON PLAN TO REMOVE THE NEEDLE PIECE FROM THE PATIENT? WHAT IS THE SURGEON OPINION AS TO THE CAUSATIVE FACTOR FOR THE NEEDLE BREAKAGE? DO YOU HAVE THE OTHER HALF OF THE NEEDLE FOR EVALUATION? WHAT ARE THE PATIENT AGE, WEIGHT AND MEDICAL HISTORY? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6061898&pc=GAM | 2210968-2016-14542 | 2016-10-06 | Malfunction | ETHICON INC. | 2016-10-27 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. IT WAS REPORTED THAT THE PACKAGE HAD BEEN ALREADY OPENED BEFORE IT WAS OPENED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: DATE SENT TO THE FDA: 12/01/2016. THE OPENED FOIL WAS EXAMINED AND IT WAS OBSERVED WRINKLES BY USE ON THE FOIL. THE FOIL WAS COMPLETELY OPEN. THE SEAL AREA WAS EXAMINED AND THE SEALED WAS CONTINUOUS AND NO DEFECTS WERE NOTED. THE NEEDLE/SUTURES COMBINATION WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS OR SUTURE DEGRADATION WERE FOUND. ACCORDING THE SAMPLES CONDITIONS, NO DEFECTS WERE OBSERVED ON THE SEAL AND THE SAMPLES MET THE FG REQUIREMENTS OF NEEDLE PULL AND TENSILE STRENGTH. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6055617&pc=GAM | 2210968-2016-14506 |  | Malfunction | ETHICON INC. | 2016-10-25 | GAM | VICRYL RAPIDE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. IT WAS ALSO REPORTED THAT THE PARTICULAR PACKAGE CONTAINED A DATAMATRIX IN AN UPPER RIGHT HAND CORNER THAT DOES NOT HAVE ENOUGH CONTRAST. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THERE WAS NO SAMPLE RECEIVED FOR ANALYSIS. ONLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, IT WAS NOT POSSIBLE TO PERFORM ANY EVALUATION ON THE PHOTO PROVIDED. DATAMATRIX READING TEST IS ONLY POSSIBLE WITH THE ACTUAL DEVICE RETURNED. NO CONCLUSION CAN BE REACHED FOR THE CODE READABILITY BY USING A PHOTO. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6045634&pc=GAM | 2210968-2016-14462 | 2016-09-29 | Malfunction | ETHICON INC. | 2016-10-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS FOUND THAT THE THICKNESS OF THE SUTURE HAD BEEN CHANGED AT ABOUT 1 CM FROM THE TIP OF THE SUTURE WHEN THE PACKAGE WAS OPENED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: THE ACTUAL SAMPLE WAS RETURNED FOR ANALYSIS. THE SUTURE WAS VISUALLY EXAMINED AND IT WAS OBSERVED THAT IN THE END NEAR THE TIP THE SUTURE WAS INSERTING IN THE CENTER OF THE SUTURE, CAUSING A BUNCHING DEFECT . IT WAS APPLIED TENSION IN THE SUTURE AND IT WAS OBSERVED THAT THE INSERTED AREA CAN BE EXTRACT, HOWEVER THIS PART WAS THICKER THAN REST OF THE SUTURE. ACCORDING TO THE SAMPLE CONDITION, THE ASSIGNABLE CAUSE IS A BUNCHING DEFECT. COMPLAINT INFORMATION WILL BE INCLUDED IN COMPLAINT TRENDING TO DETERMINE IF FURTHER ACTION IS NECESSARY. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5943135&pc=GAM | 2210968-2016-13504 |  | Malfunction | ETHICON INC. | 2016-09-12 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN CARDIOVASCULAR SURGICAL PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE EASILY GOT BENT. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4) DATE SENT TO THE FDA: 09/12/2016. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5943153&pc=GAM | 2210968-2016-13505 |  | Malfunction | ETHICON INC. | 2016-09-12 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN CARDIOVASCULAR SURGICAL PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE EASILY GOT BENT. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5943166&pc=GAM | 2210968-2016-13507 |  | Malfunction | ETHICON INC. | 2016-09-12 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN CARDIOVASCULAR SURGICAL PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE EASILY GOT BENT. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4) DATE SENT TO THE FDA: 09/12/2016. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926010&pc=GAM | 2210968-2016-12736 | 2016-08-15 | Malfunction | ETHICON INC. | 2016-09-02 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Unsealed Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. IT WAS ALSO REPORTED THAT, DURING THE PROCEDURE, THE TECH WAS UNSURE IF IT WAS COMPLETELY SEALED/STERILE; ERROR IN SIDE OF CAUTION. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5922735&pc=GAM | 2210968-2016-12518 | 2016-07-28 | Malfunction | ETHICON INC. | 2016-09-01 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HYSTERECTOMY ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE CLOSING OF THE VAGINAL VAULT, THE NEEDLE BENT. THERE WERE NO REPORTED PATIENT CONSEQUENCES. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: ONE REPRESENTATIVE SAMPLE WAS RECEIVED. THE REPRESENTATIVE SAMPLE NEEDLE WAS INSPECTED AND NO DEFECTS WERE OBSERVED ON THE NEEDLE OF THE SAMPLE. THE NEEDLE RETAINED SAMPLES WERE VISUALLY INSPECTED FOR PHYSICAL APPEARANCE LIKE CURVATURE, TIP, ETC. ATTRIBUTE DEFECTS SUCH AS BEND, CRACKED BARREL AND WERE FOUND SATISFACTORY. NO DEFECTS WERE OBSERVED IN THE RETAIN SAMPLE. THESE RETAIN SAMPLES WERE TESTED FOR STIFFNESS & DUCTILITY AND FOUND TO MEET SPECIFICATION. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917892&pc=GAM | 2210968-2016-12177 | 2016-07-24 | Malfunction | ETHICON INC. | 2016-08-31 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EYE PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED IN THE SCLERA TISSUE LAYER. DURING THE PROCEDURE, THE DOCTOR FELT THE NEEDLE WAS MORE DUCTILE OR BENT. THE PROCEDURE WAS COMPLETED USING THE PRODUCT ON HAND. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917895&pc=GAM | 2210968-2016-12178 | 2016-07-24 | Malfunction | ETHICON INC. | 2016-08-31 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EYE PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED IN THE SCLERA TISSUE LAYER. DURING THE PROCEDURE, THE DOCTOR FELT THE NEEDLE WAS MORE DUCTILE OR BENT. THE PROCEDURE WAS COMPLETED USING THE PRODUCT ON HAND. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5888829&pc=GAM | 2210968-2016-11453 |  | Malfunction | ETHICON INC. | 2016-08-19 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT AN UNKNOWN ANIMAL PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. THE SUTURE BROKE POST-OPERATIVELY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW MANY PATIENTS HAD INTRA OP AND POST OB BREAKAGE (PLEASE LIST EVENTS SEPARATELY)? HOW MANY SUTURE BROKE PER PATIENT? HOW WAS THE BREAKAGE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE SURGICAL/OPERATION NOTES. HOW IS/ARE THE ANIMALS TODAY? WILL ANY PRODUCT BE RETURNED FOR ANALYSIS? ADDITIONAL INFORMATION WAS OBTAINED: THE VET STATED THAT THE SUTURE BROKE INTRA OP AND POST OP. PRODUCT CODE IS J548G, LOT# JPZ900.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5884898&pc=GAM | 2210968-2016-11412 | 2016-07-15 | Malfunction | ETHICON INC. | 2016-08-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS REPORTED THAT THE NEEDLE BECAME BENT AT THE CENTER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: THE ACTUAL SAMPLE WAS RETURNED FOR ANALYSIS. THE USED NEEDLE WAS VISUALLY INSPECTED. IT WAS OBSERVED THAT THE NEEDLE WAS BROKEN AND WAS NOT OBSERVED NEEDLE BENDING. A FRACTURE WAS OBSERVED NEAR THE SUTURE ATTACHMENT OF THE NEEDLE. THE NEEDLE EXHIBITED AMOUNTS OF INTERGRANULAR FAILURE. COMPLAINT INFORMATION WILL BE INCLUDED IN COMPLAINT TRENDING THAT IS REVIEWED ON A REGULAR BASIS TO DETERMINE IF FURTHER ACTION IS NECESSARY. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5879495&pc=GAM | 2210968-2016-11345 | 2016-07-25 | Malfunction | ETHICON INC. | 2016-08-16 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GENERAL SURGICAL PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE NEEDLE WAS FOUND TO BE BENT WHEN IT WAS PULLED FROM THE PACKAGE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL EXAMINATION REVEALED THAT THE TIP OF THE NEEDLE WAS DAMAGED. REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL EXAMINATION REVEALED THAT THE TIP OF THE NEEDLES WERE DAMAGED. THE CONDITION OF THE SAMPLES INDICATE THAT THE ASSIGNABLE CAUSE IS NEEDLE TIP DEFECT. COMPLAINT INFORMATION WILL BE INCLUDED IN COMPLAINT TRENDING THAT IS REVIEWED ON A REGULAR BASIS TO DETERMINE IF FURTHER ACTION IS NECESSARY. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5871194&pc=GAM | 5871194 | 2016-04-08 | Malfunction | ETHICON INC | 2016-08-12 | GAM | VICRYL PLUS | Component Missing | Device Embedded In Tissue or Plaque |  |  | 1 | Event Description: PT WAS UNDERGOING AN EXPLORATORY LAPAROTOMY WITH RIGHT PARATUBAL CYST REMOVAL, PARTIAL RIGHT SALPINGECTOMY, RIGHT OVARIAN CYSTECTOMY, AND LEFT OVARIAN CYSTECTOMY. DURING THE USE OF THE 3-0 VICRYL SH, THE SURGICAL TECH NOTICED THE TIP OF THE NEEDLE MISSING. THE OVARY, AFTER CLOSURE ON THE LEFT SIDE, WAS OPENED AND INSPECTED TWICE WITH DISSECTION. NO NEEDLE TIP NOTED. PRIOR TO THE CLOSURE OF THE ABDOMEN AN X-RAY WAS TAKEN OF BOTH THE ABDOMEN AND PELVIS, 2 VIEWS. NO DEFINITE EVIDENCE OF A RETAINED FOREIGN BODY NOTED ON THE RADIOLOGY REPORT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5855131&pc=GAM | 2210968-2016-10365 | 2016-05-07 | Malfunction | ETHICON INC. | 2016-08-08 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VET THAT A HORSE UNDERWENT AN ABDOMINAL PROCEDURE WITH REMOVAL OF SMALL INTESTINE STRANGULATED BY A LIPOMA ON (B)(6) 2016 AND THE SUTURE WAS USED IN CONTINUOUS TECHNIQUE TO CLOSE THE LINEA ALBA AND THE BODY WALL. THE VETERINARIAN NOTED THAT THERE WAS NO ISSUE OR ANOMALY NOTED WITH THE SUTURE PRIOR TO/ OR DURING USE AND INSTRUMENTS ARE USED TO TIE KNOTS. AS REPORTED THERE WAS NO SUTURE BREAKAGE DURING THE PROCEDURE. FOLLOWING THE PROCEDURE, THE SUTURE WAS NOTED TO BE BROKEN IN THE MIDDLE OF THE SUTURE LINE WITH KNOTS INTACT IN THE LINEA ALBA. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CAN YOU CLARIFY THE PROCEDURE DATE WAS (B)(6) 2016? WERE THERE ANY ANOMALIES WITH THE SUTURE NOTED DURING USE? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? WAS THE SUTURE BREAKING DURING THE PROCEDURE? CAN YOU DESCRIBE THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WHAT IS THE SURGEON OPINION FOR THE SUTURE BREAKAGE? DR. STATED PATIENT HAD ABDOMINAL PROCEDURE WITH REMOVAL OF SMALL INTESTINE BY THIRD YEAR RESIDENT ON (B)(6) VICRYL SUTURE, CODE J880 USING CONTINUOUS TECHNIQUE TO CLOSE THE LINEA ALBA AND THE BODY WALL. THE VETERINARIAN NOTED THAT THERE WAS NO ISSUE OR ANOMALY NOTED WITH THE SUTURE PRIOR TO/ OR DURING USE AND INSTRUMENTS ARE USED TO TIE KNOTS. SHE ADVISED THAT THERE WAS NO SUTURE BREAKAGE DURING THE PROCEDURE. DR. REMOVED THE BODY WALL AND REPORTED THAT THE SUTURE WAS NOTED TO BE BROKEN IN THE MIDDLE OF THE SUTURE LINE WITH KNOTS INTACT IN THE LINEA ALBA. REPRESENTATIVE SAMPLES WERE EXAMINED FOR VISUAL DEFECTS AND NONE WAS FOUND. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGE WERE FOUND. ACCORDING THE SAMPLES CONDITIONS, NO SUTURE BREAKAGE WERE FOUND ON THE UNOPENED SAMPLE AND MET REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5843096&pc=GAM | 2210968-2016-11043 | 2016-07-15 | Malfunction | ETHICON INC. | 2016-08-03 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING UTERINE CLOSURE, THE NEEDLE GOT DISFIGURED . NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY WHAT IS MEANT BY NEEDLE GOT DISFIGURED. THE NEEDLE WAS BENT DURING SURGERY. DID THE NEEDLE BEND WHILE IN USE? YES. IS THE ACTUAL SAMPLE BEING RETURNED FOR EVALUATION? YES. WHAT IS THE LOT NUMBER? T6001. Manufacturer Narrative: DATE SENT TO FDA: 9/13/2016. THE ACTUAL SAMPLE WAS RETURNED FOR ANALYSIS. THE ACTUAL, USED SAMPLE PRESENTED A DEFORMED NEEDLE WITH ATTACHED SUTURE SEGMENT OF APPROX. 4 CM. VISUAL EXAMINATION REVEALED MULTIPLE SEVERE USER INSTRUMENT MARKS ON NEEDLE BODY, AND ALSO SOME MARKS IN THE NEEDLE ATTACHMENT AREA. THE NEEDLE HAD BEEN BENT UP FROM THE MIDDLE TO THE ATTACHMENT END, VERY LIKELY DUE TO EXCESSIVE USER HANDLING. ETHICON IFU RECOMMENDS: GRASPING AT THE BUTT OR ATTACHMENT END COULD CAUSE BENDING OR BREAKAGE. GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. VISUAL INSPECTION AND DUCTILITY AND STIFFNESS TEST OF THE RETAINED SAMPLES MEET THE SPECIFIED QUALITY REQUIREMENTS. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY WHAT IS MEANT BY ¿NEEDLE GOT DISFIGURED¿? DID THE NEEDLE BEND WHILE IN USE? IS THE ACTUAL SAMPLE BEING RETURNED FOR EVALUATION? WHAT IS THE LOT NUMBER? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5843171&pc=GAM | 2210968-2016-11045 | 2016-07-15 | Malfunction | ETHICON INC. | 2016-08-03 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING UTERINE CLOSURE, THE NEEDLE GOT DISFIGURED. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY WHAT IS MEANT BY NEEDLE GOT DISFIGURED THE NEEDLE WAS BENT DURING SURGERY. DID THE NEEDLE BEND WHILE IN USE? YES. IS THE ACTUAL SAMPLE BEING RETURNED FOR EVALUATION? YES. WHAT IS THE LOT NUMBER? T5004. Manufacturer Narrative: DATE SENT TO FDA: 9/13/2016. THE ACTUAL SAMPLE WAS RECEIVED FOR ANALYSIS. THE NEEDLE WAS RECEIVED IN SLIGHTLY BENT CONDITION. VISUAL EXAMINATION REVEALED SEVERAL INSTRUMENT MARKS AT THE ATTACHMENT AREA WHICH INDICATES THAT THE NEEDLE WAS HANDLED THERE. USER HANDLING WAS DETERMINED TO BE THE CAUSE OF THE NEEDLE BENDING BASED ON THE OBSERVED INSTRUMENT MARKS. VISUAL INSPECTION AND DUCTILITY AND STIFFNESS TEST OF THE RETAINED SAMPLES MEET THE SPECIFIED QUALITY REQUIREMENTS. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 08/03/2016. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY WHAT IS MEANT BY ¿NEEDLE GOT DISFIGURED¿. DID THE NEEDLE BEND WHILE IN USE? IS THE ACTUAL SAMPLE BEING RETURNED FOR EVALUATION? WHAT IS THE LOT NUMBER? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5837011&pc=GAM | 5837011 | 2016-07-21 | Malfunction | ETHICON INC | 2016-08-01 | GAM | VICRYL PLUS 3-0 CT-1 | Component Missing; Defective Device | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: DURING TOTAL VAGINAL HYSTERECTOMY AND BILATERAL SALPINGO-OOPHORECTOMY, WHILE SURGEON SUTURING THE IP LIGAMENT, HE NOTICED THAT THE TIP OF THE SUTURE NEEDLE WAS MISSING. THE SUTURE NEEDLE WAS IMMEDIATELY REMOVED FROM THE SURGICAL FIELD AND X-RAY PERFORMED WHICH WAS NEGATIVE FOR BROKEN TIP. SURGEON LATER STATED THAT HE STRONGLY BELIEVES THAT SUTURE WAS DEFECTIVE AND THERE MAY HAVE NOT BEEN A TIP TO BEGIN WITH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5827855&pc=GAM | 2210968-2016-10830 | 2015-10-07 | Malfunction | ETHICON INC. | 2016-07-27 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORAL PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. TWO WEEKS AFTER THE SURGERY THE SUTURE HAD NOT YET DISSOLVED, SO THE SURGEON REMOVED IT. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816292&pc=GAM | 2210968-2016-10726 | 2016-06-07 | Malfunction | ETHICON INC. | 2016-07-22 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY: DID BOTH EVENTS OCCUR DURING THE SAME PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5808216&pc=GAM | 5808216 | 2016-03-11 | Malfunction | ETHICON, INC. | 2016-07-20 | GAM | VICRYL | Break; Component Falling | Complaint, Ill-Defined |  |  | 1 | Event Description: A #3-0 VICRYL NEEDLE BROKE OFF INSIDE PATIENT'S RECTUM DURING PROCEDURE TO REPAIR RECTAL PROLAPSE. THE NEEDLE BROKE OFF IN THE BOWEL TISSUE. THE PHYSICIAN WAS ABLE TO RETRIEVE THE BROKEN PIECE. THIS DELAYED THE CASE FOR SOME TIME. NO FURTHER ACTION NEEDED. WE DON'T HAVE THE PACKAGE FOR THE NEEDLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5806117&pc=GAM | 2210968-2016-10600 | 2016-06-22 | Malfunction | ETHICON INC. | 2016-07-19 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Unsealed Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND THE SUTURE WAS USED. DURING DISPENSING, A SUTURE PART WAS STUCK IN THE SEALING PART OF THE ALUMINUM PACKAGE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: DATE SENT TO THE FDA: 08/18/2016. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS THE NAME OF THE PROCEDURE? NO INFORMATION. WHEN THE SUTURE WAS COMING OFF OF THE TRAY WAS THE SUTURE DETACHED FROM THE NEEDLE AT ANY POINT? THE SUTURE HAS BEEN SEALED BY THE ALUMNI PACKAGE AND IT WAS UNABLE TO TAKE THE SUTURE OUT . WAS THE SUTURE COMING OFF OF THE TRAY NOTICED WHEN OPENING THE PACKAGE OR DURING DISPENSING? NO INFORMATION. THE OPEN FOIL WAS EXAMINED AND IT WAS OBSERVED MARKS OF THE SUTURE CAUGHT ON THE SEAL AREA, BUT DOES NOT EXTEND COMPLETELY THROUGH SEAL AREA. ONE SECTION OF THE BODY SUTURE WAS FLAT WHICH INDICATING THAT SUTURE WAS TRAPPED IN THE SEAL AREA AND DUE TO THIS CONDITION THE SUTURE CANNOT BE DISPENSED CORRECTLY. ADDITIONALLY, THE STERILE BARRIER WAS NOT COMPROMISED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS THE NAME OF THE PROCEDURE? NO INFORMATION. WHEN THE SUTURE WAS COMING OFF OF THE TRAY WAS THE SUTURE DETACHED FROM THE NEEDLE AT ANY POINT? NO. THE SUTURE WAS NOT DETACHED. WHEN THE SITUATION HAPPENED IS UNKNOWN. WAS THE SUTURE COMING OFF OF THE TRAY NOTICED WHEN OPENING THE PACKAGE OR DURING DISPENSING? DISPENSING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5794478&pc=GAM | 2210968-2016-10448 | 2016-06-22 | Malfunction | ETHICON INC. | 2016-07-14 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UMBILICAL HERNIA REPAIR ON (B)(6) 2016 AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS BENDING WHILE SUTURING THROUGH FASCIA. OTHER SUTURE WAS USED TO CONTINUE THE PROCEDURE WITH NO CONSEQUENCE TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. REPRESENTATIVE SAMPLES WERE EXAMINED FOR VISUAL DEFECTS ON THE PACKETS NONE WAS FOUND; NO ATTACHMENT DEFECTS OR NEEDLE BENDING WERE OBSERVED ON THE REPRESENTATIVE SAMPLES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5794629&pc=GAM | 2210968-2016-10449 |  | Malfunction | ETHICON INC. | 2016-07-14 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS BENDING. OTHER SUTURE WAS USED TO CONTINUE THE PROCEDURE WITH NO CONSEQUENCE TO THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5742710&pc=GAM | 2210968-2016-09936 |  | Malfunction | ETHICON INC. | 2016-06-22 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE BENT WHILE SUTURING. THE PROCEDURE WAS COMPLETED USING A LIKE DEVICE. THERE WAS NO PATIENT CONSEQUENCE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE NEEDLE FOUND BENT IN THE PACKAGE? NO. DID THE NEEDLE BEND DURING DISPENSING/ REMOVAL FROM PACKAGE? NO. WHERE WAS THE NEEDLE GRASPED (SWAGE, MIDDLE, BODY)? UNKNOWN. WHERE DID THE NEEDLE BEND (SWAGE, BODY, MIDDLE, TIP)? UNKNOWN. DID THE NEEDLE BREAK INTO TWO PIECES? NO. WHAT IS THE PROCEDURE NAME? UNKNOWN. WHAT IS THE PROCEDURE DATE? UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5738571&pc=GAM | 2210968-2016-09887 | 2016-04-30 | Malfunction | ETHICON INC. | 2016-06-21 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LATERAL PERINEAL REPAIR SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. SEVEN DAYS AFTER THE PROCEDURE, THE SUTURE BROKE. ADDITIONAL INFORMATION WAS REQUESTED BUT COULD NOT BE PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE WHAT TISSUE WAS THE SUTURE USED ON? HOW WAS THE SUTURE PLACED, INTERRUPTED OR CONTINUOUS? WHAT OTHER SKIN CLOSURE WAS USED WITH THE SUTURE? WHAT DATE DID THE PATIENT PRESENT WITH SYMPTOMS? WHAT IS THE SURGEON OPINION AS TO THE CONTRIBUTING FACTORS TO THE SYMPTOMS? WAS THERE ANY PATIENT EVENT PRIOR TO SYMPTOMS (COUGHING, FALLING, MOVING)? WHAT MEDICAL INTERVENTION WAS PERFORMED FOR THE PATIENT SYMPTOMS? WHAT WAS THE DATE OF THE SECOND PROCEDURE? CAN YOU DESCRIBE THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? WAS THE SUTURE BROKEN IN THE MIDDLE? WERE THE KNOTS INTACT OR BROKEN ON THE ENDS? DO YOU HAVE ANY PICTURES OF THE APPEARANCE OF THE SUTURE? WHAT IS THE CURRENT CONDITION OF THE PATIENT? THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5726508&pc=GAM | 2210968-2016-09779 | 2015-11-14 | Malfunction | ETHICON INC. | 2016-06-15 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. THE SUTURE BROKE EASILY WHILE TYING KNOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. TWO POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: AA3424 MFG. DATE: 06/01/2015, EXP.: 05/31/2020. AA6501 MFG. DATE: 06/01/2015, EXP.: 05/31/2020. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY SUTURES BROKE DURING THE PROCEDURE? MANY THREADS BROKE IN THE KNOT DURING SOME PROCEDURES AND OUTSIDE OF PROCEDURES ONLY USING THE HANDS TO TEST THE SUTURE. THERE WAS NO QUANTITATIVE COUNTING. THE SYNERGY FORM STATED THAT THE QUANTITY INVOLVED IN 480. ARE 480 SUTURES BEING RETURNED FOR EVALUATION? THE SAMPLE WAS SENT ON 28/APR/2016. WHAT WAS THE NAME OF THE PROCEDURE? SALES REP. IS UNABLE TO TRACK THE NAME ALL PROCEDURES. WAS THERE ANY DAMAGE TO THE TISSUE? BESIDES OF THE NECESSITY OF DOING A NEW CLOSURE USING OTHER THREAD, THERE WAS NO TISSUE DAMAGE. REPRESENTATIVE SAMPLES WERE RECEIVED FOR EVALUATION. VISUAL AND FUNCTIONAL EXAMINATIONS WERE PERFORMED AND SAMPLES MET THE SPECIFIED QUALITY REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5722729&pc=GAM | 2210968-2016-05851 |  | Malfunction | ETHICON INC. | 2016-06-14 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Detachment Of Device Component | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF AT THE FIRST PASSAGE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. UPON EVALUATION, THE SAMPLE RETURNED SHOWED A LOOSE DETACHED NEEDLE. THE CAUSE OF DETACHING COULD NOT BE DETERMINED BASED ON VISUAL INSPECTION AT THE ACTUAL SINGLE NEEDLE WITHOUT THREAD.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5719686&pc=GAM | 5719686 | 2016-06-03 | Malfunction | ETHICON, INC. | 2016-06-13 | GAM | VICRYL BRAIDED 2-0 | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: WHEN USING A VICRYL BRAIDED 2-0 SUTURE, THE NEEDLE BROKE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5709052&pc=GAM | 2210968-2016-09636 | 2016-04-07 | Malfunction | ETHICON INC. | 2016-06-09 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ONCOLOGICAL PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, A HOLE WAS OBSERVED IN THE SUTURE PACKAGING. THERE WERE NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WERE THE PACKAGES PROPERLY SEALED? YES. WERE THERE ANY DEFICIENCIES OR HOLES NOTE. D IN THE PACKAGES? YES, IT WAS OBSERVED A HOLE IN THE PACKAGING. THE ACTUAL DEVICE WAS RETURNED FOR EVALUATION. IT WAS POSSIBLE TO OBSERVE SMALL TEARS IN THE PACKAGE RECEIVED, HOWEVER IT COULD NOT BE DETERMINED IF THESE WERE PRESENT BEFORE USE OR IF THEY WERE CAUSED DURING OPENING OF THE ENVELOPE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686640&pc=GAM | 2210968-2016-09307 |  | Malfunction | ETHICON INC. | 2016-05-27 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A RESEARCHER THAT A RODENT UNDERWENT A RESEARCH-RELATED SURGICAL PROCEDURE BETWEEN (B)(6) 2016 AND SUTURE WAS USED IN THE MUSCLE IN THE HIND LIMB. AT THE TIME OF EVALUATION, BETWEEN (B)(6) 2016, THE SUTURE WAS STILL THERE AND NOT RESORBED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WE HAVE BEEN USING THESE IN RATS FOR RESEARCH-RELATED SURGERIES. WE USE THE SUTURE IN MUSCLE IN THE HIND LIMB AND THEN WHEN WE LOOK AT THE ANIMALS 2-3 MONTHS LATER, THE SUTURE IS STILL THERE. WE DID THESE PROCEDURES OVER THE LAST MONTH. WE UNFORTUNATELY DO NOT HAVE THE CODE OR LOT NUMBERS. WE HAVE USED THE ENTIRE BOX THAT WE INITIALLY BOUGHT. THE EVENT DATES FOR THE PROCEDURE WERE BETWEEN (B)(6) AND EVALUATION IN (B)(6).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686672&pc=GAM | 2210968-2016-09308 |  | Malfunction | ETHICON INC. | 2016-05-27 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A RESEARCHER THAT A RODENT UNDERWENT A RESEARCH-RELATED SURGICAL PROCEDURE BETWEEN (B)(6) 2016 AND SUTURE WAS USED IN THE MUSCLE IN THE HIND LIMB. AT THE TIME OF EVALUATION, BETWEEN (B)(6) 2016, THE SUTURE WAS STILL THERE AND NOT RESORBED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WE HAVE BEEN USING THESE IN RATS FOR RESEARCH-RELATED SURGERIES. WE USE THE SUTURE IN MUSCLE IN THE HIND LIMB AND THEN WHEN WE LOOK AT THE ANIMALS 2-3 MONTHS LATER, THE SUTURE IS STILL THERE. WE DID THESE PROCEDURES OVER THE LAST MONTH. WE UNFORTUNATELY DO NOT HAVE THE CODE OR LOT NUMBERS. WE HAVE USED THE ENTIRE BOX THAT WE INITIALLY BOUGHT. THE EVENT DATES FOR THE PROCEDURE WERE BETWEEN (B)(6) AND EVALUATION IN (B)(6).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686700&pc=GAM | 2210968-2016-09309 |  | Malfunction | ETHICON INC. | 2016-05-27 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A RESEARCHER THAT A RODENT UNDERWENT A RESEARCH-RELATED SURGICAL PROCEDURE BETWEEN (B)(6) 2016 AND SUTURE WAS USED IN THE MUSCLE IN THE HIND LIMB. AT THE TIME OF EVALUATION, BETWEEN (B)(6) 2016, THE SUTURE WAS STILL THERE AND NOT RESORBED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WE HAVE BEEN USING THESE IN RATS FOR RESEARCH-RELATED SURGERIES. WE USE THE SUTURE IN MUSCLE IN THE HIND LIMB AND THEN WHEN WE LOOK AT THE ANIMALS 2-3 MONTHS LATER, THE SUTURE IS STILL THERE. WE DID THESE PROCEDURES OVER THE LAST MONTH. WE UNFORTUNATELY DO NOT HAVE THE CODE OR LOT NUMBERS. WE HAVE USED THE ENTIRE BOX THAT WE INITIALLY BOUGHT. THE EVENT DATES FOR THE PROCEDURE WERE BETWEEN (B)(6) AND EVALUATION IN APRIL 1-25TH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686713&pc=GAM | 2210968-2016-09310 |  | Malfunction | ETHICON INC. | 2016-05-27 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A RESEARCHER THAT A RODENT UNDERWENT A RESEARCH-RELATED SURGICAL PROCEDURE BETWEEN (B)(6) 2016 AND SUTURE WAS USED IN THE MUSCLE IN THE HIND LIMB. AT THE TIME OF EVALUATION, BETWEEN (B)(6) 2016, THE SUTURE WAS STILL THERE AND NOT RESORBED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WE HAVE BEEN USING THESE IN RATS FOR RESEARCH-RELATED SURGERIES. WE USE THE SUTURE IN MUSCLE IN THE HIND LIMB AND THEN WHEN WE LOOK AT THE ANIMALS 2-3 MONTHS LATER, THE SUTURE IS STILL THERE. WE DID THESE PROCEDURES OVER THE LAST MONTH. WE UNFORTUNATELY DO NOT HAVE THE CODE OR LOT NUMBERS. WE HAVE USED THE ENTIRE BOX THAT WE INITIALLY BOUGHT. THE EVENT DATES FOR THE PROCEDURE WERE BETWEEN (B)(6) AND EVALUATION IN (B)(6).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680238&pc=GAM | 2210968-2016-09241 | 2016-05-03 | Malfunction | ETHICON INC. | 2016-05-25 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING SUTURING THE WOUND, THE NEEDLE BENT. THE PROCEDURE WAS COMPLETED WITH ANOTHER LIKE SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5675745&pc=GAM | 2210968-2016-09197 | 2016-04-27 | Malfunction | ETHICON INC. | 2016-05-24 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A SKIN CLOSURE PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGEON NOTED THE END OF THE NEEDLE HAD BECOME BENT. THERE WAS NO HARM OR ADVERSE OUTCOME REPORTED. OTHER SUTURE WAS USED TO COMPLETE THE PROCEDURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THIS OCCUR WITH 3 PATIENTS / CASES OR 3 BOXES/CASES? ALL THREE ISSUES OCCURRED WITH THREE DIFFERENT PATIENTS. TWO CASES WERE BOTH LAPAROSCOPIC CHOLECYSTECTOMY'S. BOTH CARRIED OUT ON THE (B)(6) 2016. Manufacturer Narrative: IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5675759&pc=GAM | 2210968-2016-09198 | 2016-04-27 | Malfunction | ETHICON INC. | 2016-05-24 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A SKIN CLOSURE PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGEON NOTED THE END OF THE NEEDLE HAD BECOME BENT. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. OTHER SUTURE WAS USED TO COMPLETE THE PROCEDURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THIS OCCUR WITH 3 PATIENTS / CASES OR 3 BOXES/CASES? ALL THREE ISSUES OCCURRED WITH THREE DIFFERENT PATIENTS. TWO CASES WERE BOTH LAPAROSCOPIC CHOLECYSTECTOMY'S. BOTH CARRIED OUT ON THE (B)(6) 2016. Manufacturer Narrative: IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THIS OCCUR WITH 3 PATIENTS / CASES OR 3 BOXES/CASES? IF 3 PATIENTS, PLEASE GIVE SPECIFIC INFORMATION FOR EACH EVENT, DATE, PROCEDURE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5674980&pc=GAM | 2210968-2016-09190 | 2016-05-03 | Malfunction | ETHICON INC. | 2016-05-23 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE DESCRIPTION OF SUTURE PRODUCT AND THE NEEDLE IMAGE ON THE OUTER PACKAGING WERE DEFORMED, AND THE SHAPE AND THE SIZE OF THE NEEDLE WERE NOT EASILY READABLE FROM THE PACKAGING. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. THE ACTUAL DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION. ONE PHOTO OF THE FRONT VIEW OF THE PACKAGING WAS RECEIVED FOR ANALYSIS. IN THE VISUAL INSPECTION OF PHOTO THE PRINT WITH ALL CORRECT INFORMATION INCLUDING NEEDLE TYPE WAS FOUND TO BE WELL-READABLE. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE LABEL INCORRECT? -SHAPE AND THE SIZE OF THE NEEDLE WERE NOT EASILY READABLE. WAS THE LABELING HARD TO READ? -ACCORDING THE INSTRUMENTALIST THE SHAPE AND THE SIZE OF THE NEEDLE WERE NOT EASILY READABLE FROM THE PACKAGING. WAS THE PACKAGING NOT PROPERLY SEALED?- PACKAGING WAS OK. WAS THE PACKAGING TORN?- NO. IS THE FOIL PACKET DEFORMED OR IS THE OUTER PACKAGING DEFORMED?- NO. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5669526&pc=GAM | 2210968-2016-09129 |  | Malfunction | ETHICON INC. | 2016-05-20 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A RESEARCHER THAT A RODENT UNDERWENT A RESEARCH-RELATED SURGICAL PROCEDURE BETWEEN (B)(6) 2016 AND SUTURE WAS USED IN THE MUSCLE IN THE HIND LIMB. AT THE TIME OF EVALUATION, BETWEEN (B)(6) 2016, THE SUTURE WAS STILL THERE AND NOT RESORBED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661082&pc=GAM | 2210968-2016-09014 |  | Malfunction | ETHICON INC. | 2016-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN EYE SURGERY ON UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, AFTER THREE OR FOUR INTRA-TISSUE PASSAGES, THE NEEDLE BECAME DEFORMED AND TWISTED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. THE SAMPLES WERE VISUALLY INSPECTED . THE RESULTS MET THE SPECIFIED QUALITY REQUIREMENTS. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661137&pc=GAM | 2210968-2016-09015 |  | Malfunction | ETHICON INC. | 2016-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT AN EYE SURGERY ON UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, AFTER THREE OR FOUR INTRA-TISSUE PASSAGES, THE NEEDLE BECAME DEFORMED AND TWISTED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661156&pc=GAM | 2210968-2016-09016 |  | Malfunction | ETHICON INC. | 2016-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT AN EYE SURGERY ON UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, AFTER THREE OR FOUR INTRA-TISSUE PASSAGES, THE NEEDLE BECAME DEFORMED AND TWISTED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661190&pc=GAM | 2210968-2016-09017 |  | Malfunction | ETHICON INC. | 2016-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT AN EYE SURGERY ON UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, AFTER THREE OR FOUR INTRA-TISSUE PASSAGES, THE NEEDLE BECAME DEFORMED AND TWISTED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661224&pc=GAM | 2210968-2016-09018 |  | Malfunction | ETHICON INC. | 2016-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT AN EYE SURGERY ON UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, AFTER THREE OR FOUR INTRA-TISSUE PASSAGES, THE NEEDLE BECAME DEFORMED AND TWISTED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661250&pc=GAM | 2210968-2016-09019 |  | Malfunction | ETHICON INC. | 2016-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT AN EYE SURGERY ON UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, AFTER THREE OR FOUR INTRA-TISSUE PASSAGES, THE NEEDLE BECAME DEFORMED AND TWISTED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5662209&pc=GAM | 2210968-2016-09027 | 2016-04-26 | Malfunction | ETHICON INC. | 2016-05-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE AND SUTURE WAS USED. DURING THE PROCEDURE AT THE TIME OF CLOSING THE INCISION PORT THE NEEDLE WAS BENT. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WAS NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ONE REPRESENTATIVE SAMPLE RETURNED AND RETAINED SAMPLES WERE EVALUATED. THE INTACT REPRESENTATIVE SAMPLE WAS INSPECTED FOR APPEARANCE AND IT WAS FOUND TO MEET THE SPECIFICATION. THE NEEDLE RETAIN SAMPLES WERE VISUALLY INSPECTED FOR PHYSICAL APPEARANCE AND WERE FOUND SATISFACTORY. THESE RETAIN SAMPLES WERE TESTED FOR STIFFNESS & DUCTILITY AND FOUND TO MEET SPECIFICATION. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5648555&pc=GAM | 2210968-2016-08878 | 2016-04-06 | Malfunction | ETHICON INC. | 2016-05-11 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED POST 14-15 DAYS OPERATIVELY THE SUTURE DID NOT DISSOLVE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628798&pc=GAM | 2210968-2016-08681 | 2016-04-11 | Malfunction | ETHICON INC. | 2016-05-03 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. POST OPERATIVELY, THE KNOTS CAME LOOSE FROM THE INCISION. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5606470&pc=GAM | 2210968-2016-08399 |  | Malfunction | ETHICON INC. | 2016-04-25 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break; Detachment Of Device Component | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DEFORMED AND/OR BROKE OFF DURING THE PASSAGES INTO THE TISSUE. THE NEEDLE POSSIBLY DETACHED FROM THE SUTURE WHEN MODERATE TENSION WAS APPLIED. THE PROCEDURE WAS COMPLETED WITH ANOTHER LIKE SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5595062&pc=GAM | 2210968-2016-08344 |  | Malfunction | ETHICON INC. | 2016-04-21 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break; Detachment Of Device Component | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DEFORMED AND/OR BROKE OFF DURING THE PASSAGES INTO THE TISSUE. THE NEEDLE POSSIBLY DETACHED FROM THE SUTURE WHEN MODERATE TENSION WAS APPLIED. THE PROCEDURE WAS COMPLETED WITH ANOTHER LIKE SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5588324&pc=GAM | 2210968-2016-08310 | 2016-04-01 | Malfunction | ETHICON INC. | 2016-04-20 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE DURING USE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). CONCLUSION CODE: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5591067&pc=GAM | 2210968-2016-08320 |  | Malfunction | ETHICON INC. | 2016-04-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break; Detachment Of Device Component | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DEFORMS. THE NEEDLE BROKE OFF AND WAS BLUNT DURING THE PASSAGES INTO THE TISSUE. THE NEEDLE ALSO DETACHED FROM THE SUTURE WHEN MODERATE TENSION WAS APPLIED. THE PROBLEM WAS SOLVED BY REPLACING THE SUTURE WITH A NEW ONE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5591480&pc=GAM | 2210968-2016-08321 |  | Malfunction | ETHICON INC. | 2016-04-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break; Detachment Of Device Component | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DEFORMED AND/OR BROKE OFF DURING THE PASSAGES INTO THE TISSUE. THE NEEDLE POSSIBLY DETACHED FROM THE SUTURE WHEN MODERATE TENSION WAS APPLIED. THE PROCEDURE WAS COMPLETED WITH ANOTHER LIKE SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5591669&pc=GAM | 2210968-2016-08323 |  | Malfunction | ETHICON INC. | 2016-04-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break; Detachment Of Device Component | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DEFORMED AND/OR BROKE OFF DURING THE PASSAGES INTO THE TISSUE. THE NEEDLE POSSIBLY DETACHED FROM THE SUTURE WHEN MODERATE TENSION WAS APPLIED. THE PROCEDURE WAS COMPLETED WITH ANOTHER LIKE SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5591697&pc=GAM | 2210968-2016-08324 |  | Malfunction | ETHICON INC. | 2016-04-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break; Detachment Of Device Component | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DEFORMED AND/OR BROKE OFF DURING THE PASSAGES INTO THE TISSUE. THE NEEDLE POSSIBLY DETACHED FROM THE SUTURE WHEN MODERATE TENSION WAS APPLIED. THE PROCEDURE WAS COMPLETED WITH ANOTHER LIKE SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5591718&pc=GAM | 2210968-2016-08325 |  | Malfunction | ETHICON INC. | 2016-04-20 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break; Detachment Of Device Component | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DEFORMED AND/OR BROKE OFF DURING THE PASSAGES INTO THE TISSUE. THE NEEDLE POSSIBLY DETACHED FROM THE SUTURE WHEN MODERATE TENSION WAS APPLIED. THE PROCEDURE WAS COMPLETED WITH ANOTHER LIKE SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5584163&pc=GAM | 2210968-2016-08269 | 2016-03-30 | Malfunction | ETHICON INC. | 2016-04-18 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, PRIOR TO USE ON THE PATIENT, IT WAS NOTED THE NEEDLE TIP WAS BENT IN A TRANSVERSE DIRECTION. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: THE ACTUAL DEVICE INVOLVED IN THIS EVENT WAS RETURNED FOR EVALUATION. VISUAL INSPECTION REVEALED THAT THE NEEDLE HAS A DENT IN AT THE AREA WHERE THE BEND STARTS. THE DENT APPEARS TO HAVE BEEN CAUSED DURING MANUFACTURE OF THE NEEDLES. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBER: BATCH JP2320, MFG DATE: 12/07/2015, EXP DATE: 11/30/2020. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBER WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577137&pc=GAM | 2210968-2016-08068 | 2016-02-05 | Malfunction | ETHICON INC. | 2016-04-14 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A HORSE UNDERWENT AN EQUINE CASTRATION PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED FOR SUBCUTICULAR LAYER. THE SUTURE BROKE IN MULTIPLE PLACES POST-OPERATIVELY AT THE SAME DAY AS THE PROCEDURE, DURING RECOVERY, AND THE SPERMATIC CORD WITH SKIN SUPERFICIAL LAYERS WERE RE-CLOSED. IT WAS ALSO REPORTED THAT KNOTS WERE INTACT AND THE SUTURE WAS PLACED AS A CONTINUOUS STITCH. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577188&pc=GAM | 2210968-2016-08065 | 2016-02-05 | Malfunction | ETHICON INC. | 2016-04-14 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT A HORSE UNDERWENT AN EQUINE CASTRATION PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED FOR SUBCUTICULAR LAYER. THE SUTURE BROKE IN MULTIPLE PLACES POST-OPERATIVELY AT THE SAME DAY AS THE PROCEDURE, DURING RECOVERY, AND THE SPERMATIC CORD WITH SKIN SUPERFICIAL LAYERS WERE RE-CLOSED. IT WAS ALSO REPORTED THAT KNOTS WERE INTACT AND THE SUTURE WAS PLACED AS A CONTINUOUS STITCH. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577464&pc=GAM | 2210968-2016-08059 | 2016-02-05 | Malfunction | ETHICON INC. | 2016-04-14 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A HORSE UNDERWENT AN EQUINE CASTRATION PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED FOR SUBCUTICULAR LAYER. THE SUTURE BROKE IN MULTIPLE PLACES POST-OPERATIVELY AT THE SAME DAY AS THE PROCEDURE, DURING RECOVERY, AND THE SPERMATIC CORD WITH SKIN SUPERFICIAL LAYERS WERE RE-CLOSED. IT WAS ALSO REPORTED THAT KNOTS WERE INTACT AND THE SUTURE WAS PLACED AS A CONTINUOUS STITCH. Manufacturer Narrative: (B)(4). CONCLUSION: THE REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED THAT SAMPLES MET ALL SPECIFIED REQUIREMENTS. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5565370&pc=GAM | 2210968-2016-07919 | 2015-12-02 | Malfunction | ETHICON INC. | 2016-04-08 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT MOH'S SURGERY IN THE FACE, ONE INCH BELOW LEFT EYE ON (B)(6) 2015 AND SUTURE WAS USED. THE PATIENT REPORTS EXPERIENCING FLU-LIKE SYMPTOMS IMMEDIATELY FOLLOWING SURGERY AND HAS HAD SYMPTOMS SINCE. THE PATIENT REPORTS SEEING PIECES OF UNDISSOLVED SUTURES IN FACE WITH A STRONG LED FLASHLIGHT. THE PATIENT REPORTS STILL HAVING A NEGATIVE AND OR AN ALLERGIC REACTION TO THE FOREIGN SUBSTANCE THAT HAS BEEN IN FACE FOR APPROXIMATELY FOURTEEN WEEKS AND IS STILL NOT TOTALLY DISSOLVED. THE PATIENT HAS HAD COMPLETE PHYSICAL AND ALL TESTS HAVE BEEN NEGATIVE. THERE HAS BEEN NO MEDICAL OR SURGICAL INTERVENTION REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553793&pc=GAM | 5553793 | 2016-01-21 | Malfunction | ETHICON, INC. | 2016-04-07 | GAM | VICRYL PLUS | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: THE NEEDLE BROKE OFF DURING THE PROCEDURE. AN X-RAY WAS USED TO RETRIEVE BROKEN PART. THE COMPLETE PART WAS RETRIEVED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5532845&pc=GAM | 2210968-2016-07470 |  | Malfunction | ETHICON INC. | 2016-03-29 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT A HORSE UNDERWENT AN EQUINE CASTRATION PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. THE SUTURE BROKE POST-OPERATIVELY AND THE SPERMATIC CORD WITH SKIN SUPERFICIAL LAYERS WERE RE-CLOSED. Manufacturer Narrative: ADDITIONAL NARRATIVE: IT WAS REPORTED THAT THE SUTURE WAS USED FOR SUBCUTICULAR LAYER DURING THE INITIAL PROCEDURE AND THE SUTURE BROKE IN MULTIPLE PLACES AT THE SAME DAY DURING RECOVERY. IT WAS ALSO REPORTED THAT KNOTS WERE INTACT AND THE SUTURE WAS PLACED AS A CONTINUOUS STITCH. REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED THAT SAMPLES MET ALL SPECIFIED REQUIREMENTS. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5535147&pc=GAM | MW5061323 | 2016-03-09 | Malfunction | ETHICON US, LLC | 2016-03-28 | GAM | VICRYL PLUS 3.0 SH 26MM 1/2C TAPER ABSORBABLE SUTURE | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: THE PT HAD A VAGINAL DELIVERY AND SUSTAINED A VAGINAL LACERATION. WHILE THE PRACTITIONER WAS REPAIRING THE LACERATION, THE SUTURE NEEDLE BROKE IN THE PT'S TISSUE. THE PRACTITIONER HAD TO TAKE DOWN THE WHOLE REPAIR AND SEARCH FOR THE NEEDLE IN THE TISSUE. THE NEEDLE WAS FOUND, REMOVED AND THE LACERATION WAS RESUTURED. THE PT TOLERATED THE PROCEDURE WELL. THE SUTURE USED WAS: VICRYL PLUS 3.0 ON SH ((B)(4)) BY ETHICON. UNABLE TO DETERMINE THE LOT NUMBER BECAUSE THE EXTERIOR PACKAGING WAS NOT SAVED AND THERE WERE APPROXIMATELY 10 DIFFERENT LOT NUMBERS IN USE ON (B)(6). THE ROOM MAY HAVE BEEN RESTOCKED WITH A DIFFERENT LOT NUMBER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5527451&pc=GAM | 2210968-2016-07136 | 2016-03-01 | Malfunction | ETHICON INC. | 2016-03-25 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A RECTAL PROLAPSE PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE WAS BENT WHEN SUTURING. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE WAS RETURNED FOR EVALUATION AND IT WAS OBSERVED THAT THERE WERE SEVERAL MARKS ON THE BODY OF THE NEEDLES. BESIDES, THE NEEDLE WAS DAMAGED PROBABLY DUE TO RUN INTO A HARD SURFACE, WHICH CAUSED THE BENT TIP. THEREFORE, THE EVENT CANNOT BE CONFIRMED SINCE THE SAMPLE CONDITIONS SUGGEST AN IMPROPER HANDLING OF THE PRODUCT. Manufacturer Narrative: (B)(4). CONCLUSION CODE: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5527609&pc=GAM | 2210968-2016-07137 | 2016-03-02 | Malfunction | ETHICON INC. | 2016-03-25 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC GASTRIC BYPASS PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. AFTER SEVERAL PASSES THROUGH THICK TISSUE THE TIP OF THE NEEDLE BENT. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4).TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5527708&pc=GAM | 2210968-2016-07139 |  | Malfunction | ETHICON INC. | 2016-03-25 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC SLEEVE GASTRECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE HAS BENT AT THE TIP. THE DOCTOR USED A SECOND NEEDLE AND SUTURE TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5513447&pc=GAM | 2210968-2016-06680 | 2016-02-23 | Malfunction | ETHICON INC. | 2016-03-21 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GASTRIC CANCER PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN OPENED THE PACKAGE IN THE OPERATING ROOM, PRIOR TO USE ON THE PATIENT, THE NEEDLE WAS FOUND PUNCTURED FROM THE PACKAGE. IT WAS REPORTED ONE HOLE FROM THE POPPED OUT NEEDLE. THE PROCEDURE WAS COMPLETED WITH ANOTHER LIKE SUTURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: CONCLUSION: THE ACTUAL DEVICE WAS RETURNED FOR EVALUATION. VISUAL EXAMINATION OBSERVED THE NEEDLE PASSED THROUGH THE PAPER LID, CAUSING A HOLE IN THE PAPER LID. THE OPENED FOIL WAS INSPECTED AND A HOLE WAS FOUND ON THE TOP FOIL AND THERE WAS DAMAGE INSIDE THE FOIL CAUSED BY THE TIP OF THE NEEDLE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5494723&pc=GAM | 2210968-2016-05976 | 2016-02-15 | Malfunction | ETHICON INC. | 2016-03-11 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC SLEEVE GASTRECTOMY PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, WHILE USING ON THE FIELD, THE NEEDLE FOUND BENT AT THE TIP. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL EXAMINATION REVEALED THE POINT OF THE NEEDLE IS BENT AND NO SIGNIFICANT MARKS WERE NOTED IN THE BEND AREA SUGGESTING THIS WAS A MANUFACTURING DEFECT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5494044&pc=GAM | 2210968-2016-05971 |  | Malfunction | ETHICON INC. | 2016-03-10 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORAL/ DENTAL PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. TWO WEEKS FOLLOWING THE PROCEDURE, THE SUTURE HAS NOT DISSOLVED SO THE SURGEON HAD TO TAKE THE SUTURE OFF HIMSELF. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: THE INITIAL PROCEDURE WAS A TOOTH EXTRACTION. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5494056&pc=GAM | 2210968-2016-05972 |  | Malfunction | ETHICON INC. | 2016-03-10 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORAL/ DENTAL PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. TWO WEEKS FOLLOWING THE PROCEDURE, THE SUTURE HAS NOT DISSOLVED SO THE SURGEON HAD TO TAKE THE SUTURE OFF HIMSELF. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: THE INITIAL PROCEDURE WAS A TOOTH EXTRACTION. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5494058&pc=GAM | 2210968-2016-05973 |  | Malfunction | ETHICON INC. | 2016-03-10 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORAL/ DENTAL PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. TWO WEEKS FOLLOWING THE PROCEDURE, THE SUTURE HAS NOT DISSOLVED SO THE SURGEON HAD TO TAKE THE SUTURE OFF HIMSELF. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: THE INITIAL PROCEDURE WAS A TOOTH EXTRACTION. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5475660&pc=GAM | 2210968-2016-05609 | 2016-01-27 | Malfunction | ETHICON INC. | 2016-03-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DOCTOR REPORTED THAT NEEDLE BENT DURING PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: THE INITIAL PROCEDURE WAS A KERATOPLASTY. THE NEEDLE FAILED TO MAINTAIN ITS ORIGINAL BEND IN THE MIDDLE OF THE BODY. Manufacturer Narrative: THE ACTUAL DEVICE INVOLVED IN THIS EVENT WAS RETURNED FOR EVALUATION. THE SAMPLE WAS VISUALLY INSPECTED. DISTORTION WAS OBSERVED IN ONE NEEDLE AT THE MIDDLE OF THE BODY OF THE NEEDLE BY USE OF THE NEEDLE HOLDER. THE SAMPLE CONDITIONS SUGGEST AN IMPROPER HANDLING OF THE SAMPLE. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5467229&pc=GAM | 2210968-2016-05044 | 2016-01-21 | Malfunction | ETHICON INC. | 2016-02-29 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. THE NEEDLE BENT DURING THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: THE ACTUAL DEVICE INVOLVED IN THIS EVENT WAS RETURNED FOR EVALUATION. THE EVALUATION FOUND SEVERAL MARKS ON THE BODY OF THE NEEDLE AND IT WAS BENT AT THE MIDDLE OF THE BODY OF THE NEEDLE BY THE USE OF A NEEDLE HOLDER. THE SAMPLE CONDITION SUGGESTS AN IMPROPER HANDLING OF THE NEEDLE/SUTURE. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5456803&pc=GAM | 2210968-2016-05366 |  | Malfunction | ETHICON INC. | 2016-02-24 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE NEEDLE WAS CRASHED AND COULD NOT PASS THOUGH THE TISSUE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5457802&pc=GAM | 2210968-2016-05389 | 2016-01-27 | Malfunction | ETHICON INC. | 2016-02-24 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A HIP REPLACEMENT PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE SEWING THE CAPSULE, THE NEEDLE BENT AT A FIRST STICH. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. THE NEEDLE /SUTURE WAS VISUALLY INSPECTED AND THE INSPECTION REVEALED THAT THE NEEDLE WAS BENT. THIS OCCURRED DURING THE WINDING PROCESS. THE ZIPPER TRAY (WINDING FORMER) WAS VISUALLY INSPECTED AND THE TRAY HAS SLIGHT MARKS, WHICH INDICATES THAT THE NEEDLE TIP MADE CONTACT WITH THE TRAY WHEN IT WAS PLACED IN THE SLOT OF THE TRAY. Manufacturer Narrative: (B)(4). PHOTO OF THE NEEDLE WAS RECEIVED FOR EVALUATION. VISUAL INSPECTION REVEALED THAT THE NEEDLE WAS BENDING AND NO ASSIGNABLE CAUSE COULD BE DETERMINED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5447597&pc=GAM | 2210968-2016-05206 | 2016-01-28 | Malfunction | ETHICON INC. | 2016-02-19 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EAR SURGERY ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS FOUND TO BEND WHILE SEWING. THERE WERE NO REPORTED PATIENT CONSEQUENCES. ADDITIONAL INFORMATION IS NOT AVAILBALE. Manufacturer Narrative: DATE SENT TO THE FDA: 3/24/2016. THE ACTUAL DEVICE INVOLVED IN THIS EVENT WAS RETURNED FOR EVALUATION. THE EVALUATION FOUND THE NEEDLE HAS A BENT POINT AND SCRATCH MARKS WERE NOTED ON THE BAYONET CLOSE TO THE POINT. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE WAS NOT RETURNED FOR EVALUATION. A SINGLE LOW RESOLUTION PICTURE WAS RETURNED. THE NEEDLE APPEARS TO BE BENT NEAR THE POINT OR TIP BUT THE CAUSE OF THE BENDING CAN NOT BE DETERMINED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5444487&pc=GAM | 2210968-2016-05090 | 2015-10-07 | Malfunction | ETHICON INC. | 2016-02-18 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORAL/ DENTAL PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. TWO WEEKS FOLLOWING THE PROCEDURE, THE SUTURE HAS NOT DISSOLVED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THIS MEDWATCH REPORT IS BEING VOIDED AS IT IS A DUPLICATE OF MEDWATCH REPORT # 2210968-2016-05972. PLEASE SEE MEDWATCH REPORT # 2210968-2016-05972 FOR ALL INFORMATION REGARDING THIS EVENT. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5444509&pc=GAM | 2210968-2016-05091 | 2015-10-07 | Malfunction | ETHICON INC. | 2016-02-18 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORAL/ DENTAL PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. TWO WEEKS FOLLOWING THE PROCEDURE, THE SUTURE HAS NOT DISSOLVED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THIS MEDWATCH REPORT IS BEING VOIDED AS IT IS A DUPLICATE OF MEDWATCH REPORT # 2210968-2016-05973. PLEASE SEE MEDWATCH REPORT # 2210968-2016-05973 FOR ALL INFORMATION REGARDING THIS EVENT. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5445321&pc=GAM | 2210968-2016-05089 | 2015-10-07 | Malfunction | ETHICON INC. | 2016-02-18 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORAL/ DENTAL PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. TWO WEEKS FOLLOWING THE PROCEDURE, THE SUTURE HAS NOT DISSOLVED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THIS MEDWATCH REPORT IS BEING VOIDED AS IT IS A DUPLICATE OF MEDWATCH REPORT # 2210968-2016-05971. PLEASE SEE MEDWATCH REPORT # 2210968-2016-05971 FOR ALL INFORMATION REGARDING THIS EVENT. Manufacturer Narrative: (B)(4). CONCLUSION : TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5445539&pc=GAM | 2210968-2016-05088 | 2015-10-07 | Malfunction | ETHICON INC. | 2016-02-18 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORAL/DENTAL PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. TWO WEEKS FOLLOWING THE PROCEDURE, THE SUTURE WAS NOT DISSOLVED. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: DATE SENT TO THE FDA: 03/18/2016, IT WAS REPORTED THAT NO MEDICAL INTERVENTION WAS APPLIED. IT WAS ALSO REPORTED THAT SUTURE WAS REMOVED. REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL EXAMINATIONS REVEALED NO DEFECTS ON SAMPLES. IFU STATES THAT NORMALLY THE REMOVAL OF THE SUTURE IS NOT REQUIRED. ABSORPTION BEGINS AS A LOSS OF TENSILE STRENGTH FOLLOWED BY A LOSS OF MASS. ALL OF THE ORIGINAL TENSILE STRENGTH IS LOST APPROXIMATELY 10 TO 14 DAYS POST-IMPLANTATION. THE ABSORPTION OF THIS SUTURE OCCURS THEREAFTER AND IS ESSENTIALLY COMPLETE 42 DAYS. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5441432&pc=GAM | 2210968-2016-05018 |  | Malfunction | ETHICON INC. | 2016-02-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE NEEDLE WAS CRASHED AND COULD NOT PASS THOUGH THE TISSUE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: DATE SENT TO THE FDA: 02/25/2016. CORRECTED INFORMATION: IT WAS REPORTED THAT THIS DEVICE EVENT IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5438254&pc=GAM | 2210968-2016-04938 | 2015-12-01 | Malfunction | ETHICON INC. | 2016-02-16 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN INCISIONAL HERNIA REPAIR PROCEDURE IN 2015 AND SUTURE WAS USED. SIX TO EIGHT DAYS AFTER THE PROCEDURE, THE THREAD BROKE IN LENGTH BUT THE KNOTS WERE STILL INTACT. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5438255&pc=GAM | 2210968-2016-04937 | 2015-12-01 | Malfunction | ETHICON INC. | 2016-02-16 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN INCISIONAL HERNIA REPAIR PROCEDURE IN 2015 AND SUTURE WAS USED. SIX TO EIGHT DAYS AFTER THE PROCEDURE, THE THREAD BROKE IN LENGTH BUT THE KNOTS WERE STILL INTACT. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5438274&pc=GAM | 2210968-2016-04939 | 2015-12-01 | Malfunction | ETHICON INC. | 2016-02-16 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN INCISIONAL HERNIA REPAIR PROCEDURE IN 2015 AND SUTURE WAS USED. SIX TO EIGHT DAYS AFTER THE PROCEDURE, THE THREAD BROKE IN LENGTH BUT THE KNOTS WERE STILL INTACT. Manufacturer Narrative: (B)(4) DATE SENT TO THE FDA: 02/16/2016. (B)(4). IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5438276&pc=GAM | 2210968-2016-04941 | 2015-12-01 | Malfunction | ETHICON INC. | 2016-02-16 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN INCISIONAL HERNIA REPAIR PROCEDURE IN 2015 AND SUTURE WAS USED. SIX TO EIGHT DAYS AFTER THE PROCEDURE, THE THREAD BROKE IN LENGTH BUT THE KNOTS WERE STILL INTACT. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5438277&pc=GAM | 2210968-2016-04940 | 2015-12-01 | Malfunction | ETHICON INC. | 2016-02-16 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN INCISIONAL HERNIA REPAIR PROCEDURE IN 2015 AND SUTURE WAS USED. SIX TO EIGHT DAYS AFTER THE PROCEDURE, THE THREAD BROKE IN LENGTH BUT THE KNOTS WERE STILL INTACT. Manufacturer Narrative: CONCLUSION: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL AND FUNCTIONAL EXAMINATIONS WERE PERFORMED AND SAMPLES MET THE SPECIFIED QUALITY REQUIREMENTS. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5447812&pc=GAM | MW5060283 | 2016-02-07 | Malfunction | ETHICON | 2016-02-16 | GAM | VICRYL SUTURE VCP370H SIZE 0, CTX NEEDLE | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: DURING CLOSURE OF THE UTERUS AFTER A C-SECTION, THE TIP/POINT OF THE SUTURE NEEDLE BROKE OFF. AFTER SEARCHING AND XRAYS OF THE AREA, THE TIP WAS LOCATED AND DISCARDED. NO FURTHER INTERVENTIONS WERE NECESSARY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5417420&pc=GAM | 2210968-2016-04277 | 2016-01-20 | Malfunction | ETHICON INC. | 2016-02-08 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT GRAPH VESSEL PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BENT WHEN SUTURING TRAPEZIUS MUSCLE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4): REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL INSPECTION WAS PERFORMED AND NO DEFECTS WERE FOUND. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5406558&pc=GAM | 2210968-2016-03483 | 2016-01-11 | Malfunction | ETHICON INC. | 2016-02-03 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2016 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: CONCLUSION: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED THAT SAMPLES DID NOT MEET THE RESHAPE REQUIREMENT. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5397478&pc=GAM | 2210968-2016-03035 | 2011-07-14 | Malfunction | ETHICON INC. | 2016-01-28 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED FOR SLIPPED MEDIAL RECTUS MUSCLE. ON THE FIRST DAY OF POST-OP, THE SUTURE BROKE, BUT THE KNOT WAS STILL INTACT AND THE CONJUNCTIVA WAS STILL CLOSED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. THIS MEDWATCH REPORT IS IN RESPONSE TO RECEIPT OF A VOLUNTARY MEDWATCH REPORT (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5384989&pc=GAM | 5384989 | 2015-11-18 | Malfunction | ETHICON, INC. | 2016-01-22 | GAM | VICRYL PLUS | Break | No Code Available |  |  | 1 | Event Description: THE OPERATING DOCTOR NOTED THAT NEEDLE FROM SUTURE BROKE WHILE IN THE PATIENT. RADIOLOGY WAS ALREADY PRESENT FOR CHOLANGIOGRAM AND OBTAINED FILMS FOR NEEDLE. AFTER EXAMINING THE FILMS , THE RADIOLOGIST AND DOCTOR VERIFIED THAT ENTIRE NEEDLE (WHICH HAD BROKEN INTO 2 PIECES) WAS REMOVED BY SURGEON.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5369335&pc=GAM | 2210968-2016-01152 | 2015-12-22 | Malfunction | ETHICON INC. | 2016-01-15 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. DURING THE PROCEDURE, TWO PACKAGE LABELS WERE IN THE PACKAGE BEFORE USE. THERE WAS NO ADVERSE CONSEQUENCE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5359949&pc=GAM | 2210968-2016-00579 | 2015-12-21 | Malfunction | ETHICON INC. | 2016-01-12 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2015 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BECAME BENT IN THE MIDDLE WHEN SEWING THE UTERUS. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5357450&pc=GAM | 2210968-2016-00499 | 2015-12-01 | Malfunction | ETHICON INC. | 2016-01-11 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A COW UNDERWENT A CAESAREAN SECTION IN 2015 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN SUTURING THE SKIN, THE NEEDLE BENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE CONSEQUENCES TO THE ANIMAL. Manufacturer Narrative: (B)(4). IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE INVOLVED IN THIS EVENT WAS RETURNED FOR EVALUATION. THE DEVICE WAS VISUALLY AND FUNCTIONALLY EVALUATED. THE DEVICE PERFORMED ACCORDING TO SPECIFICATION. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5357540&pc=GAM | 2210968-2016-00500 | 2015-12-01 | Malfunction | ETHICON INC. | 2016-01-11 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A COW UNDERWENT A CAESAREAN SECTION IN 2015 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN SUTURING THE SKIN, THE NEEDLE BENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE CONSEQUENCES TO THE ANIMAL. Manufacturer Narrative: IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE INVOLVED IN THIS EVENT WAS RETURNED FOR EVALUATION. THE DEVICE WAS VISUALLY AND FUNCTIONALLY EVALUATED. THE DEVICE PERFORMED ACCORDING TO SPECIFICATION. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5344074&pc=GAM | 2210968-2016-00075 |  | Malfunction | ETHICON INC. | 2016-01-05 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient; Device Embedded In Tissue or Plaque | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BREAKAGE OCCURRED AT THE SWAGE PART WHEN SUTURING THE ALIMENTARY CANAL IN THE ABDOMINAL CAVITY. THE BROKEN NEEDLE PIECE FELL INTO THE PATIENT&#38112;BODY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IT WAS REPORTED THAT THE PIECE OF NEEDLE WITH SUTURE WAS RETRIEVED WITHOUT ANY MEASUREMENTS AND NO X-RAY WAS PERFORMED TO DETERMINE THE NEEDLE LOCATION. DURING THE PROCEDURE, THE NEEDLE WAS GRASPED AT SWAGE PART. CONCLUSION: THE ACTUAL SAMPLE WAS RECEIVED FOR EVALUATION. VISUAL EXAMINATION REVEALED THAT THE NEEDLE EXHIBITED AMOUNTS OF INTERGRANULAR FAILURE. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). CONCLUSION CODE: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5314039&pc=GAM | 2210968-2015-20973 |  | Malfunction | ETHICON INC. | 2015-12-21 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY THE CUSTOMER THAT THE DESCRIPTION OF THE NEEDLE ON THE BOX WAS DIFFERENT THAN ASSIGNED TO THE PRODUCT CODE. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION THE ACTUAL SAMPLE WAS PERFORMED. THIS PRODUCT CODE CARRIES A CUTTING EDGE REVERSE NEEDLE. THE GRAPHIC PRINTED ON THE BOX WAS CORRECT. THE LOT # JLZ182 WAS MANUFACTURED ON 10/12/2015 WITH THE CORRECT INFORMATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5314041&pc=GAM | 2210968-2015-20974 |  | Malfunction | ETHICON INC. | 2015-12-21 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY THE CUSTOMER THAT THE DESCRIPTION OF THE NEEDLE ON THE BOX WAS DIFFERENT THAN ASSIGNED TO THE PRODUCT CODE. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION OF THE ACTUAL SAMPLE WAS PERFORMED. THIS PRODUCT CODE CARRIES A CUTTING EDGE REVERSE NEEDLE. THE GRAPHIC PRINTED ON THE BOX WAS CORRECT. THE LOT # JLZ182 WAS MANUFACTURED ON 10/12/2015 WITH THE CORRECT INFORMATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5314043&pc=GAM | 2210968-2015-20975 |  | Malfunction | ETHICON INC. | 2015-12-21 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY THE CUSTOMER THAT THE DESCRIPTION OF THE NEEDLE ON THE BOX WAS DIFFERENT THAN ASSIGNED TO THE PRODUCT CODE. THERE WAS NO PATIENT INVOLVEMENT AND NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION OF THE ACTUAL SAMPLE WAS PERFORMED. THIS PRODUCT CODE CARRIES A CUTTING EDGE REVERSE NEEDLE. THE GRAPHIC PRINTED ON THE BOX WAS CORRECT. THE LOT # JLZ182 WAS MANUFACTURED ON 10/12/2015 WITH THE CORRECT INFORMATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5306833&pc=GAM | 2210968-2015-20761 | 2015-12-04 | Malfunction | ETHICON INC. | 2015-12-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. THE NEEDLE BENT WHILE USING. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO REPORTED ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL EXAMINATION REVEALED THAT THE NEEDLE WAS BENT AND NEEDLE TIP WAS DAMAGED. THE NEEDLE HAD PINCH MARKS ON THE BODY AND ATTACHMENT REGION, WHICH IS HOLLOW AND CONTAINS LESS MATERIAL. THE NEEDLE WAS FUNCTIONALLY TESTED AND ALL THE RESULTS MET THE SPECIFICATIONS. IN ORDER TO AVOID DAMAGING NEEDLE POINTS AND SWAGE AREAS GRASP THE NEEDLE IN AREA ONE-THIRD(1/3) TO ONE HALF (1/2) OF THE DISTANCE FROM THE SWAGED END TO THE POINT . RESHAPING NEEDLES MAY CAUSE THEM TO LOSE STRENGTH AND BE LESS RESISTANT TO BENDING AND BREAKING. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5309139&pc=GAM | 2210968-2015-20830 | 2015-12-05 | Malfunction | ETHICON INC. | 2015-12-17 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. DURING SUTURING THE UTERUS, THE NEEDLE TIP WAS BENT. ANOTHER LIKE A DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: CONCLUSION: THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL EXAMINATION REVEALED THAT THE NEEDLE HAS A BENT/DAMAGED POINT AND MARKS WERE NOTED ON THE NEEDLE. NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5297423&pc=GAM | 2210968-2015-20448 | 2015-11-26 | Malfunction | ETHICON INC. | 2015-12-15 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. DURING PICKING OF THE PACKAGE, A PRINTING FAILURE WAS NOTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: DATE SENT TO THE FDA: 1/19/2016. IT WAS REPORTED THAT THE YELLOW LINE WAS NOTICED ON THE PACKAGE. THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL EXAMINATION REVEALED A SLIGHTLY PALE/DISCOLORED 1/2 INCH STRIPE THAT RUNS THE WIDTH OF THE PACKAGE LABEL. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261746&pc=GAM | 2210968-2015-19460 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY AND THE NEEDLE WAS SOFT AND BENT EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4): THE RETAIN SAMPLES WERE EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET ALL REQUIREMENTS AND SPECIFICATIONS. Manufacturer Narrative: (B)(4) CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261747&pc=GAM | 2210968-2015-19459 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY AND THE NEEDLE WAS SOFT AND BENT EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4): THE RETAIN SAMPLES WERE EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET ALL REQUIREMENTS AND SPECIFICATIONS. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261765&pc=GAM | 2210968-2015-19461 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY AND THE NEEDLE WAS SOFT AND BENT EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: DATE SENT TO THE FDA:02/18/2016. THE RETAIN SAMPLES WERE EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET ALL REQUIREMENTS AND SPECIFICATIONS. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261767&pc=GAM | 2210968-2015-19464 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY AND THE NEEDLE WAS SOFT AND BENT EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: DATE SENT TO THE FDA: 02/18/2016. THE RETAIN SAMPLES WERE EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET ALL REQUIREMENTS AND SPECIFICATIONS. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261768&pc=GAM | 2210968-2015-19465 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY AND THE NEEDLE WAS SOFT AND BENT EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4): RETAIN SAMPLES WERE EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET ALL SPECIFICATIONS. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261769&pc=GAM | 2210968-2015-19463 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY AND THE NEEDLE WAS SOFT AND BENT EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: DATE SENT TO THE FDA: 02/19/2016. THE RETAIN SAMPLES WERE EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET ALL REQUIREMENTS AND SPECIFICATIONS. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261771&pc=GAM | 2210968-2015-19466 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY AND THE NEEDLE WAS SOFT AND BENT EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4): THE RETAIN SAMPLES WERE EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET ALL REQUIREMENTS AND SPECIFICATIONS. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261772&pc=GAM | 2210968-2015-19467 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY AND THE NEEDLE WAS SOFT AND BENT EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4): RETAIN SAMPLES WERE EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET SPECIFICATIONS. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261783&pc=GAM | 2210968-2015-18411 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent; Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY AND THE NEEDLE WAS SOFT AND BENT EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4): THE RETAIN SAMPLES WERE EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET ALL REQUIREMENTS AND SPECIFICATIONS. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 12/02/2015 (B)(4). IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261978&pc=GAM | 2210968-2015-19475 | 2015-11-12 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE DESCRIPTION ON THE BOX OF NEEDLES WAS DIFFERENT THAN ASSIGNED TO THE PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ACTUAL PHOTOGRAPHIC SAMPLE WAS VISUALLY INSPECTED. THE GRAPHIC FOR THE ASSIGNED NEEDLE PRINTED ON THE BOX WAS INCORRECT. THE CORRECTION WAS PERFORMED PREVIOUSLY. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: CONCLUSION: ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION OF THE PICTURES AND THE ACTUAL SAMPLE WAS PERFORMED. THIS PRODUCT CODE MUST CARRY THE CUTTING EDGE REVERSE NEEDLE. THE GRAPHIC PRINTED ON THE BOX WAS CORRECT. THE LOT # JLZ182 WAS MANUFACTURED ON 10/12/2015 WITH THE CORRECT INFORMATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5261979&pc=GAM | 2210968-2015-19476 | 2015-11-12 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE DESCRIPTION ON THE BOX OF NEEDLES WAS DIFFERENT THAN ASSIGNED TO THE PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ACTUAL PHOTOGRAPHIC SAMPLE WAS VISUALLY INSPECTED. THE GRAPHIC FOR THE ASSIGNED NEEDLE PRINTED ON THE BOX WAS INCORRECT. THE CORRECTION WAS PERFORMED PREVIOUSLY. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: ACTUAL SAMPLES WERE RETURNED FOR EVALUATION. VISUAL INSPECTION OF THE PICTURES AND THE ACTUAL SAMPLES WAS PERFORMED. THIS PRODUCT CODE MUST CARRY THE CUTTING EDGE REVERSE NEEDLE. THE GRAPHIC PRINTED ON THE BOXES WAS CORRECT. THE LOT # JLZ182 WAS MANUFACTURED ON 10/12/2015 WITH THE CORRECT INFORMATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4) DATE SENT TO THE FDA: 12/02/2015. (B)(4) CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5262101&pc=GAM | 2210968-2015-19477 | 2015-11-12 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE DESCRIPTION ON THE BOX OF NEEDLES WAS DIFFERENT THAN ASSIGNED TO THE PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ACTUAL PHOTOGRAPHIC SAMPLE WAS VISUALLY INSPECTED. THE GRAPHIC FOR THE ASSIGNED NEEDLE PRINTED ON THE BOX WAS INCORRECT. THE CORRECTION WAS PERFORMED PREVIOUSLY. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: CONCLUSION: ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION OF THE PICTURES AND THE ACTUAL SAMPLE WAS PERFORMED. THIS PRODUCT CODE MUST CARRY THE CUTTING EDGE REVERSE NEEDLE. THE GRAPHIC PRINTED ON THE BOX WAS CORRECT. THE LOT # JLZ182 WAS MANUFACTURED ON 10/12/2015 WITH THE CORRECT INFORMATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 12/02/2015. (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5262128&pc=GAM | 2210968-2015-19481 | 2015-11-12 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE DESCRIPTION ON THE BOX OF NEEDLES WAS DIFFERENT THAN ASSIGNED TO THE PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ACTUAL PHOTOGRAPHIC SAMPLE WAS VISUALLY INSPECTED. THE GRAPHIC FOR THE ASSIGNED NEEDLE PRINTED ON THE BOX WAS INCORRECT. THE CORRECTION WAS PERFORMED PREVIOUSLY. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: CONCLUSION: ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION OF THE PICTURES AND THE ACTUAL SAMPLE WAS PERFORMED. THIS PRODUCT CODE MUST CARRY THE CUTTING EDGE REVERSE NEEDLE. THE GRAPHIC PRINTED ON THE BOX WAS CORRECT. THE LOT # JLZ182 WAS MANUFACTURED ON 10/12/2015 WITH THE CORRECT INFORMATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5262145&pc=GAM | 2210968-2015-19483 | 2015-11-12 | Malfunction | ETHICON INC. | 2015-12-02 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE DESCRIPTION ON THE BOX OF NEEDLES WAS DIFFERENT THAN ASSIGNED TO THE PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ACTUAL PHOTOGRAPHIC SAMPLE WAS VISUALLY INSPECTED. THE GRAPHIC FOR THE ASSIGNED NEEDLE PRINTED ON THE BOX WAS INCORRECT. THE CORRECTION WAS PERFORMED PREVIOUSLY. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: CONCLUSION: ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION OF THE PICTURES AND THE ACTUAL SAMPLE WAS PERFORMED. THIS PRODUCT CODE MUST CARRY THE CUTTING EDGE REVERSE NEEDLE. THE GRAPHIC PRINTED ON THE BOX WAS CORRECT. THE LOT # JLZ182 WAS MANUFACTURED ON 10/12/2015 WITH THE CORRECT INFORMATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5258147&pc=GAM | 2210968-2015-19358 | 2015-11-03 | Malfunction | ETHICON INC. | 2015-12-01 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OOPHORECTOMY ON UNKNOWN DATE AND SUTURE WAS USED TO SUTURE OVARIAN PEDICLES AND CLOSE THE ABDOMINAL MUSCULAR WALL. TWO DAYS LATER, THE RUPTURE WAS OBSERVED ON THE LENGTH OF THE THREAD AND KNOTS ON EXTREMITIES WERE STILL PRESENT. Manufacturer Narrative: (B)(4): TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247156&pc=GAM | 2210968-2015-19105 | 2015-11-05 | Malfunction | ETHICON INC. | 2015-11-24 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FELINE UNDERWENT AN OVARIOHYSTERECTOMY ON (B)(6) 2015 AND SUTURE WAS USED. THE SUTURE KNOTS UNTIED POST-OPERATIVELY. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5247218&pc=GAM | 2210968-2015-19107 | 2015-11-05 | Malfunction | ETHICON INC. | 2015-11-24 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FELINE UNDERWENT AN OVARIOHYSTERECTOMY ON (B)(6) 2015 AND SUTURE WAS USED. THE SUTURE KNOTS UNTIED POST-OPERATIVELY. Manufacturer Narrative: (B)(4). A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL AND FUNCTIONAL INSPECTIONS FOR STRENGTH WERE PERFORMED AND THE SAMPLES MET THE REQUIREMENTS. Manufacturer Narrative: (B)(4) SUTURE KNOTS UNTYING POST-OP. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5233586&pc=GAM | 2210968-2015-18658 | 2015-11-01 | Malfunction | ETHICON INC. | 2015-11-18 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT A HORSE UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. FOLLOWING THE PROCEDURE, THE ANIMAL EXPERIENCED POST-OPERATIVE SUTURE BREAKAGE. Manufacturer Narrative: (B)(4). CONCLUSION CODE: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5231198&pc=GAM | 2210968-2015-16586 | 2015-10-01 | Malfunction | ETHICON INC. | 2015-11-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient; Device Embedded In Tissue or Plaque | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC HERNIA REPAIR PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. DURING SUTURING THE FASCIA, THE NEEDLE BROKE WITHOUT BEND. THE BROKEN NEEDLE WAS RETAINED IN THE PATIENT. IT IS UNKNOWN IF ANY INTERVENTION PERFORMED TO REMOVE THE NEEDLE PIECE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC HERNIA REPAIR PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. DURING SUTURING THE FASCIA, THE NEEDLE BROKE WITHOUT BEND. THE BROKEN NEEDLE PIECE WAS REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE AND NO ADDITIONAL INCISION WAS NECESSARY. Manufacturer Narrative: (B)(4). REPRESENTATIVE SAMPLES WERE RETURNED AND EVALUATED. VISUAL AND FUNCTIONAL EXAMINATIONS REVEALED NO DEFECTS AND SAMPLES MET THE SPECIFIED QUALITY REQUIREMENTS. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5225658&pc=GAM | 2210968-2015-18394 | 2015-10-30 | Malfunction | ETHICON INC. | 2015-11-16 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Bent | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CESAREAN SECTION ON (B)(6) 2015 AND SUTURE WAS USED. WHEN REMOVING SUTURE FROM THE PACKAGING, THE NURSE NOTICED THAT THE NEEDLE LOOKED TO BE BENT. IT APPEARED ALMOST BARBED OR BROKEN. THE REMAINING SUTURE STRAND WAS NOT REMOVED FROM PACKAGING. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5202232&pc=GAM | 2210968-2015-16115 | 2015-10-12 | Malfunction | ETHICON INC. | 2015-11-04 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT A HORSE UNDERWENT A SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. ON (B)(6) 2015, THE ANIMAL EXPERIENCED POST-OPERATIVE SUTURE BREAKAGE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5190098&pc=GAM | 2210968-2015-16803 | 2015-10-05 | Malfunction | ETHICON INC. | 2015-10-29 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND ABSORBABLE SUTURE WAS USED. DURING THE PROCEDURE, THE USER SAW THAT THE INNER PACKAGING DID NOT MATCH THE OUTER PACKAGING. ADDITIONAL INFORMATION WAS REQUESTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: FOLLOW-UP INFORMATION WAS RECEIVED THAT STATED THE OUTER PACKAGE (PAPER-PLASTIC ENVELOPE), THE INNER PACKAGE (PAPER FOLDER), AND THE PRODUCT WERE ALL RELATED TO PRODUCT CODE J341. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5152824&pc=GAM | 2210968-2015-15037 | 2015-08-27 | Malfunction | ETHICON INC. | 2015-10-15 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. PRIOR TO USING ON THE PATIENT, A CRACK WAS FOUND ON THE SURFACE SIDE OF THE OUTER PACKAGE BEFORE OPENING. THE DEVICE WAS NOT USED. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT. Manufacturer Narrative: (B)(4). CONCLUSION: THE ACTUAL DEVICE WAS RETURNED FOR EVALUATION. THE FOIL CAVITY WAS DAMAGED AND TORN ON THE BACK OF THE PACKET. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5112536&pc=GAM | 5112536 | 2015-04-22 | Malfunction | ETHICON | 2015-09-30 | GAM | VICRYL | Break | Device Embedded In Tissue or Plaque |  |  | 1 | Event Description: WHEN BEING SUTURED, SUTURE BROKE FROM NEEDLE AND BECAME EMBEDDED IN PT WOUND. DOCTOR TRIED TO REMOVE IT BUT WAS UNABLE TO FIND NEEDLE. FLUOROSCOPY WAS USED WITHOUT RESULT. PATIENT TO COME IN THE AM AND HAVE FOREIGN BODY REMOVED BY RADIOLOGY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5112567&pc=GAM | 5112567 | 2015-09-03 | Malfunction | ETHICON ENDO-SURGERY LLC. | 2015-09-30 | GAM | VICRYL SUTURE | Corroded; Device Packaging Compromised | No Patient Involvement |  |  | 1 | Event Description: MALE PATIENT HAVING ORTHOPEDIC SURGERY FOR END-STAGE ARTHRITIS OF (L) KNEE TOTAL ARTHROPLASTY: AT INITIAL COUNTING OF SHARPS IT WAS NOTED THAT A PACKAGE OF 0 VICRYL SUTURES HAD RUST ON ONE OF THE NEEDLES. PACKAGE REMOVED AND A NEW ONE OPENED. THERE WAS ZERO PATIENT INVOLVEMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5083590&pc=GAM | 5083590 | 2015-09-03 | Malfunction | ETHICON | 2015-09-17 | GAM | VICRYL | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: DURING USE OF SUTURE NEEDLE IT BROKE AND BOTH PIECES RECOVERED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5068391&pc=GAM | 2210968-2015-13016 |  | Malfunction | ETHICON INC. | 2015-09-10 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A DENTAL GRAFTS PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. ONE OR TWO DAYS FOLLOWING THE PROCEDURE, THE PATIENT EXPERIENCED SUTURE BREAKAGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5068455&pc=GAM | 2210968-2015-13024 |  | Malfunction | ETHICON INC. | 2015-09-10 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A DENTAL GRAFTS PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. ONE OR TWO DAYS FOLLOWING THE PROCEDURE, THE PATIENT EXPERIENCED SUTURE BREAKAGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5068554&pc=GAM | 2210968-2015-13032 |  | Malfunction | ETHICON INC. | 2015-09-10 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A BONE GRAFT PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. ONE OR TWO DAYS FOLLOWING THE PROCEDURE, THE PATIENT EXPERIENCED SUTURE BREAKAGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. THEY WERE VISUALLY AND FUNCTIONALLY EXAMINED FOR STRENGTH AND THEY MET REQUIREMENTS. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5041841&pc=GAM | 2210968-2015-12425 | 2015-07-11 | Malfunction | ETHICON INC. | 2015-08-31 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CAESAREAN-SECTION ON (B)(6) 2015 AND SUTURE WAS USED. BEFORE USE ON THE PATIENT, IT WAS REPORTED THAT THE PACKAGING WAS COMPROMISED. THERE WAS NO PINHOLE DETECTED IN THE PACKAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE RETAIN SAMPLES WERE EVALUATED. VISUAL EXAMINATION WAS REVEALED NO DEFECTS. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5019405&pc=GAM | 2210968-2015-11969 | 2015-07-17 | Malfunction | ETHICON INC. | 2015-08-20 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Incomplete or Missing Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CAESAREAN SECTION ON (B)(6) 2015 AND SUTURE WAS USED. WHEN THE PACKAGE WAS OPENED, BEFORE USE ON THE PATIENT, THE PAPER WITH THE WRITTEN KINDS OF SUTURE WAS MISSING ON THE BASE IN THE PACKAGE. THE SUTURE WAS UNABLE TO BE DISTINGUISHED FROM OTHERS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: CONCLUSION: THE ACTUAL DEVICE INVOLVED IN THIS EVENT WAS RETURNED FOR EVALUATION. THE TRAY WAS VISUALLY EXAMINED. THE INSERTION TAB OF THE PAPER COVER WAS NOT OBSERVED IN THE TRAY. THIS CONDITION INDICATES THAT THE PAPER LID WAS NOT SECURED IN THE TRAY. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5008311&pc=GAM | 2210968-2015-11492 | 2015-07-22 | Malfunction | ETHICON INC. | 2015-08-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. BEFORE THE PACKAGE WAS OPENED, A HOLE SEEMED TO BE STUCK ON THE PACKAGE. THE DEVICE WAS NOT USED ON THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4965068&pc=GAM | 2210968-2015-10803 | 2015-07-01 | Malfunction | ETHICON INC. | 2015-08-04 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. APPROXIMATELY 4 DAYS FOLLOWING THE PROCEDURE, THE SUTURE WAS BROKEN AND IT WAS SUSPECTED THAT THE SUTURE DEGRADED OR ABSORBED IN ADVANCE. NO MORE ACTION HAS BEEN TAKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: DATE SENT TO THE FDA: 08/31/2015, ADDITIONAL NARRATIVE: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION AT THE END OF (B)(6), BEGINNING OF (B)(6) 2015. THE SUTURE WAS USED ON THE PERINEUM. IT WAS REPORTED THAT NO ADDITIONAL PROCEDURE WAS PERFORMED AND THE SURGEON PUT SOME MEDICINE ON THE PATIENT'S WOUND.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4965074&pc=GAM | 2210968-2015-10802 | 2015-07-01 | Malfunction | ETHICON INC. | 2015-08-04 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. APPROXIMATELY 4 DAYS FOLLOWING THE PROCEDURE, THE SUTURE WAS BROKEN AND IT WAS SUSPECTED THAT THE SUTURE DEGRADED OR ABSORBED IN ADVANCE. NO MORE ACTION HAS BEEN TAKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: DATE SENT TO THE FDA: 08/31/2015. ADDITIONALNARRATIVE: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION AT THE END OF (B)(6), BEGINNING OF (B)(6) 2015. THE SUTURE WAS USED ON THE PERINEUM. IT WAS REPORTED THAT NO ADDITIONAL PROCEDURE WAS PERFORMED AND THE SURGEON PUT SOME MEDICINE ON THE PATIENT'S WOUND.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4965138&pc=GAM | 2210968-2015-10807 | 2015-07-01 | Malfunction | ETHICON INC. | 2015-08-04 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. APPROXIMATELY 4 DAYS FOLLOWING THE PROCEDURE, THE SUTURE WAS BROKEN AND IT WAS SUSPECTED THAT THE SUTURE DEGRADED OR ABSORBED IN ADVANCE. NO MORE ACTION HAS BEEN TAKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: DATE SENT TO THE FDA: 08/31/2015. ADDITIONAL NARRATIVE: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION AT THE END OF (B)(6), BEGINNING OF (B)(6) 2015. THE SUTURE WAS USED ON THE PERINEUM. IT WAS REPORTED THAT NO ADDITIONAL PROCEDURE WAS PERFORMED AND THE SURGEON PUT SOME MEDICINE ON THE PATIENT'S WOUND.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4965141&pc=GAM | 2210968-2015-10805 | 2015-07-01 | Malfunction | ETHICON INC. | 2015-08-04 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. APPROXIMATELY 4 DAYS FOLLOWING THE PROCEDURE, THE SUTURE WAS BROKEN AND IT WAS SUSPECTED THAT THE SUTURE DEGRADED OR ABSORBED IN ADVANCE. NO MORE ACTION HAS BEEN TAKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: DATE SENT TO THE FDA: 08/31/2015. ADDITIONAL NARRATIVE: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION AROUND THE END OF (B)(6), BEGINNING OF (B)(6) 2015. THE SUTURE WAS USED ON THE PERINEUM. IT WAS REPORTED THAT NO ADDITIONAL PROCEDURE WAS PERFORMED AND THE SURGEON PUT SOME MEDICINE ON THE PATIENT'S WOUND.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4945319&pc=GAM | 4945319 | 2015-01-08 | Malfunction | ETHICON, INC. | 2015-07-24 | GAM | VICRYL (POLYGLACTIN 910) | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: SUTURE NEEDLE TIP BROKE DURING SUTURING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4894989&pc=GAM | 2210968-2015-07937 | 2015-06-15 | Malfunction | ETHICON INC. | 2015-07-07 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT DURING UNKNOWN PROCEDURE, THE PACKAGE OF THE DEVICE WAS FOUND BROKEN BEFORE USE ON THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED AT THIS TIME. Manufacturer Narrative: A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4911572&pc=GAM | 4911572 | 2015-05-15 | Malfunction | ETHICON, INC | 2015-06-25 | GAM | VICRYL | Break; Detachment Of Device Component | No Information |  |  | 1 | Event Description: A 2MM TIP OF 26MM SH NEEDLE ASSOCIATED WITH ETHICON COATED VICRYL 2-0 SUTURE BROKE OFF WHEN NEEDLE DRIVER (HAND TOOL) WAS PLACED ON SHARP END TO PROTECT SURGEON. THIS IS A CUSTOMARY TECHNIQUE BUT DOCTOR HAS BEEN SEEING MORE NEEDLES BREAKING THAN HE USED TO SEE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4857079&pc=GAM | 2210968-2015-07306 | 2015-05-27 | Malfunction | ETHICON INC. | 2015-06-19 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT A DOG UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE. DURING FOLLOW-UP EXAMINATION, IT WAS NOTED THAT THE SUTURE WAS BROKEN. Manufacturer Narrative: (B)(4). CONCLUSION - TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. (B)(4). Manufacturer Narrative: CONCLUSION: REPRESENTATIVE SAMPLES WERE RETURNED. THE VISUAL AND FUNCTIONAL EVALUATION WAS PERFORMED. THE SAMPLES MET REQUIREMENTS AND NO DEFECTS WERE FOUND.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4857198&pc=GAM | 2210968-2015-07307 | 2015-05-27 | Malfunction | ETHICON INC. | 2015-06-19 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT A DOG UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE. DURING FOLLOW-UP EXAMINATION, IT WAS NOTED THAT THE SUTURE WAS BROKEN. Manufacturer Narrative: (B)(4). CONCLUSION - TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. (B)(4). Manufacturer Narrative: CONCLUSION: REPRESENTATIVE SAMPLES WERE RETURNED. THE VISUAL AND FUNCTIONAL EVALUATION WAS PERFORMED. THE SAMPLES MET REQUIREMENTS AND NO DEFECTS WERE FOUND.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4850751&pc=GAM | 2210968-2015-07188 | 2015-05-25 | Malfunction | ETHICON INC. | 2015-06-17 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Other (for use when an appropriate device code cannot be identified) | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO USE ON A PATIENT, AN ISSUE WAS ENCOUNTERED WITH BARCODE SCANNING FOR THE SUTURE, WHEN THE BARCODE WAS PRINTED IN THE GRAY AREA. A SHC400 BARCODE SCANNER WAS USED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4850770&pc=GAM | 2210968-2015-07189 | 2015-05-25 | Malfunction | ETHICON INC. | 2015-06-17 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Other (for use when an appropriate device code cannot be identified) | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO USE ON A PATIENT, AN ISSUE WAS ENCOUNTERED WITH BARCODE SCANNING FOR THE SUTURE, WHEN THE BARCODE WAS PRINTED IN THE GRAY AREA. A SHC400 BARCODE SCANNER WAS USED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. Manufacturer Narrative: CONCLUSION: PHOTOS WERE SUBMITTED FOR VISUAL EVALUATION AND REVEALED NO DEFECT COULD BE IDENTIFIED IN THE ENVELOPE OR THE BARCODE PRINTING. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4850849&pc=GAM | 2210968-2015-07190 | 2015-05-25 | Malfunction | ETHICON INC. | 2015-06-17 | GAM | VICRYL RAPIDE (POLYGLACTIN 910) SUTURE | Other (for use when an appropriate device code cannot be identified) | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO USE ON A PATIENT, AN ISSUE WAS ENCOUNTERED WITH BARCODE SCANNING FOR THE SUTURE, WHEN THE BARCODE WAS PRINTED IN THE GRAY AREA. A SHC400 BARCODE SCANNER WAS USED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4841683&pc=GAM | 2210968-2015-07042 | 2015-05-13 | Malfunction | ETHICON INC. | 2015-06-12 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT AN ANIMAL UNDERWENT A CATARACT PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE APPROXIMATELY ONE WEEK POST-OPERATIVELY. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4827323&pc=GAM | 2210968-2015-06877 | 2015-05-13 | Malfunction | ETHICON INC. | 2015-06-09 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT A CATARACT PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. ON THE ONE DAY RECHECK, IT WAS NOTED THE SUTURE STRAND WAS BROKEN BUT THE KNOT WAS INTACT. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4924410&pc=GAM | 4924410 | 2015-01-08 | Malfunction | ETHICON, INC. | 2015-06-09 | GAM | VICRYL | Break; Component Falling | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: THE SUTURE NEEDLE BROKE AND WE RETRIEVED THE TIP FROM THE PATIENT WOUND.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4815527&pc=GAM | 2210968-2015-06651 | 2015-03-26 | Malfunction | ETHICON INC. | 2015-06-03 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Unsealed Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. DURING THE PROCEDURE, THE CIRCULATING NURSE WHO WAS OPENING THIS SUTURE ON TO THE SCRUB TABLE, NOTICED THAT THE FOIL PACKAGING HAD BEEN COMPROMISED AND THAT IT WAS OPEN AT THE END. THE NURSE COULD SEE THAT AIR WOULD HAVE GOT IN TO THE PACKET. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. (B)(6). Manufacturer Narrative: CONCLUSION: ACTUAL DEVICE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION WAS PERFORMED AND REVEALED THAT THE INNER FOLDER WAS TRAPPED IN THE SEAL OF THE FOIL PACK.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4817679&pc=GAM | 4817679 | 2015-04-29 | Malfunction | ETHICON, INC | 2015-05-22 | GAM | VICRYL | Break; Detachment Of Device Component | Device Embedded In Tissue or Plaque |  |  | 1 | Event Description: WHILE CLOSING THE HIP, THE TIP OF A 2 VICRYL NEEDLE BROKE OFF IN THE BONE. AN X-RAY WAS CONDUCTED TO ENSURE THE NEEDLE STILL RESIDED IN THE BONE. REMOVAL OF THE PIECE WAS DETERMINED TO BE TOO RISKY, AND IT WAS DECIDED IT WAS SAFER TO LEAVE IT IN THE BONE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4705469&pc=GAM | 4705469 | 2015-03-24 | Malfunction | ETHICON INC. | 2015-04-09 | GAM | VICRYL | Break; Difficult to Insert; Detachment of Device or Device Component | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: PHYSICIAN ASSISTANT (PA) WAS CLOSING LAPAROSCOPIC SITES AND WHEN THEY PULLED THE SUTURE THROUGH THE SKIN THE NEEDLE BROKE OFF THREAD AND WAS STUCK INSIDE SUBCUTANEOUS LAYER. DR WAS CALLED BACK INTO OR.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4637907&pc=GAM | 4637907 | 2015-01-01 | Malfunction | ETHICON, INC | 2015-03-19 | GAM | VICRYL | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: SURGEON WAS SUTURING PATIENT'S INCISION WHEN TIP OF SUTURE NEEDLE BROKE OFF REST OF NEEDLE. SURGEON STATES SHE SAW IT IN THE NEEDLE HOLDER AND THEN IT WAS GONE. SURGEON STATES THAT IT COULD NOT HAVE FALLEN IN THE PATIENT BUT DID A WOUND CHECK TO VERIFY. SHE DID NOT SEE IT AND CONTINUED TO CLOSE. CIRCULATOR STATED THAT WE NEED TO TAKE AN X-RAY PER PROTOCOL. SURGEON REPLIED EVEN IF FOUND SHE WOULD NOT GO IN AND RETRIEVE IT BECAUSE THE RISK OUTWEIGHED THE BENEFITS AND TO TAKE THE PICTURE IN RECOVERY DURING THE POST-OP X-RAY. NO RADIOPAQUE FOREIGN BODY IDENTIFIED IN POST OP X-RAY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4637907&pc=GAM | 4637907 | 2015-01-01 | Malfunction | ETHICON, INC | 2015-03-19 | GAM | VICRYL | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: SURGEON WAS SUTURING PATIENT'S INCISION WHEN TIP OF SUTURE NEEDLE BROKE OFF REST OF NEEDLE. SURGEON STATES SHE SAW IT IN THE NEEDLE HOLDER AND THEN IT WAS GONE. SURGEON STATES THAT IT COULD NOT HAVE FALLEN IN THE PATIENT BUT DID A WOUND CHECK TO VERIFY. SHE DID NOT SEE IT AND CONTINUED TO CLOSE. CIRCULATOR STATED THAT WE NEED TO TAKE AN X-RAY PER PROTOCOL. SURGEON REPLIED EVEN IF FOUND SHE WOULD NOT GO IN AND RETRIEVE IT BECAUSE THE RISK OUTWEIGHED THE BENEFITS AND TO TAKE THE PICTURE IN RECOVERY DURING THE POST-OP X-RAY. NO RADIOPAQUE FOREIGN BODY IDENTIFIED IN POST OP X-RAY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4590978&pc=GAM | 2210968-2015-02780 | 2015-02-18 | Malfunction | ETHICON INC. | 2015-03-11 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Unsealed Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2015, IT WAS DIFFICULT TO OPEN THE PACKAGE BECAUSE THE HEAT SEALING PART WAS PASTED TO THE END OF THE PACKAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH HH5280, MFG DATE: 07/31/2014, EXP DATE: 07/31/2019. BATCH HH6728, MFG DATE: 07/31/2014, EXP DATE: 07/31/2019. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: CONCLUSION: REPRESENTATIVE SAMPLE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION REVEALED THE PACKAGE HAD THE TEAR OPEN FLAP OF THE FOIL PACKAGE STUCK TOGETHER. THE SAMPLE APPEARS TO HAVE BEEN RESTERILIZED USING AN AUTOCLAVE. THIS WOULD CAUSE THE PACKAGE TO HEAT UP SWELL AND THEN DEFLATE. THERE WERE NO MANUFACTURING DEFECTS NOTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4591070&pc=GAM | 2210968-2015-02781 | 2015-02-18 | Malfunction | ETHICON INC. | 2015-03-11 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Unsealed Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2015, IT WAS DIFFICULT TO OPEN THE PACKAGE BECAUSE THE HEAT SEALING PART WAS PASTED TO THE END OF THE PACKAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH HH5280, MFG DATE: 07/31/2014, EXP DATE: 07/31/2019. BATCH HH6728, MFG DATE: 07/31/2014, EXP DATE: 07/31/2019. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: REPRESENTATIVE SAMPLE WAS RETURNED FOR EVALUATION. VISUAL INSPECTION REVEALED THE PACKAGE HAD THE TEAR OPEN FLAP OF THE FOIL PACKAGE STUCK TOGETHER. THE SAMPLE APPEARS TO HAVE BEEN RESTERILIZED USING AN AUTOCLAVE. THIS WOULD CAUSE THE PACKAGE TO HEAT UP SWELL AND THEN DEFLATE. THERE WERE NO MANUFACTURING DEFECTS NOTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4591151&pc=GAM | 2210968-2015-02783 | 2015-02-18 | Malfunction | ETHICON INC. | 2015-03-11 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Unsealed Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2015, IT WAS DIFFICULT TO OPEN THE PACKAGE BECAUSE THE HEAT SEALING PART WAS PASTED TO THE END OF THE PACKAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: CONCLUSION: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL INSPECTION WAS PERFORMED. TWO OF THE REPRESENTATIVE SAMPLE PACKAGES HAD THE TEAR OPEN FLAPS OF THE FOIL PACKAGE STUCK TOGETHER. THE SAMPLES APPEAR TO HAVE BEEN RESTERILIZED USING AN AUTOCLAVE. THIS WOULD CAUSE THE PACKAGES TO HEAT UP SWELL AND THEN DEFLATE. THERE WERE NO MANUFACTURING DEFECTS NOTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4591646&pc=GAM | 2210968-2015-02792 | 2015-02-18 | Malfunction | ETHICON INC. | 2015-03-11 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Unsealed Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2015, IT WAS DIFFICULT TO OPEN THE PACKAGE BECAUSE THE HEAT SEALING PART WAS PASTED TO THE END OF THE PACKAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: CONCLUSION: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL INSPECTION WAS PERFORMED. TWO OF THE REPRESENTATIVE SAMPLE PACKAGES HAD THE TEAR OPEN FLAPS OF THE FOIL PACKAGE STUCK TOGETHER. THE SAMPLES APPEAR TO HAVE BEEN RESTERILIZED USING AN AUTOCLAVE. THIS WOULD CAUSE THE PACKAGES TO HEAT UP SWELL AND THEN DEFLATE. THERE WERE NO MANUFACTURING DEFECTS NOTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4564380&pc=GAM | 2210968-2015-02417 | 2015-02-07 | Malfunction | ETHICON INC. | 2015-03-03 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break; Other (for use when an appropriate device code cannot be identified) | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SURGICAL PROCEDURE FOR GLAUCOMA ON (B)(6) 2015 AND SUTURE WAS USED. WITHIN A WEEK AFTER THE SURGERY, THE SUTURE BROKE OFF AND KNOTS WERE BECOMING FRAGILE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4536834&pc=GAM | 2210968-2015-01897 | 2015-01-30 | Malfunction | ETHICON INC. | 2015-02-23 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT PHARYNGEAL FLAP PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. ON (B)(6) 2015, THE PATIENT EXPERIENCED POST-OPERATIVE SUTURE BREAKAGE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH HL6479, BATCH HMK279. Manufacturer Narrative: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH HL6479 MFG DATE: 10/16/2014, EXP DATE: 07/31/2019. BATCH HMK279 MFG DATE: 11/24/2014, EXP DATE: 07/31/2019. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4536837&pc=GAM | 2210968-2015-01893 | 2015-01-30 | Malfunction | ETHICON INC. | 2015-02-23 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT PHARYNGEAL FLAP PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. ON (B)(6) 2015, THE PATIENT EXPERIENCED POST-OPERATIVE SUTURE BREAKAGE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH HG6189, BATCH HLK377. Manufacturer Narrative: ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH HG6189, EXP DATE: 01/31/2019, BATCH HLK377, EXP DATE: 07/31/2019. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4536913&pc=GAM | 2210968-2015-01900 | 2015-01-30 | Malfunction | ETHICON INC. | 2015-02-23 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT PHARYNGEAL FLAP PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. ON (B)(6) 2015, THE PATIENT EXPERIENCED POST-OPERATIVE SUTURE BREAKAGE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). CONCLUSION: NO CONCLUSION CAN BE DRAWN AT THIS TIME. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A SUPPLEMENTAL 3500A FORM WILL BE SUBMITTED ACCORDINGLY. ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH HC2369, BATCH EM2217. Manufacturer Narrative: IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS FOR THE POSSIBLE BATCH NUMBERS WAS CONDUCTED AND THE BATCHES MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4521126&pc=GAM | 2210968-2015-01627 |  | Malfunction | ETHICON INC. | 2015-02-17 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT AN ANIMAL UNDERWENT OVARIECTOMY FOR OVARIAN PEDICLE LIGATION AND CLOSURE OF ABDOMINAL MUSCLE ON AN UNKNOWN DATE AND SUTURE WAS USED. APPROXIMATELY TWO DAYS FOLLOWING THE PROCEDURE, THE ANIMAL EXPERIENCED POST-OPERATIVE SUTURE BREAKAGE. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA. Manufacturer Narrative: (B)(4): REPRESENTATIVE SAMPLE WAS RETURNED FOR EVALUATION. VISUAL AND FUNCTIONAL EVALUATION WAS PERFORMED AND THE SAMPLE MET THE REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4494550&pc=GAM | 2210968-2015-01227 | 2014-11-01 | Malfunction | ETHICON INC. | 2015-02-09 | GAM | VICRYL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON (B)(6) 2014 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). CONCLUSION: TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. Manufacturer Narrative: CONCLUSION: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. VISUAL AND FUNCTIONAL INSPECTIONS WERE PERFORMED AND THE SAMPLES MET THE REQUIREMENTS. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4494579&pc=GAM | 2210968-2015-01228 | 2014-11-01 | Malfunction | ETHICON INC. | 2015-02-09 | GAM | VICRYL PLUS ANTIBACTERIAL (POLYGLACTIN 910) SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON (B)(6) 2014 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFO DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
The MAUDE database houses medical device reports submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers. Each year | the FDA receives several hundred thousand medical device reports (MDRs) of suspected device-associated deaths | serious injuries and malfunctions. The FDA uses MDRs to monitor device performance | detect potential device-related safety issues | and contribute to benefit-risk assessments of these products. The MAUDE database houses MDRs submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers.
Although MDRs are a valuable source of information | this passive surveillance system has limitations | including the potential submission of incomplete | inaccurate | untimely | unverified | or biased data. In addition | the incidence or prevalence of an event cannot be determined from this reporting system alone due to potential under-reporting of events and lack of information about frequency of device use. Because of this | MDRs comprise only one of the FDA's several important postmarket surveillance data sources.
MDR data alone cannot be used to establish rates of events | evaluate a change in event rates over time or compare event rates between devices. The number of reports cannot be interpreted or used in isolation to reach conclusions about the existence | severity | or frequency of problems associated with devices. Confirming whether a device actually caused a specific event can be difficult based solely on information provided in a given report. Establishing a cause-and-effect relationship is especially difficult if circumstances surrounding the event have not been verified or if the device in question has not been directly evaluated.
MAUDE data does not represent all known safety information for a reported medical device and should be interpreted in the context of other available information when making device-related or treatment decisions. Variations in trade | product | and company names affect search results. Searches only retrieve records that contain the search term(s) provided by the requester.Submission of a medical device report and the FDA's release of that information is not necessarily an admission that a product | user facility | importer | distributor | manufacturer | or medical personnel caused or contributed to the event. Certain types of report information are protected from public disclosure under the Freedom of Information Act (FOIA).
If a report contains trade secret or confidential business information | that text is replaced by "(b)(4)". If a report contains personnel or medical files information | that text is replaced by "(b)(6)". The designations "(b)(4)" and "(b)(6)" refer to the exemptions in the FOIA. For example | "(b)(4)" may be found in place of the product's composition and "(b)(6)" may be found in place of a patient's age. MAUDE is updated monthly and the search page reflects the date of the most recent update. The FDA seeks to include all reports received prior to the update but the inclusion of some reports may be delayed.
Manufacturers and importers must submit reports when they become aware of information that reasonably suggests that one of their marketed devices may have caused or contributed to a death or serious injury or has malfunctioned and the malfunction of the device or a similar device that they market would be likely to cause or contribute to a death or serious injury if the malfunction were to recur. Manufacturers must send reports of such deaths | serious injuries and malfunctions to the FDA. Importers must send reports of deaths and serious injuries to the FDA and the manufacturer | and reports of malfunctions to the manufacturer.
Device user facilities include hospitals | outpatient diagnostic or treatment facilities | nursing homes and ambulatory surgical facilities. Device user facilities must submit reports when they become aware of information that reasonably suggests that a device may have caused or contributed to a death or serious injury of a patient in their facility. Death reports must be sent to the FDA and the manufacturer | if known. Serious injury reports must be sent to the manufacturer or to the FDA | if the manufacturer is not known.