## Sheet1
| Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8735291&pc=GAM | 2210968-2019-83261 | 2019-06-02 | Malfunction | ETHICON INC. | 2019-06-26 | GAM | VICRYL POLYGLACTIN 910 2/0 SGLE ARMED CT PLUSSUTURE | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE TORE DURING THE KNOT FORMATION WHILE CLOSING THE UTERUS. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8735294&pc=GAM | 2210968-2019-83262 | 2019-06-02 | Malfunction | ETHICON INC. | 2019-06-26 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE TORE DURING THE KNOT FORMATION WHILE CLOSING THE UTERUS. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MLBDWQR0 NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8735365&pc=GAM | 2210968-2019-83305 |  | Malfunction | ETHICON INC. | 2019-06-26 | GAM | VICRYL RAP UND 75CM M2 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED AWAY FROM THE THREAD. THE NEEDLE MAY HAVE FALLEN INSIDE THE ABDOMEN OF THE PATIENT AND WAS SEARCHED BY SCOPE. THE PROCEDURE WAS FINISHED WITH ANOTHER SUTURE WITHOUT ISSUE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE NEEDLE RETAINED IN PATIENT TISSUE? IF YES, IN WHAT PATIENT STRUCTURE/TISSUE? WHAT PRODUCT CODE AND LOT NUMBER WERE RETAINED IN THE PATIENT? IF NEEDLE IS RETAINED, ARE THERE PLANS TO REMOVE THE NEEDLE? IF NEEDLE IS NOT RETAINED, WAS THE NEEDLE REMOVED DURING THE INITIAL PROCEDURE OR DID THE PATIENT UNDERGO A SECOND SURGICAL PROCEDURE TO REMOVE THE NEEDLE? WHAT IS THE PATIENT¿S CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736896&pc=GAM | 2210968-2019-83319 | 2019-06-03 | Malfunction | ETHICON INC. | 2019-06-26 | GAM | VICRYL* UND 4-0 70CM (1)SC-20 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT PLASTIC SURGERY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. AT THE MOMENT OF KNOTTING AND PULLING THE SUTURE, THE THREAD BROKE. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM DIFFERENT LOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736899&pc=GAM | 2210968-2019-83318 | 2019-06-03 | Malfunction | ETHICON INC. | 2019-06-26 | GAM | VICRYL* UND 4-0 70CM (1)SC-20 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT PLASTIC SURGERY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. AT THE MOMENT OF KNOTTING AND PULLING THE SUTURE, THE THREAD BROKE. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM DIFFERENT LOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8730833&pc=GAM | 2210968-2019-83233 | 2019-05-31 | Malfunction | ETHICON INC. | 2019-06-25 | GAM | VICRYL PLUS SUTURE 27"(70CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CAESAREAN SECTION ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS BROKEN WHEN PULLING THE SUTURE DURING CONTINUOUS SUTURING ON THE FASCIA. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MP2295, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8730834&pc=GAM | 2210968-2019-83234 | 2019-05-31 | Malfunction | ETHICON INC. | 2019-06-25 | GAM | VICRYL PLUS SUTURE 27"(70CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CAESAREAN SECTION ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS BROKEN WHEN PULLING THE SUTURE DURING CONTINUOUS SUTURING ON THE FASCIA. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MP2295, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8731401&pc=GAM | 2210968-2019-83251 | 2019-05-27 | Malfunction | ETHICON INC. | 2019-06-25 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS BROKEN DURING INTERRUPTED SUTURING. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8732827&pc=GAM | 2210968-2019-83260 | 2019-05-25 | Malfunction | ETHICON INC. | 2019-06-25 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING USE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ME2338, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8726569&pc=GAM | 2210968-2019-83171 | 2019-05-23 | Malfunction | ETHICON INC. | 2019-06-24 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN INTESTINAL ADHESION LYSIS ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8726594&pc=GAM | 2210968-2019-83172 | 2019-04-21 | Malfunction | ETHICON INC. | 2019-06-24 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SKIN STITCH PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8726807&pc=GAM | 2210968-2019-83176 | 2019-05-29 | Malfunction | ETHICON INC. | 2019-06-24 | GAM | VICRYL SUTURE 18"(45CM) 1 UND | Device Contamination with Chemical or Other Material | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, BEFORE OPENING THE PACKAGE, IT WAS FOUND THAT THERE HAD BEEN A FOREIGN SUBSTANCE LIKE 2 ¿ 3 CM HAIR IN THE PACKAGE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). A SEALED BOX OF PRODUCT CODE JB725, LOT # PBK081 WAS RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF SEALED BOX, A FOREIGN MATTER (HAIR) WAS OBSERVED UNDER FILM OF THE BOX. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO SAMPLE CONDITION THE ASSIGNABLE CAUSE OF THE FOREIGN MATTER IS A HAIR. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8727599&pc=GAM | 2210968-2019-83190 | 2019-05-24 | Malfunction | ETHICON INC. | 2019-06-24 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. CHANGED TO ANOTHER SUTURE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8727716&pc=GAM | 2210968-2019-83192 | 2019-05-23 | Malfunction | ETHICON INC. | 2019-06-24 | GAM | VICRYL RAP UND 90CM 2-0 V-39 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PERINEAL REPAIR ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE DETACHED FROM THE THREAD. THE PRODUCT WAS RETRIEVED AND A NEW SUTURE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8729693&pc=GAM | 2210968-2019-83221 |  | Malfunction | ETHICON INC. | 2019-06-24 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING USE. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8729694&pc=GAM | 2210968-2019-83219 |  | Malfunction | ETHICON INC. | 2019-06-24 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING USE. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8729695&pc=GAM | 2210968-2019-83220 |  | Malfunction | ETHICON INC. | 2019-06-24 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING USE. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8724152&pc=GAM | 2210968-2019-83117 | 2019-05-29 | Malfunction | ETHICON INC. | 2019-06-21 | GAM | VICRYL RAPIDE SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. WHEN THE SUTURE IS USED, THE FILAMENTS BREAK. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8724621&pc=GAM | 2210968-2019-83114 | 2019-05-29 | Malfunction | ETHICON INC. | 2019-06-21 | GAM | VICRYL VIO 75CM M1.5 4/0 SH-2 PLUS | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT DENTAL SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE DETACHED FROM SUTURE WHEN KNOT TYING AT THE GUM AREA. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE W9113, ME8CJXA0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8724668&pc=GAM | 2210968-2019-83122 | 2019-05-10 | Malfunction | ETHICON INC. | 2019-06-21 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN CARDIOVASCULAR SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS FOUND THAT THE NEEDLE HAD BEEN BROKEN AT THE SWAGE PART BEFORE USE. THE PRODUCT WAS NOT USED FOR THE PATIENT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8718657&pc=GAM | 2210968-2019-83058 | 2019-05-22 | Malfunction | ETHICON INC. | 2019-06-20 | GAM | VICRYL CT VIO 30IN USP2-0 | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS ALLEGED THAT THE TIP OF THE NEEDLE BROKE OFF. THE TIP WAS FOUND AND THERE WAS NO PATIENT EFFECT. THERE WAS NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8719359&pc=GAM | 2210968-2019-83068 | 2019-05-06 | Malfunction | ETHICON INC. | 2019-06-20 | GAM | VICRYL PLUS SUTURE 36"(90CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED BY HEALTH AUTHORITY THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE BEFORE PERINEAL STITCH. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MH2240, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8713390&pc=GAM | 2210968-2019-82951 | 2019-05-08 | Malfunction | ETHICON INC. | 2019-06-19 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION AND TOTAL HYSTERECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE TIP BROKE DURING USE. NO BROKEN PIECE WAS FOUND IN PATIENT BODY BY X-RAY. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THE PROCEDURE WAS PROLONGED 2 HOURS. PATIENT CONDITION LISTED AS STABLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ML2593, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? NO. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8713516&pc=GAM | 2210968-2019-82954 | 2019-05-22 | Malfunction | ETHICON INC. | 2019-06-19 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULL OFF THE SUTURE WHEN OUT OF PACKAGE DURING USE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8707649&pc=GAM | 2210968-2019-82852 |  | Malfunction | ETHICON INC. | 2019-06-18 | GAM | VICRYL CT UD 27IN USP4-0 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE OFF AFTER THE FIRST USE. THE NEEDLE WENT OFF THREAD WHEN THE NEEDLE DRIVER WAS GETTING PUT ON. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8708385&pc=GAM | 2210968-2019-82859 |  | Malfunction | ETHICON INC. | 2019-06-18 | GAM | VICRYL SUTURE 18"(45CM) 1 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE WAS BROKEN AT THE BODY PART WHEN THE NEEDLE PASSED FROM THE NEEDLE TIP TO THE BODY PART HELD THROUGH THE FASCIA. NO DEVICES WERE LEFT IN THE PATIENT¿S BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8708864&pc=GAM | 2210968-2019-82867 | 2019-05-24 | Malfunction | ETHICON INC. | 2019-06-18 | GAM | VICRYL PLUS SUTURE 18"(45CM) 4-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE NEEDLE BROKE WHEN SUTURING THE DUCT. THE SURGEON DID NOT HOLD THE TIP WITH THE NEEDLE HOLDER. WHEN THE NURSE CHECKED THE NEEDLE COUNT AFTER SURGERY, THEY FOUND THE NEEDLE TIP WAS BROKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8708865&pc=GAM | 2210968-2019-82866 | 2019-05-24 | Malfunction | ETHICON INC. | 2019-06-18 | GAM | VICRYL PLUS SUTURE 18"(45CM) 4-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE NEEDLE BROKE WHEN SUTURING THE DUCT. THE SURGEON DID NOT HOLD THE TIP WITH THE NEEDLE HOLDER. WHEN THE NURSE CHECKED THE NEEDLE COUNT AFTER SURGERY, THEY FOUND THE NEEDLE TIP WAS BROKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8708867&pc=GAM | 2210968-2019-82865 | 2019-05-24 | Malfunction | ETHICON INC. | 2019-06-18 | GAM | VICRYL PLUS SUTURE 18"(45CM) 4-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE NEEDLE BROKE WHEN SUTURING THE DUCT. THE SURGEON DID NOT HOLD THE TIP WITH THE NEEDLE HOLDER. WHEN THE NURSE CHECKED THE NEEDLE COUNT AFTER SURGERY, THEY FOUND THE NEEDLE TIP WAS BROKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8703727&pc=GAM | 2210968-2019-82791 | 2019-05-10 | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN BREAST PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE DETACHED FROM SUTURE DURING USE. THE NEEDLE WAS REMOVED FROM THE SURGICAL FIELD. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ARE DEVICES AVAILABLE FOR EVALUATION? PLEASE FOLLOW UP AND PROVIDE DEVICE RETURN STATUS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8703752&pc=GAM | 2210968-2019-82793 | 2019-05-10 | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN BREAST PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE DETACHED FROM SUTURE DURING USE. THE NEEDLE WAS REMOVED FROM THE SURGICAL FIELD. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. LOT NUMBER? ARE DEVICES AVAILABLE FOR EVALUATION? PLEASE FOLLOW UP AND PROVIDE DEVICE RETURN STATUS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8703801&pc=GAM | 2210968-2019-82794 | 2019-05-20 | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Break | No Known Impact Or Consequence To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL HIP ARTHROPLASTY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS BROKEN EASILY DURING INTERRUPTED SUTURING ON THE SUBCUTIS. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PAK576 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8704195&pc=GAM | 2210968-2019-82808 | 2019-05-15 | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF. THE NEEDLE WAS RETRIEVED FROM PATIENT. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8704222&pc=GAM | 2210968-2019-82811 | 2019-05-20 | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS DETACHED FROM THE SUTURE WHEN PULLING DURING INTERRUPTED SUTURING. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8704230&pc=GAM | 2210968-2019-82812 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL VIO 70CM M3.5 USP0 SGL ARM CT-3 PLUS | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING FASCIA CLOSURE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: THE ACTUAL DEVICE COMPONENT WAS IDENTIFIED AS PRODUCT CODE V329H AND RECEIVED FOR EVALUATION. A FRACTURE WAS OBSERVED AT THE TIP OF THE NEEDLE. ONE SIDE OF THE NEEDLE WAS RECEIVED, THE MATING FRACTURE SURFACE WAS NOT PROVIDED FOR THIS EVALUATION. A SCANNING ELECTRON MICROSCOPE (SEM) WAS USED TO EXAMINE THE FRACTURE SURFACES AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACES WERE EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION OF (B)(4) REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. MECHANICAL DAMAGE OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDES ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE TIP OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: AN UNOPENED SAMPLE OF PRODUCT CODE V329 LOT MMBGCLQ0 WAS RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF THE SAMPLE, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLE WAS OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. NO NEEDLE BREAKAGE WAS OBSERVED. THE SUTURE WAS DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS WERE OBSERVED. THE SAMPLE WAS TESTED BY RESISTANCE TEST TO NEEDLE AND MET THE REQUIREMENTS. PER THE CONDITIONS OF THE SAMPLE RECEIVED, NO PERFORMANCE BREAKAGE NEEDLE WERE FOUND. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MMBGCLQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8704387&pc=GAM | 2210968-2019-82816 |  | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN OBGYN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE SUTURING ON THE FASCIA USING AN INSTRUMENT AT 3RD LIGATURE, THE SUTURE WAS DETACHED FROM THE NEEDLE. IT WAS NOT A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LPZ717 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHEN WE SEND THE SAMPLE TO YOU, WE WILL LET YOU KNOW ITS RETURN DATE AND TRACKING NUMBER. PLEASE CONFIRM THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY ACTUALLY INVOLVED WITH REPORTED ISSUE) THAT FAILED DURING THIS ONE EVENT REPORT/PROCEDURE ? IT IS REPORTED 3 NEEDLE-SUTURES. OR WAS THE REPORTED EVENT OF "OCCURRED IN 3 NEEDLE-SUTURES IN THE SAME LOT" NOTICED WITH MULTIPLE PROCEDURES? ONE PROCEDURE. NOTE: EVENTS REPORTED VIA MW 2210968-2019-82815, 2210968-2019-82816, 2210968-2019-82817.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8704389&pc=GAM | 2210968-2019-82817 |  | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN OBGYN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE SUTURING ON THE FASCIA USING AN INSTRUMENT AT 3RD LIGATURE, THE SUTURE WAS DETACHED FROM THE NEEDLE. IT WAS NOT A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LPZ717 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHEN WE SEND THE SAMPLE TO YOU, WE WILL LET YOU KNOW ITS RETURN DATE AND TRACKING NUMBER. PLEASE CONFIRM THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY ACTUALLY INVOLVED WITH REPORTED ISSUE) THAT FAILED DURING THIS ONE EVENT REPORT/PROCEDURE ? IT IS REPORTED 3 NEEDLE-SUTURES. OR WAS THE REPORTED EVENT OF "OCCURRED IN 3 NEEDLE-SUTURES IN THE SAME LOT" NOTICED WITH MULTIPLE PROCEDURES? ONE PROCEDURE. NOTE: EVENTS REPORTED VIA MW 2210968-2019-82815, 2210968-2019-82816, 2210968-2019-82817.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8704391&pc=GAM | 2210968-2019-82815 |  | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN OBGYN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE SUTURING ON THE FASCIA USING AN INSTRUMENT AT 3RD LIGATURE, THE SUTURE WAS DETACHED FROM THE NEEDLE. IT WAS NOT A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LPZ717 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHEN WE SEND THE SAMPLE TO YOU, WE WILL LET YOU KNOW ITS RETURN DATE AND TRACKING NUMBER. PLEASE CONFIRM THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY ACTUALLY INVOLVED WITH REPORTED ISSUE) THAT FAILED DURING THIS ONE EVENT REPORT/PROCEDURE ? IT IS REPORTED 3 NEEDLE-SUTURES. OR WAS THE REPORTED EVENT OF "OCCURRED IN 3 NEEDLE-SUTURES IN THE SAME LOT" NOTICED WITH MULTIPLE PROCEDURES? NE PROCEDURE. EVENTS REPORTED VIA MW 2210968-2019-82815, 2210968-2019-82816, 2210968-2019-82817.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8704440&pc=GAM | 2210968-2019-82818 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL RAP UND 90CM M3 2/0 V-34 | Degraded; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A SKIN LACERATION REPAIR ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE SUTURE WAS DISPENSED AND IT BROKE INTO TWO PIECES WHEN REMOVED FROM THE PACKAGE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. UPON EVALUATION OF DEVICE, SUTURE HAS BEGUN WITH THE PROCESS OF DEGRADATION. THE FOIL WAS INSPECTED AND MULTIPLE HOLES IN CAVITY WERE NOTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: AN OPENED SAMPLE OF PRODUCT CODE V5170, LOT # JB8HHMQ0 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE, THE NEEDLE WAS RECEIVED DETACHED AND SUTURE REMNANT WAS NOTED INTO THE BARREL HOLE. THE SUTURE HAS BEGUN WITH THE PROCESS OF DEGRADATION, SINCE, THE STRAND WAS BROKEN IN MULTIPLES PIECES DURING THE INSPECTION, THIS IS THE CAUSE THAT BREAKAGE SUTURE. THE PRODUCT CODE V5170 CONTAINS AN ABSORBABLE SUTURE AND AS THE SAMPLE WAS RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED AND NO FUNCTIONAL TEST CAN BE PERFORMED. THE FOIL WAS INSPECTED AND MULTIPLES HOLES IN CAVITY WERE NOTED. THIS CONDITION CONTRIBUTED TO DEGRADATION OF THE SUTURES. THE CONDITIONS OF THE HOLES COULD NOT BE DETERMINED. PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE PRE-OP, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT . A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V5170H, JB8HHMQ0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8704489&pc=GAM | 2210968-2019-82819 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON 02/01/2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF. IT WAS RETRIEVED FROM PATIENT. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE IS NO REPORT ON PATIENT'S INJURY. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8704545&pc=GAM | 2210968-2019-82820 | 2019-03-03 | Malfunction | ETHICON INC. | 2019-06-17 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C- SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE DURING USE. THE NEEDLE WAS RETRIEVED FROM THE PATIENT DURING THE PROCEDURE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LP8CRMR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699249&pc=GAM | 2210968-2019-82741 | 2019-04-25 | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL CT UD 30IN USP2-0 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL KNEE REPLACEMENT PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BENT AND BROKE DURING USE. THE PROCEDURE WAS NOT DELAYED AND WAS COMPLETED SUCCESSFULLY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LLM083 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: THERE WAS NO SAMPLE RECEIVED FOR ANALYSIS. ONLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, AN OPENED SAMPLE OF PRODUCT CODE W9582T WITH LOT# LLM083 COULD BE OBSERVED. HOWEVER, NO BREAKAGE OR BENDING ISSUES COULD BE FOUND. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE FOLLOWING ADDITIONAL INFORMATION WAS OBTAINED: THE ATTACHED PICTURE SHOWS LOT LLM083, PLEASE CONFIRM THE EVENT REPORT IS FOR WHICH LOT? LLM083 (AS PER PIC) OR MLZ524 (AS REPORTED IN FILE)? EVENT REPORT FOR BOTH BATCHES. SURGEON EXPERIENCED NEEDLE CROOKED FOR TWICE. CONFIRM THE TOTAL QTY INVOLVED WITH REPORTED ISSUE? WAS 1 OR 2 NEEDLES INVOLVED? 2 NEEDLES INVOLVED. CONFIRM THE EVENT - DID THE SAME SINGLE REPORTED NEEDLE FIRST BEND AND THEN BREAK? OR 2 NEEDLES INVOLVED WHERE 1 NEEDLE BENT AND OTHER BROKE? FOR BOTH NEEDLES, SURGEONS COMMENTED NEEDLES CROOKED EASILY, NEEDLES WASN¿T SHARP (NOT SMOOTH WHEN HE SUTURED), AND NEEDLES TEND TO BREAK EASILY. CONFIRM THE LOT #'S FOR BOTH NEEDLES? BOTH BATCHES. HOWEVER, THE BATCH IN THE PHOTO WAS NOT AVAILABLE. REGARDING THE BROKEN NEEDLE, PLEASE ANSWER THE FOLLOWING: DID THE BROKEN NEEDLE OR NEEDLE PIECE FALL INTO PATIENT? NO. IF YES, WAS IT REMOVED? N/A. AND HOW? N/A. ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THE EVENT REPORT? UNKNOWN. IT IS MENTIONED BELOW THAT THE PATIENT CONSEQUENCE IS UNKNOWN. PLEASE ADVISE. DID THE HOSPITAL SURGEON/NURSE REPORT TO YOU THAT ANY PATIENT HARM OCCURRED AS A RESULT OF THE NEEDLE BREAKING? NO PATIENT HARM REPORTED. CAN YOU CONFIRM IF BOTH BATCHES OCCURRED IN THE SAME SURGERY AND SAME PATIENT? OR SEPARATE SURGERY AND DIFFERENT PATIENT? BOTH BATCHES WERE FOR SAME SURGERY AND SAME PATIENTS. NOTE: EVENTS REPORTED VIA MEDWATCH 2210968-2019-81518.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699323&pc=GAM | 2210968-2019-82742 | 2019-04-18 | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON (B)(6) AND SUTURE WAS USED. DURING SEWING SKIN IN PROCEDURE, THE NEEDLE PULLED OFF. THE NEEDLE WAS RETRIEVED FROM PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MG2643, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699827&pc=GAM | 2210968-2019-82754 |  | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL SUTURE 36"(90CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING CONTINUOUS SUTURING. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LDM025 BATCH NUMBER, AND NO NON-CONFORMANCE'S RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699849&pc=GAM | 2210968-2019-82755 |  | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL SUTURE 36"(90CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING CONTINUOUS SUTURING. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LMM520 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8700887&pc=GAM | 2210968-2019-82761 | 2019-05-20 | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL* UND 3-0 70CM (1)SC-20 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A PHOTO UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: LOT #? AL1221. PRODUCT CODE? J123. PROCEDURE/SURGERY NAME? INFORMATION NOT PROVIDED. DID THE NEEDLE PIECE FALL INTO PATIENT? NO. IF NOT FOUND, DOES A PIECE OF THE NEEDLE REMAIN IN THE PATIENT¿S TISSUE? NO. WAS THE PATIENT BROUGHT BACK TO OR TO HAVE NEEDLE PIECE REMOVED OR WAS IT REMOVED DURING THE INITIAL PROCEDURE? NO. MEDICAL /SURGICAL INTERVENTION TO HAVE THE NEEDLE PIECE REMOVED? NO. ANY PATIENT CONSEQUENCE/AE OUTCOME AS A RESULT OF EVENT REPORT? NO. WAS THE PROCEDURE COMPLETED SUCCESSFULLY? IF YES, HOW? YES, ANOTHER SUTURE WAS USED PATIENT CURRENT CONDITION? INFORMATION NOT PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8701156&pc=GAM | 2210968-2019-82765 | 2019-04-26 | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING SUTURING. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MK2826, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8701158&pc=GAM | 2210968-2019-82764 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED AWAY FROM THE THREAD. THE NEEDLE MAY HAVE FALLEN INSIDE THE ABDOMEN OF THE PATIENT AND WAS SEARCHED BY SCOPIE. THE PROCEDURE WAS FINISHED WITH ANOTHER SUTURE WITHOUT ISSUE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV2017, MMBDGMR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING ADDITIONAL INFORMATION WAS OBTAINED: FOR ONE PROCEDURE, THE NEEDLE IS FALLEN INSIDE THE ABDOMEN OF THE PATIENT AND HAS BEEN SEARCHED BY SCOPIE. THE PROCEDURES WERE FINISHED WITH ANOTHER SUTURE WITHOUT ISSUE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE NEEDLE RETAINED IN PATIENT TISSUE? IF YES, IN WHAT PATIENT STRUCTURE/TISSUE? WHAT PRODUCT CODE AND LOT NUMBER WERE RETAINED IN THE PATIENT? IF NEEDLE IS RETAINED, ARE THERE PLANS TO REMOVE THE NEEDLE? IF NEEDLE IS NOT RETAINED, WAS THE NEEDLE REMOVED DURING THE INITIAL PROCEDURE OR DID THE PATIENT UNDERGO A SECOND SURGICAL PROCEDURE TO REMOVE THE NEEDLE? WHAT IS THE PATIENT¿S CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8701253&pc=GAM | 2210968-2019-82767 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL PLUS SUTURE 27"(70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CESAREAN SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD DETACHED FROM THE NEEDLE WHEN REPAIRING THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJM737 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8701282&pc=GAM | 2210968-2019-82770 | 2019-05-09 | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL PLUS SUTURE 27"(70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CESAREAN SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD WAS FOUND DETACHED FROM THE NEEDLE WHEN REPAIRING THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJM737 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8701684&pc=GAM | 2210968-2019-82783 | 2019-05-10 | Malfunction | ETHICON INC. | 2019-06-14 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT GYNECOLOGICAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE AT THE TIP DURING SEWING SUBCUTANEOUS TISSUE. THE PROCEDURE WAS EXTENDED OVER 30 MINUTES TO SEARCH FOR THE BROKEN PIECE. THE PIECE WAS RETRIEVED FROM INSIDE FAT LAYER. IT IS UNKNOWN IF X-RAY WAS USED. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ML7180, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN. DEVICE RETURN STATUS/FOLLOW UP.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8697290&pc=GAM | 2210968-2019-82720 | 2019-05-04 | Malfunction | ETHICON INC. | 2019-06-13 | GAM | VICRYL PLUS SUTURE 27"(70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CAESAREAN SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD DETACHED FROM THE NEEDLE WHEN REPAIRING THE UTERUS. THE PROCEDURE WAS COMPLETED SUCCESSFULLY BY OPENING ANOTHER NEW PACK OF SUTURE TO CONTINUE THE WOUND CLOSURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4) A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJM737 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: THERE WAS NO SAMPLE RECEIVED FOR ANALYSIS. ONLY PICTURES OF THE SAMPLE WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, A USED SAMPLES OF PRODUCT CODE VCP364 LOT# MJM73 WITH THE NEEDLE DETACHED COULD BE OBSERVED, HOWEVER, NO CONCLUSION COULD BE REACH ON HOW THIS HAPPEN AS THE SAMPLES WERE NOT RETURNED FOR ANALYSIS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE FOLLOWING ADDITIONAL INFORMATION WAS OBTAINED: CONFIRM THE LOT FOR THE 2ND DEVICE INVOLVED, IS IT THE SAME AS REPORTED? MJM737 WAS PROCEDURE COMPLETED SUCCESSFULLY? AND HOW? THE PROCEDURE WAS COMPLETED SUCCESSFULLY BY OPENING ANOTHER NEW PACK OF SUTURE TO CONTINUE THE WOUND CLOSURE. ANY PATIENT CONSEQUENCES? NO. ANY SAMPLE AVAILABLE TO BE RETURNED FOR EVALUATION? NO DEFECTED PRODUCT WERE RETURNED AS THE NURSES DISPOSED ALL. EVENTS REPORTED VIA MEDWATCH 2210968-2019-82719.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8697299&pc=GAM | 2210968-2019-82719 | 2019-05-04 | Malfunction | ETHICON INC. | 2019-06-13 | GAM | VICRYL PLUS SUTURE 27"(70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CAESAREAN SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD DETACHED FROM THE NEEDLE WHEN REPAIRING THE UTERUS. THE PROCEDURE WAS COMPLETED SUCCESSFULLY BY OPENING ANOTHER NEW PACK OF SUTURE TO CONTINUE THE WOUND CLOSURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJM737 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: THERE WAS NO SAMPLE RECEIVED FOR ANALYSIS. ONLY PICTURES OF THE SAMPLE WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, A USED SAMPLES OF PRODUCT CODE VCP364 LOT# MJM73 WITH THE NEEDLE DETACHED COULD BE OBSERVED, HOWEVER, NO CONCLUSION COULD BE REACH ON HOW THIS HAPPEN AS THE SAMPLES WERE NOT RETURNED FOR ANALYSIS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE FOLLOWING ADDITIONAL INFORMATION WAS OBTAINED: CONFIRM THE LOT FOR THE 2ND DEVICE INVOLVED, IS IT THE SAME AS REPORTED? MJM737. WAS PROCEDURE COMPLETED SUCCESSFULLY? AND HOW? THE PROCEDURE WAS COMPLETED SUCCESSFULLY BY OPENING ANOTHER NEW PACK OF SUTURE TO CONTINUE THE WOUND CLOSURE. ANY PATIENT CONSEQUENCES? NO. ANY SAMPLE AVAILABLE TO BE RETURNED FOR EVALUATION? NO DEFECTED PRODUCT WERE RETURNED AS THE NURSES DISPOSED ALL. NOTE: EVENTS REPORTED VIA MEDWATCH 2210968-2019-82720.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8697637&pc=GAM | 2210968-2019-82725 | 2019-04-10 | Malfunction | ETHICON INC. | 2019-06-13 | GAM | VICRYL SUTURE 27"(70CM) 0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GENSLING, LAPAROSCOPIC PORT SITE PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING CLOSURE THE SUTURE MATERIAL BROKE WHEN ATTEMPTED TO HAND TIE. NO ADVERSE PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ME2926, AND NO NON-CONFORMANCES WERE IDENTIFIED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION: HE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH MM2162 DATE OF MFG.: 11/01/2018. EXPIRATION DATE: 10/31/2023. BATCH UNKNOWN. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MM2162 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: LOT M2926 IS INVALID; PLEASE PROVIDE CORRECT LOT NUMBER . DID 3 DEVICES, (2 OF LOT MM2162 AND 1 OF THE OTHER LOT) HAD SUTURE BREAKAGE ISSUES IN A SINGLE PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8697668&pc=GAM | 2210968-2019-82723 | 2019-04-10 | Malfunction | ETHICON INC. | 2019-06-13 | GAM | VICRYL SUTURE 27"(70CM) 0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GENSLING, LAPAROSCOPIC PORT SITE PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING CLOSURE THE SUTURE MATERIAL BROKE WHEN ATTEMPTED TO HAND TIE. NO ADVERSE PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MM2162 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: LOT M2926 IS INVALID; PLEASE PROVIDE CORRECT LOT NUMBER. DID 3 DEVICES, (2 OF LOT MM2162 AND 1 OF THE OTHER LOT) HAD SUTURE BREAKAGE ISSUES IN A SINGLE PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8697671&pc=GAM | 2210968-2019-82726 | 2019-04-10 | Malfunction | ETHICON INC. | 2019-06-13 | GAM | VICRYL SUTURE 27"(70CM) 0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GENSLING, LAPAROSCOPIC PORT SITE PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING CLOSURE THE SUTURE MATERIAL BROKE WHEN ATTEMPTED TO HAND TIE. NO ADVERSE PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MM2162 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: LOT M2926 IS INVALID; PLEASE PROVIDE CORRECT LOT NUMBER. DID 3 DEVICES, (2 OF LOT MM2162 AND 1 OF THE OTHER LOT) HAD SUTURE BREAKAGE ISSUES IN A SINGLE PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8691279&pc=GAM | 2210968-2019-82651 | 2019-05-14 | Malfunction | ETHICON INC. | 2019-06-12 | GAM | VICRYL SUTURE 18" (45CM) 6-0 UND | Component or Accessory Incompatibility | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS NOTED THAT THE WRONG SUTURE WAS INSIDE ONE PACK OF BOX OF 12. OUTSIDE LABEL SAYS J489 AND INSIDE SUTURE SAYS J544. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE DEVICE HAD PACKAGING ISSUES. THERE WAS NO SAMPLE RECEIVED FOR ANALYSIS. ONLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, AN OPENED FOIL OF PRODUCT CODE J489 LOT# MEZ014 AND A FOLDER OF PRODUCT CODE J544 COULD BE OBSERVED. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692261&pc=GAM | 2210968-2019-82668 | 2019-05-17 | Malfunction | ETHICON INC. | 2019-06-12 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CORONARY ARTERY BYPASS GRAFT PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS DETACHED FROM THE NEEDLE EASILY DURING USE ON THE MUSCLE LAYER. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692293&pc=GAM | 2210968-2019-82669 | 2019-05-16 | Malfunction | ETHICON INC. | 2019-06-12 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A PLASTIC SURGERY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE TIP WAS BREAKING OFF. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692297&pc=GAM | 2210968-2019-82670 | 2019-05-16 | Malfunction | ETHICON INC. | 2019-06-12 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A PLASTIC SURGERY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE TIP WAS BREAKING OFF. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692837&pc=GAM | 2210968-2019-82681 | 2019-05-04 | Malfunction | ETHICON INC. | 2019-06-12 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A RESECTION OF PERIANAL ABSCESS PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE WHEN SEWING ON SKIN. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MGK372 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY ACTUALLY INVOLVED WITH REPORTED ISSUE) THAT FAILED DURING THIS ONE EVENT REPORT/PROCEDURE ? WAS IT ONLY 1 NEEDLE SUTURE DEVICE WITH ISSUE AND 2 BEING RETURNED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692845&pc=GAM | 2210968-2019-82679 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-12 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT IT IS POSSIBLE THAT AN ABSORBABLE SUTURE WAS NOT ABSORBED. SUTURE WAS USED ON A PORT-SITE CLOSURE AND A PIECE OF SUTURE WAS THE RETRIEVED BY THE PATIENT FROM THE WOUND. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PRODUCT CODE. LOT NUMBER. HOW MANY DAYS POST-OP DID THE PATIENT RETRIEVE WHITE SUTURE FROM THE WOUND?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692888&pc=GAM | 2210968-2019-82678 | 2019-05-04 | Malfunction | ETHICON INC. | 2019-06-12 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A RESECTION OF PERIANAL ABSCESS PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE WHEN SEWING ON SKIN. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MGK372 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY ACTUALLY INVOLVED WITH REPORTED ISSUE) THAT FAILED DURING THIS ONE EVENT REPORT/PROCEDURE ? WAS IT ONLY 1 NEEDLE SUTURE DEVICE WITH ISSUE AND 2 BEING RETURNED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8687754&pc=GAM | 2210968-2019-82615 |  | Malfunction | ETHICON INC. | 2019-06-11 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OP, THE SUTURE FRAYED AND THEN BROKE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 11/15/2019. SEVENTEEN UNOPENED SAMPLES OF PRODUCT CODE JV453, LOT MMBBWCR0 WERE RECEIVED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THIRTEEN SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT.THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH.PER THE CONDITION OF THE REPRESENTATIVE SAMPLES, NO PERFORMANCE - BREAKAGE SUTURE WAS FOUND AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8688891&pc=GAM | 2210968-2019-82630 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-11 | GAM | VICRYL CT UD 30IN USP3-0 | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING USE, OFF FROM THE NEEDLE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MHZ847 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8689022&pc=GAM | 2210968-2019-82631 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-11 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 UND | Material Frayed; Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE SUTURE WAS FRAYING DURING USE, AND ALMOST POPPING OFF FROM THE NEEDLE. THE SUTURE FELT THINNER THAN NORMAL AND WEAK COMPOSITION. THE SUTURE FELT AS IF IT WAS GOING TO RIP WHEN SECURING TO TIE THE KNOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: AN UNOPENED SAMPLE OF PRODUCT CODE VCP215, LOT PBZ938 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLE WAS OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURE WAS DISPENSED WITHOUT PROBLEM AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, FRAYING SUTURE OR DIAMETER ISSUE WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED TO STRANDS USING A FEDERAL GAUGE TO MEASURE THE SUTURE DIAMETER AND THE RESULT MEET THE REQUIREMENTS. PER THE CONDITION OF THE REPRESENTATIVE SAMPLE, NO PERFORMANCE FRAYING SUTURE INTRA-OP OR SUTURE DIAMETER WAS FOUND AND THE TESTED SAMPLE MET THE FINISHED GOODS REQUIREMENTS. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER PBZ938 AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8689111&pc=GAM | 2210968-2019-82633 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-11 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 UND | Material Frayed; Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE SUTURE WAS FRAYING DURING USE, AND ALMOST POPPING OFF FROM THE NEEDLE. THE SUTURE FELT THINNER THAN NORMAL AND WEAK COMPOSITION. THE SUTURE FELT AS IF IT WAS GOING TO RIP WHEN SECURING TO TIE THE KNOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: AN UNOPENED SAMPLE OF PRODUCT CODE VCP110, LOT PAM821 WAS RETURNED FOR ANALYSIS. THE PRODUCT CODE IS MULTI-STRAND NON-NEEDLE AND REQUIRED TWELVE STRANDS PER PACKET. AS TOTAL TWELVE SUTURE WERE EVALUATED. DURING THE VISUAL INSPECTION OF THE SAMPLE, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLE WAS OPENED, AND THE TWELVE SUTURES WERE DISPENSED WITHOUT PROBLEM AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS, FRAYING SUTURE OR DIAMETER ISSUE WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED TO STRANDS USING A FEDERAL GAUGE TO MEASURE THE SUTURE DIAMETER AND THE RESULT MEET THE REQUIREMENTS. PER THE CONDITION OF THE REPRESENTATIVE SAMPLE, NO PERFORMANCE FRAYING SUTURE INTRA-OP OR SUTURE DIAMETER WAS FOUND AND THE TESTED SUTURES MET THE FINISHED GOODS REQUIREMENTS. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER PAM821 AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8689156&pc=GAM | 2210968-2019-82634 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-11 | GAM | VICRYL* RAPIDE* VLT 45 6-0 P-1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ENDOSCOPIC TURBINOPLASTY AND NASAL RECONSTRUCTION WITH GRAFT ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE NEEDLE WAS DISASSEMBLED FROM THE THREAD DURING USE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8689866&pc=GAM | 2210968-2019-82636 | 2019-05-13 | Malfunction | ETHICON INC. | 2019-06-11 | GAM | VICRYL SUTURE 18"(45CM) 4-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A SUBCUTANEOUS PORT PLACEMENT PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE FELL OFF THE SUTURE WHILE SUTURING THE TISSUE AROUND THE PORT. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MLM346 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8689899&pc=GAM | 2210968-2019-82639 |  | Malfunction | ETHICON INC. | 2019-06-11 | GAM | VICRYL SUTURE 27"(70CM) 4-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE NEEDLE WAS BROKEN AT THE TIP DURING USE. NO BROKEN PIECES WERE DROPPED IN THE PATIENT BODY. THERE WERE NO PATIENT CONSEQUENCES WERE REPORTED. THERE IS NO FURTHER INFORMATION. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJZ443 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8684036&pc=GAM | 2210968-2019-82576 |  | Malfunction | ETHICON INC. | 2019-06-10 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE CAME OFF THE SUTURE. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE DEVICE HAD PULL OFF - SUTURE NEEDLE. SIX UNOPENED SAMPLES OF PRODUCT CODE J497, LOT MJK059 WERE RETURNED FOR ANALYSIS. THE ISSUE SAMPLE WAS NOT RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF SIX SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8684251&pc=GAM | 2210968-2019-82577 | 2019-04-22 | Malfunction | ETHICON INC. | 2019-06-10 | GAM | VICRYL SUTURE 36"(90CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND A SUTURE WAS USED. DURING SURGERY, THE SUTURE WAS BREAKING. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8684940&pc=GAM | 2210968-2019-82597 |  | Malfunction | ETHICON INC. | 2019-06-10 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE FELL OFF THE SUTURE DURING USE. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: FOUR SEALED BOXES WITH TWELVE UNOPENED SAMPLES EACH OF PRODUCT CODE J497, LOT MJK059 WERE RETURNED FOR ANALYSIS. AS TOTAL 48 UNOPENED SAMPLES WERE RECEIVED FOR EVALUATION. THE ISSUE SAMPLE WAS NOT RETURNED. DURING THE VISUAL INSPECTION OF THIRTEEN SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT 4 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-82593, 2210968-2019-82594, 2210968-2019-82595 AND 2210968-2019-82596. ANALYSIS RESULTS HAVE BEEN INCLUDED IN REPORTS FOR EACH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8684950&pc=GAM | 2210968-2019-82594 |  | Malfunction | ETHICON INC. | 2019-06-10 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE FELL OFF THE SUTURE. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE DEVICE HAD PULL OFF - SUTURE NEEDLE. FOUR SEALED BOXES WITH TWELVE UNOPENED SAMPLES EACH OF PRODUCT CODE J497, LOT MJK059 WERE RETURNED FOR ANALYSIS. AS TOTAL 48 UNOPENED SAMPLES WERE RECEIVED FOR EVALUATION. THE ISSUE SAMPLE WAS NOT RETURNED. DURING THE VISUAL INSPECTION OF THIRTEEN SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 4 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-82593, 2210968-2019-82595, AND 2210968-2019-82596. ANALYSIS RESULTS HAVE BEEN INCLUDED IN REPORTS FOR EACH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8684953&pc=GAM | 2210968-2019-82596 |  | Malfunction | ETHICON INC. | 2019-06-10 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE FELL OFF THE SUTURE. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE DEVICE HAD PULL OFF - SUTURE NEEDLE. FOUR SEALED BOXES WITH TWELVE UNOPENED SAMPLES EACH OF PRODUCT CODE J497, LOT MJK059 WERE RETURNED FOR ANALYSIS. AS TOTAL 48 UNOPENED SAMPLES WERE RECEIVED FOR EVALUATION. THE ISSUE SAMPLE WAS NOT RETURNED. DURING THE VISUAL INSPECTION OF THIRTEEN SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 4 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-82593, 2210968-2019-82594, 2210968-2019-82595, AND 2210968-2019-82596. ANALYSIS RESULTS HAVE BEEN INCLUDED IN REPORTS FOR EACH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8684955&pc=GAM | 2210968-2019-82595 |  | Malfunction | ETHICON INC. | 2019-06-10 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE FELL OFF THE SUTURE. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE DEVICE HAD PULL OFF - SUTURE NEEDLE. FOUR SEALED BOXES WITH TWELVE UNOPENED SAMPLES EACH OF PRODUCT CODE J497, LOT MJK059 WERE RETURNED FOR ANALYSIS. AS TOTAL 48 UNOPENED SAMPLES WERE RECEIVED FOR EVALUATION. THE ISSUE SAMPLE WAS NOT RETURNED. DURING THE VISUAL INSPECTION OF THIRTEEN SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 4 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-82593, 2210968-2019-82594, 2210968-2019-82595, AND 2210968-2019-82596. ANALYSIS RESULTS HAVE BEEN INCLUDED IN REPORTS FOR EACH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8685001&pc=GAM | 2210968-2019-82598 |  | Malfunction | ETHICON INC. | 2019-06-10 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE MATERIAL BROKE BEFORE THE KNOT REACHED THE SKIN. A DIFFERENT DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PA2200, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8685029&pc=GAM | 2210968-2019-82599 |  | Malfunction | ETHICON INC. | 2019-06-10 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE MATERIAL BROKE BEFORE THE KNOT REACHED THE SKIN. A DIFFERENT DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PA2200, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8680788&pc=GAM | 2210968-2019-82536 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-07 | GAM | VICRYL RAPIDE SUTURE | Break; Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE STAFF NOTED A DISUNION OF SUTURE WAS NOTED WHEN USED ON PERINEAL SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT DO YOU MEAN BY "DISUNION OF THE SUTURE"? PLEASE EXPLAIN THE SPECIFIC QUALITY ISSUE? ALSO, PLEASE PROVIDE ANSWERS TO BELOW: DID THE WHOLE SUTURE DETACH / SEPARATE / PULL OFF FROM THE WHOLE NEEDLE? DID THE SUTURE ONLY UNRAVEL OR FRAY? DID THE SUTURE BREAK INTO 2 PIECES? WAS THE EVENT REPORT NOTICED DURING USE ON PATIENT WHILE SUTURING? ANY PATIENT CONSEQUENCES? WAS THE PROCEDURE COMPLETED SUCCESSFULLY? AND HOW?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676424&pc=GAM | 2210968-2019-82491 |  | Malfunction | ETHICON INC. | 2019-06-06 | GAM | VICRYL CT VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A VASCULAR PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE SEPARATED FROM THE SUTURE DURING TISSUE PASSAGE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ADDITIONAL EVALUATION SUMMARY: TWENTY-FIVE UNOPENED SAMPLES OF PRODUCT CODE V102, LOT PABDHSQ0 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PULL OFF SUTURE NEEDLE WAS FOUND AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676943&pc=GAM | 2210968-2019-82504 | 2019-05-09 | Malfunction | ETHICON INC. | 2019-06-06 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED BY THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE TIP OF SUTURE NEEDLE CAME OFF WHILE SEWING UP THE UTERUS. THE SURGEON NOTED THE TIP WAS MISSING AND REMOVED IT FROM THE PATIENT. THE SURGERY WAS DELAYED AN UNKNOWN NUMBER OF MINUTES DUE TO THE REPORTED EVENT. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LL7085, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676969&pc=GAM | 2210968-2019-82505 | 2019-05-09 | Malfunction | ETHICON INC. | 2019-06-06 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED BY THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE TIP OF SUTURE NEEDLE CAME OFF WHILE SEWING UP THE UTERUS. THE SURGEON NOTED THE TIP WAS MISSING AND REMOVED IT FROM THE PATIENT. THE SURGERY WAS DELAYED AN UNKNOWN NUMBER OF MINUTES DUE TO THE REPORTED EVENT. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PA2652, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8670124&pc=GAM | 2210968-2019-82422 | 2019-05-01 | Malfunction | ETHICON INC. | 2019-06-05 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF THE SUTURE. THE FALLEN PIECE WAS RETRIEVED FROM PATIENT. CHANGED ANOTHER ONE TO COMPLETE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8668411&pc=GAM | 2210968-2019-82368 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-06-04 | GAM | VICRYL SUTURE 27"(70CM) 0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING CLOSURE OF THE FASCIA. THE TIP BROKE IN AROUND 2CM IN LENGTH. ALL PIECES OF THE NEEDLE WERE RETRIEVED DURING THE PROCEDURE AND NO X-RAY WAS NEEDED. THE SURGEON OPINED THAT A CONTRIBUTING FACTOR TO THE EFFORTS PLACED ON THE NEEDLE USED IN CLOSURE WAS THAT THE PATIENT WAS OBESE AND DIFFICULT TO CLOSE THE FASCIA. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8668603&pc=GAM | 2210968-2019-82374 | 2019-04-26 | Malfunction | ETHICON INC. | 2019-06-04 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE TIP WAS BROKEN. NO BROKEN PIECE WAS DROPPED IN THE PATIENT¿S BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8669268&pc=GAM | 2210968-2019-82401 | 2019-05-08 | Malfunction | ETHICON INC. | 2019-06-04 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE FELL OFF THE SUTURE DURING USE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PAZ709 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8669288&pc=GAM | 2210968-2019-82403 | 2019-05-08 | Malfunction | ETHICON INC. | 2019-06-04 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE FELL OFF THE SUTURE DURING USE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MMK144 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8669293&pc=GAM | 2210968-2019-82404 | 2019-05-08 | Malfunction | ETHICON INC. | 2019-06-04 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE FELL OFF THE SUTURE DURING USE. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PAZ709 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8669299&pc=GAM | 2210968-2019-82405 | 2019-05-08 | Malfunction | ETHICON INC. | 2019-06-04 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A GYNECOLOGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS USED TO CLOSE THE UTERUS. DURING THE SURGERY THE SUTURE WAS TORN. THE SURGEON TOOK ANOTHER LIKE SUTURE AND FINISHED THE SURGERY. THERE WAS NO DAMAGE TO THE PATIENT OR SIGNIFICANT DELAY IN SURGERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8662462&pc=GAM | 2210968-2019-82248 | 2019-04-24 | Malfunction | ETHICON INC. | 2019-06-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Material Frayed | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THERE WAS SUTURE BIFURCATION. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8658975&pc=GAM | 2210968-2019-82098 | 2019-04-18 | Malfunction | ETHICON INC. | 2019-05-31 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Material Separation | No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NATURAL LABOR PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING SEWING THE SKIN THE NEEDLE PULLED OFF. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8659910&pc=GAM | 2210968-2019-82168 | 2019-02-10 | Malfunction | ETHICON INC. | 2019-05-31 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE WAS FOUND DETACHED FROM THE SUTURE WHEN THE PACKAGE WAS OPENED. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. UPON EVALUATION, A FRACTURE WAS OBSERVED AT THE SUTURE ATTACHMENT OF THE NEEDLE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: IT WAS REPORTED PERFORMANCE PULL OFF SUTURE NEEDLE PRE-OP. A USED NEEDLE WAS RETURNED FOR EVALUATION. A FRACTURE WAS OBSERVED AT THE SUTURE ATTACHMENT OF THE NEEDLE. THE NEEDLE WAS RECEIVED IN TWO PIECES, ONLY ONE OF THE MATING FRACTURE SURFACES WAS EXAMINED FOR THIS EVALUATION. THE FRACTURE SURFACES WERE MICROSCOPICALLY EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. THIS WAS A DUCTILE FRACTURE. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE ATTACHMENT AREA OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. THE MANUFACTURING RECORDS COULDN'T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. THE NEEDLE BREAKAGE WAS CONFIRMED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8659933&pc=GAM | 2210968-2019-82167 | 2019-04-26 | Malfunction | ETHICON INC. | 2019-05-31 | GAM | VICRYL SUTURE 8-18" (45CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING USE THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING INTERRUPTED SUTURING ON THE MYOMETRIUM. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660064&pc=GAM | 2210968-2019-82171 | 2019-04-26 | Malfunction | ETHICON INC. | 2019-05-31 | GAM | VICRYL SUTURE 8-18" (45CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING USE THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING INTERRUPTED SUTURING ON THE MYOMETRIUM. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660139&pc=GAM | 2210968-2019-82173 | 2019-04-26 | Malfunction | ETHICON INC. | 2019-05-31 | GAM | VICRYL SUTURE 8-18" (45CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING USE THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING INTERRUPTED SUTURING ON THE MYOMETRIUM. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660181&pc=GAM | 2210968-2019-82175 | 2019-04-26 | Malfunction | ETHICON INC. | 2019-05-31 | GAM | VICRYL SUTURE 8-18" (45CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING USE THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING INTERRUPTED SUTURING ON THE MYOMETRIUM. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: ONE WINDING FORMER WITH EIGHT USED NEEDLES AND A DISPENSED SUTURE OF PRODUCT CODE JB753, LOT MHM739 WERE RETURNED FOR ANALYSIS. THE PRODUCT CODE IS CONTROL RELEASE AND REQUIRED EIGHT STRANDS PER PACKET. DURING THE VISUAL INSPECTION OF EIGHT USED NEEDLES, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE OF THE NEEDLES WAS EXAMINED UNDER MAGNIFICATION, NO SUTURE REMNANT WAS NOTED AND SLIGHTLY MARKS WERE OBSERVED ON THE BODY OF NEEDLES. THE SUTURE PRESENT BODY FLUIDS AND A CORRECT INSERTION MARK; HOWEVER, THE FORCE USED TO DETACH NEEDLE FROM THE SUTURE COULD NOT BE DETERMINED. ANOTHER SEVEN SUTURES WERE NOT RETURNED FOR ANALYSIS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. DUE TO CONDITION OF THE SAMPLE, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT OF: ¿ THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING INTERRUPTED SUTURING ON THE MYOMETRIUM¿. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660207&pc=GAM | 2210968-2019-82176 | 2019-04-26 | Malfunction | ETHICON INC. | 2019-05-31 | GAM | VICRYL SUTURE 8-18" (45CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING USE THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING INTERRUPTED SUTURING ON THE MYOMETRIUM. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656436&pc=GAM | 2210968-2019-82137 | 2019-04-03 | Malfunction | ETHICON INC. | 2019-05-30 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE WHEN SEWING THE SKIN. THE BROKEN PIECE WAS RETRIEVED FROM THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LLZ964 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656473&pc=GAM | 2210968-2019-82138 | 2019-05-02 | Malfunction | ETHICON INC. | 2019-05-30 | GAM | VICRYL VIO 75CM M3 USP2/0 SGL ARM MH-1 PLUS | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED AWAY FROM THE THREAD DURING USE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV2016, MPBCPTR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8650570&pc=GAM | 2210968-2019-82089 | 2019-04-18 | Malfunction | ETHICON INC. | 2019-05-29 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NATURAL LABOR PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING SEWING THE SKIN THE NEEDLE PULLED OFF. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8650596&pc=GAM | 2210968-2019-82090 | 2019-04-18 | Malfunction | ETHICON INC. | 2019-05-29 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NATURAL LABOR PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING SEWING THE SKIN THE NEEDLE PULLED OFF. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8650615&pc=GAM | 2210968-2019-82059 | 2019-05-01 | Malfunction | ETHICON INC. | 2019-05-29 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A ROBOTIC HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE POPPED OFF THE SUTURE WHILE PUTTING THE NEEDLE INTO THE TROCAR. ADDITIONAL SUTURE FROM ANOTHER LOT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. EVALUATION: FIVE UNOPENED SAMPLES OF PRODUCT CODE J269, LOT MPZ820 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF FIVE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS AND THE REPORTED COMPLAINT COULD NOT BE OBSERVED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8650639&pc=GAM | 2210968-2019-82092 |  | Malfunction | ETHICON INC. | 2019-05-29 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED-OFF DURING TISSUE PASSAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IT WAS REPORTED PERFORMANCE PULL OFF SUTURE NEEDLE. A DETACHED NEEDLE AND A DISPENSED SUTURE OF PRODUCT CODE VCP497 RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED AND MARKS ON THE EDGE NEEDLE THAT APPEAR TO BE BY USE OF A SURGICAL INSTRUMENT COLD BE OBSERVED. TO AVOID THIS KIND OF DAMAGE, GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. THE SUTURE WAS RECEIVED DISPENSED AND THE END WAS OBSERVED DAMAGED (CUT), THIS CONDITION CAUSES THE SEPARATION OF THE NEEDLE FROM THE SUTURE. IN HANDLING THIS OR ANY OTHER SUTURE MATERIAL, CARE SHOULD BE TAKEN TO AVOID DAMAGE FROM HANDLING. AVOID CRUSHING OR CRIMPING DAMAGE DUE TO APPLICATION OF SURGICAL INSTRUMENTS SUCH AS FORCEPS OR NEEDLE HOLDERS. THE MANUFACTURING RECORDS COULDN'T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. ACCORDING TO SAMPLE CONDITION, THE ASSIGNABLE CAUSE OF THE PERFORMANCE PULL OFF SUTURE NEEDLE WAS AN IMPROPER HANDLING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8650640&pc=GAM | 2210968-2019-82091 |  | Malfunction | ETHICON INC. | 2019-05-29 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED-OFF DURING TISSUE PASSAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MFR: IT WAS REPORTED PERFORMANCE PULL OFF SUTURE NEEDLE. AN EMPTY OPENED FOLDER, A DETACHED NEEDLE AND A DISPENSED SUTURE OF PRODUCT CODE VCP497 RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED AND MARKS ON THE EDGE NEEDLE THAT APPEAR TO BE BY USE OF A SURGICAL INSTRUMENT COLD BE OBSERVED. TO AVOID THIS KIND OF DAMAGE, GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. THE SUTURE WAS RECEIVED DISPENSED AND THE END WAS OBSERVED DAMAGED (CUT), THIS CONDITION CAUSES THE SEPARATION OF THE NEEDLE FROM THE SUTURE. IN HANDLING THIS OR ANY OTHER SUTURE MATERIAL, CARE SHOULD BE TAKEN TO AVOID DAMAGE FROM HANDLING. AVOID CRUSHING OR CRIMPING DAMAGE DUE TO APPLICATION OF SURGICAL INSTRUMENTS SUCH AS FORCEPS OR NEEDLE HOLDERS. THE MANUFACTURING RECORDS COULDN'T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. ACCORDING TO SAMPLE CONDITION, THE ASSIGNABLE CAUSE OF THE PERFORMANCE PULL OFF SUTURE NEEDLE WAS AN IMPROPER HANDLING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8650854&pc=GAM | 2210968-2019-82096 | 2019-04-18 | Malfunction | ETHICON INC. | 2019-05-29 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NATURAL LABOR PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING SEWING THE SKIN THE NEEDLE PULLED OFF. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8650893&pc=GAM | 2210968-2019-82097 | 2019-04-18 | Malfunction | ETHICON INC. | 2019-05-29 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NATURAL LABOR PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING SEWING, THE SKIN THE NEEDLE PULLED OFF. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8651216&pc=GAM | 2210968-2019-82099 | 2019-05-04 | Malfunction | ETHICON INC. | 2019-05-29 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN EMBOLECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BREAKS DURING THE PROCEDURE AND NOT AS STRONG. NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. EVALUATION: RECEIVED 1 COMPLAINT SAMPLE ALONG WITH SUTURE, NEEDLE FOR CODE NW2762, LOT T8013. SUTURE RECEIVED IN PARTIALLY DISINTEGRATED CONDITION. AS THE COMPLAINT SAMPLE RECEIVED IN OPEN CONDITION FURTHER INVESTIGATION ON COMPLAINT SAMPLE CANNOT BE PERFORMED EXCEPT VISUAL INSPECTION. THE FOIL PACK WAS NOT RETURNED FOR INVESTIGATION. RETURNED NEEDLE WAS INSPECTED UNDER MAGNIFICATION AND FOUND SATISFACTORY. REFERENCE SAMPLE: 1 INTACT REFERENCE SAMPLE WAS RETURNED FOR CODE NW2762/T8013. THE FOIL PACK WAS VISUALLY INSPECTED UNDER A MAGNIFICATION FOR ANY DAMAGE AND SHOWED A MULTITUDE OF WRINKLES OVER THE WHOLE FOIL. THE RECEIVED REFERENCE SAMPLE OPENED AND VISUALLY INSPECTED. THE PRIMARY PACK WAS OPENED, AND SUTURE AND NEEDLES WERE FOUND INTACT. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS AND NO SUCH DEFECTS WERE OBSERVED. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS AND NO SUCH DEFECTS WERE OBSERVED. THE SUTURE THEN TESTED FOR KNOT PULL TENSILE STRENGTH TEST AND FOUND TO BE MEETING SPECIFICATION REQUIREMENT. KNOT PULL TENSILE STRENGTH VALUE OF REFERENCE SAMPLE WAS FOUND TO MEET SPECIFICATION. FIVE RETAIN SAMPLES OF INCIDENT CODES AND LOT NUMBER WERE RETRIEVED FROM RETAINS FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS AND NO SUCH DEFECTS WERE OBSERVED. THE PRIMARY PACKS WERE OPENED, AND SUTURES AND NEEDLES WERE FOUND INTACT. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS AND NO SUCH DEFECTS WERE OBSERVED. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTE DEFECTS AND NO SUCH DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WERE TESTED FOR KNOT PULL TENSILE STRENGTH TEST AND FOUND TO MEET THE SPECIFICATION. AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED FOR ANY PROCESS DEVIATION, BUT NO DEVIATION WAS OBSERVED. FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION. FROM THE ABOVE ANALYSIS, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. AS THE COMPLAINT IS RELATED TO PERFORMANCE-BREAKAGE SUTURE, KNOT PULL TENSILE STRENGTH TEST IS APPLICABLE & MEASURABLE PARAMETER. KNOT PULL TENSILE STRENGTH TEST WAS PERFORMED OVER FIVE RETAIN SAMPLES TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. THE AVERAGE KNOT PULL TENSILE STRENGTH VALUE OF THE RETAIN SAMPLE WAS FOUND TO MEET THE REQUIREMENT. RETAIN AS WELL AS REFERENCE SAMPLE AVERAGE AS WELL AS INDIVIDUAL KNOT PULL TENSILE STRENGTH VALUES WERE FOUND TO MEET THE IN-HOUSE REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS. BASED ON THE ANALYSIS THIS IS NOT A DUPLICATED COMPLAINT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8651380&pc=GAM | 2210968-2019-82101 | 2019-05-01 | Malfunction | ETHICON INC. | 2019-05-29 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE FELL OFF THE SUTURE, THE CRIMP WAS BAD BETWEEN THE NEEDLE AND THE THREAD. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645385&pc=GAM | 2210968-2019-82034 | 2019-04-30 | Malfunction | ETHICON INC. | 2019-05-28 | GAM | VICRYL RAPIDE SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE POPPED OFF OF THE SUTURE AT THE SWAGE. A SECOND LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE PULL OFF SUTURE NEEDLE. A PAPER LID, AN EMPTY WINDING FORMER AND A SUTURE PIECE OF PRODUCT CODE VR490 WERE RETURNED FOR ANALYSIS. DURING VISUAL INSPECTION OF THE SAMPLE, THE INSERTION MARK COULD NOT BE OBSERVED AND THE OTHER EXTREME WAS CUT THAT APPEARS TO BE BY USE OF SURGICAL INSTRUMENT. IN ADDITION, THE NEEDLE WAS NOT RETURNED FOR EVALUATION. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. DUE TO CONDITION OF THE SAMPLE, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT OF: ¿ THE NEEDLE POPPED OFF OF THE SUTURE AT THE SWAGE¿.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645523&pc=GAM | 2210968-2019-82036 |  | Malfunction | ETHICON INC. | 2019-05-28 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN OPHTHALMIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE S-29 BENDS WHEN PASSED THROUGH MUSCLE OR THE TIP BREAKS BY ANGLING 90 DEGREES. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. LOT #? PRODUCT CODE? ANY SAMPLE AVAILABLE TO BE RETURNED FOR EVALUATION FOR THIS EVENT REPORT? FOR THIS EVENT REPORT - DID THE REPORTED NEEDLE ONLY BEND? OR DID IT BEND AND BREAK? WAS PROCEDURE FOR THIS EVENT REPORT COMPLETED SUCCESSFULLY? AND HOW? BECAUSE MULTIPLE PROCEDURES MENTIONED IN EVENT REPORT- PLEASE CONFIRM THE SPECIFIC NUMBER OF PATIENT EVENTS. HAVE ANY OF THESE EVENTS BEEN PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE COMPLAINT FILE NUMBER/ AFFILIATE NUMBER. FOR EACH PATIENT EVENT PLEASE PROVIDE ADDITIONAL COMPLAINT DETAILS: EVENT DATE, PRODUCT CODE AND LOT NUMBER INVOLVED, PROCEDURE, EVENT DESCRIPTION, QTY INVOLVED? WAS THERE ANY ADVERSE PATIENT OUTCOME? IF YES, WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED, WAS PROCEDURE COMPLETED SUCCESSFULLY AND HOW? ANY SAMPLE RETURN..ETC? EVENT DESCRIPTION. EVENT OCCURRENCE. PRODUCT CODE AND LOT NUMBER. EVENT DATE. PROCEDURE NAME. PATIENT CONSEQUENCES/AE REPORT IF ANY? ANY MEDICAL/SURGICAL INTERVENTION OR TREATMENT PROVIDED TO PATIENT POST-OP? TOTAL QTY OF SUTURES INVOLVED? SAMPLE RETURN, IF ANY? PLEASE PROVIDE THE STATUS OF THE DEVICE(S) AS IT HAS NOT BEEN RECEIVED FOR ANALYSIS. IF THE DEVICE HAS BEEN SHIPPED, PLEASE PROVIDE THE SHIPMENT TRACKING DETAILS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645673&pc=GAM | 2210968-2019-82039 |  | Malfunction | ETHICON INC. | 2019-05-28 | GAM | VICRYL SUTURE 18"(45CM) 6-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WAS NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645713&pc=GAM | 2210968-2019-82041 |  | Malfunction | ETHICON INC. | 2019-05-28 | GAM | VICRYL SUTURE 18"(45CM) 6-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WAS NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8647293&pc=GAM | 2210968-2019-82058 | 2019-05-01 | Malfunction | ETHICON INC. | 2019-05-28 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A ROBOTIC HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE POPPED OFF THE SUTURE WHILE PUTTING THE NEEDLE INTO THE TROCAR. ADDITIONAL SUTURE FROM ANOTHER LOT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. EVALUATION: FIVE UNOPENED SAMPLES OF PRODUCT CODE J269, LOT MGZ159 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF FIVE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS AND THE REPORTED COMPLAINT COULD NOT BE OBSERVED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8648378&pc=GAM | 2210968-2019-82072 | 2019-05-03 | Malfunction | ETHICON INC. | 2019-05-28 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING CLOSURE, THE NEEDLE BROKE. THE NEEDLE WAS LOST IN THE PATIENT AND AN X-RAY HAD TO BE TAKEN TO LOCATE IT. THIS RESULTED IN AN EXTRA 60 MINUTES OF PROCEDURE TIME. A NEW SUTURE WAS OPENED TO COMPLETE THE CLOSURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE BROKEN NEEDLE PIECE FOUND AND RETRIEVED FROM PATIENT? IF YES, HOW WAS IT RETRIEVED? ANY PATIENT CONSEQUENCE/ AE OUTCOME AS A RESULT OF THE EVENT REPORT? WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8640290&pc=GAM | 2210968-2019-81918 | 2019-04-24 | Malfunction | ETHICON INC. | 2019-05-24 | GAM | VICRYL VIO 70CM M3.5 USP0 SGLE ARMED SH PLUS | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE INTO TWO PARTS DURING THE FIRST USE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MPBCXTQ0 NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8640412&pc=GAM | 2210968-2019-81920 |  | Malfunction | ETHICON INC. | 2019-05-24 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN OPHTHALMIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE S-29 BENDS WHEN PASSED THROUGH MUSCLE OR THE TIP BREAKS BY ANGLING 90 DEGREES. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. LOT #? PRODUCT CODE? ANY SAMPLE AVAILABLE TO BE RETURNED FOR EVALUATION FOR THIS EVENT REPORT? FOR THIS EVENT REPORT - DID THE REPORTED NEEDLE ONLY BEND? OR DID IT BEND AND BREAK? WAS PROCEDURE FOR THIS EVENT REPORT COMPLETED SUCCESSFULLY? AND HOW? BECAUSE MULTIPLE PROCEDURES MENTIONED IN EVENT REPORT. PLEASE CONFIRM THE SPECIFIC NUMBER OF PATIENT EVENTS. HAVE ANY OF THESE EVENTS BEEN PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE COMPLAINT FILE NUMBER/ AFFILIATE NUMBER. FOR EACH PATIENT EVENT PLEASE PROVIDE ADDITIONAL COMPLAINT DETAILS: EVENT DATE, PRODUCT CODE AND LOT NUMBER INVOLVED, PROCEDURE, EVENT DESCRIPTION, QTY INVOLVED? WAS THERE ANY ADVERSE PATIENT OUTCOME? IF YES, WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED, WAS PROCEDURE COMPLETED SUCCESSFULLY AND HOW? ANY SAMPLE RETURN, ETC? EVENT DESCRIPTION, EVENT OCCURRENCE, PRODUCT CODE AND LOT NUMBER, EVENT DATE, PROCEDURE NAME. PATIENT CONSEQUENCES/AE REPORT IF ANY? ANY MEDICAL/SURGICAL INTERVENTION OR TREATMENT PROVIDED TO PATIENT POST-OP? TOTAL QTY OF SUTURES INVOLVED? SAMPLE RETURN, IF ANY? PLEASE PROVIDE THE STATUS OF THE DEVICE(S) AS IT HAS NOT BEEN RECEIVED FOR ANALYSIS. IF THE DEVICE HAS BEEN SHIPPED, PLEASE PROVIDE THE SHIPMENT TRACKING DETAILS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8641122&pc=GAM | 2210968-2019-81947 | 2019-04-29 | Malfunction | ETHICON INC. | 2019-05-24 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8639053&pc=GAM | 2210968-2019-81836 | 2019-04-26 | Malfunction | ETHICON INC. | 2019-05-23 | GAM | VICRYL RAPIDE SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING USE, THE NEEDLE POPPED OFF THE SUTURE AT THE SWAGE. ANOTHER SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8632934&pc=GAM | 2210968-2019-81672 |  | Malfunction | ETHICON INC. | 2019-05-22 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A C-SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE PULLED OFF OF THE SUTURE. IT WAS REPORTED THAT DURING SEWING UTERUS IN C-SECTION. THE NEEDLE WAS RETRIEVED FROM PATIENT. CHANGED TO ANOTHER ONE TO COMPLETE SURGERY. NO ADDITIONAL INFORMATION IS AVAILABLE. THERE IS NO REPORTED PATIENT CONSEQUENCE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633460&pc=GAM | 2210968-2019-81696 |  | Malfunction | ETHICON INC. | 2019-05-22 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EPISIOTOMY REPAIR ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE DURING KNOTTING. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: INVESTIGATION SUMMARY: TO VERIFY THE COMPLAINT, RESULTS OF RETAINED SAMPLES WERE REVIEWED AND NO DISCREPANCY WAS OBSERVED FOR KNOT PULL TENSILE STRENGTH TESTING. RESULTS FOR KNOT PULL TEST WERE COMPLYING WITH THE PRE-DEFINED ACCEPTANCE CRITERIA. FIVE RETAIN SAMPLES WERE SUBJECTED TO KNOT PULL TENSILE STRENGTH TEST TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. THE AVERAGE KNOT PULL TENSILE STRENGTH VALUE OF THE RETAIN SAMPLE MEETS THE IN-HOUSE AVERAGE KNOT PULL TENSILE STRENGTH REQUIREMENT. ALL THE INDIVIDUAL AS WELL AS AVERAGE KNOT PULL VALUES WERE FOUND TO MEET THE IN-HOUSE SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS. BATCH MANUFACTURING RECORDS OF NW2777/T8038 REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE, NEEDLE PULL-OFF VALUE AND MOISTURE TEST RESULTS AND FOUND TO MEET THE SPECIFICATION REQUIREMENT. THE FOLLOWING ADDITIONAL INFORMATION WAS RECEIVED: WERE ALL THE SUTURE STRANDS BREAKING WHILE USE ON THE PATIENT DURING SUTURING? YES. WAS PROCEDURE COMPLETED SUCCESSFULLY? AND HOW? YES. IT WAS EPISIOTOMY REPAIR, DOCTOR COMPLETE THE PROCEDURE WITH CATGUT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633511&pc=GAM | 2210968-2019-81705 |  | Malfunction | ETHICON INC. | 2019-05-22 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EPISIOTOMY REPAIR ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE DURING KNOTTING. A DIFFERENT DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: INVESTIGATION SUMMARY: TO VERIFY THE COMPLAINT, RESULTS OF RETAINED SAMPLES WERE REVIEWED AND NO DISCREPANCY WAS OBSERVED FOR KNOT PULL TENSILE STRENGTH TESTING. RESULTS FOR KNOT PULL TEST WERE COMPLYING WITH THE PRE-DEFINED ACCEPTANCE CRITERIA. THE COMPLAINT IS RELATED TO PERFORMANCE-BREAKAGE SUTURE, KNOT PULL TENSILE STRENGTH TEST IS APPLICABLE AND IS A MEASURABLE PARAMETER. FIVE RETAIN SAMPLES WERE SUBJECTED TO KNOT PULL TENSILE STRENGTH TEST TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. THE AVERAGE KNOT PULL TENSILE STRENGTH VALUE WAS FOUND TO MEETS THE IN-HOUSE AVERAGE KNOT PULL TENSILE STRENGTH REQUIREMENT. ALL THE INDIVIDUAL AS WELL AS AVERAGE KNOT PULL VALUES WERE FOUND TO MEET THE IN-HOUSE SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS. BATCH MANUFACTURING RECORDS OF NW2777/T8038 REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE, NEEDLE PULL-OFF VALUE AND MOISTURE TEST RESULTS AND FOUND TO MEET THE SPECIFICATION REQUIREMENT. THE FOLLOWING ADDITIONAL INFORMATION WAS RECEIVED: WERE ALL THE SUTURE STRANDS BREAKING WHILE USE ON THE PATIENT DURING SUTURING? YES. WAS PROCEDURE COMPLETED SUCCESSFULLY? AND HOW? YES. IT WAS EPISIOTOMY REPAIR,DOCTOR COMPLETE THE PROCEDURE WITH CATGUT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633797&pc=GAM | 2210968-2019-81713 |  | Malfunction | ETHICON INC. | 2019-05-22 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EPISIOTOMY REPAIR ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE DURING KNOTTING. A DIFFERENT DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: TO VERIFY THE COMPLAINT, RESULTS OF RETAINED SAMPLES WERE REVIEWED AND NO DISCREPANCY WAS OBSERVED FOR KNOT PULL TENSILE STRENGTH TESTING. RESULTS FOR KNOT PULL TEST WERE COMPLYING WITH THE PRE-DEFINED ACCEPTANCE CRITERIA. THE COMPLAINT IS RELATED TO PERFORMANCE-BREAKAGE SUTURE, KNOT PULL TENSILE STRENGTH TEST IS APPLICABLE AND IS A MEASURABLE PARAMETER. FIVE RETAIN SAMPLES WERE SUBJECTED TO KNOT PULL TENSILE STRENGTH TEST TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. THE AVERAGE KNOT PULL TENSILE STRENGTH VALUE OF THE RETAIN SAMPLE WAS FOUND TO MEETS THE IN-HOUSE AVERAGE KNOT PULL TENSILE STRENGTH REQUIREMENT. ALL THE INDIVIDUAL AS WELL AS AVERAGE KNOT PULL VALUES WERE FOUND TO MEET THE IN-HOUSE SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS. BATCH MANUFACTURING RECORDS OF NW2777/T8038 REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE, NEEDLE PULL-OFF VALUE AND MOISTURE TEST RESULTS AND FOUND TO MEET THE SPECIFICATION REQUIREMENT. THE FOLLOWING ADDITIONAL INFORMATION WAS RECEIVED: WERE ALL THE SUTURE STRANDS BREAKING WHILE USE ON THE PATIENT DURING SUTURING? YES. WAS PROCEDURE COMPLETED SUCCESSFULLY? AND HOW? YES. IT WAS EPISIOTOMY REPAIR,DOCTOR COMPLETE THE PROCEDURE WITH CATGUT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8634033&pc=GAM | 2210968-2019-81716 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-22 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EPISIOTOMY REPAIR ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE DURING USE. THE PROCEDURE WAS COMPLETED WITH A DIFFERENT SUTURE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). BATCH MANUFACTURING RECORDS OF NW2777/T8038 REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE, NEEDLE PULL-OFF VALUE AND MOISTURE TEST RESULTS AND FOUND TO MEET THE SPECIFICATION THE FOLLOWING ADDITIONAL INFORMATION WAS RECEIVED: WERE ALL THE SUTURE STRANDS BREAKING WHILE USE ON THE PATIENT DURING SUTURING? YES. WERE ALL THE SUTURE STRANDS THAT BROKE INTRA-OP FROM THE SAME LOT/BATCH? YES. WAS PROCEDURE COMPLETED SUCCESSFULLY? AND HOW? YES. IT WAS EPISIOTOMY REPAIR, DOCTOR COMPLETE THE PROCEDURE WITH CATGUT. TO VERIFY THE COMPLAINT, RESULTS OF RETAINED SAMPLES WERE REVIEWED AND NO DISCREPANCY WAS OBSERVED FOR KNOT PULL TENSILE STRENGTH TESTING. RESULTS FOR KNOT PULL TEST WERE COMPLYING WITH THE PRE-DEFINED ACCEPTANCE CRITERIA. THE COMPLAINT IS RELATED TO PERFORMANCE-BREAKAGE SUTURE, KNOT PULL TENSILE STRENGTH TEST IS APPLICABLE AND IS A MEASURABLE PARAMETER. FIVE RETAIN SAMPLES WERE SUBJECTED TO KNOT PULL TENSILE STRENGTH TEST TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. THE AVERAGE KNOT PULL TENSILE STRENGTH VALUE OF THE RETAIN SAMPLE WAS FOUND TO MEET THE IN-HOUSE AVERAGE KNOT PULL TENSILE STRENGTH REQUIREMENT. ALL THE INDIVIDUAL AS WELL AS AVERAGE KNOT PULL VALUES WERE FOUND TO MEET THE IN-HOUSE SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8628387&pc=GAM | 2210968-2019-81592 | 2019-03-10 | Malfunction | ETHICON INC. | 2019-05-21 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Degraded; Material Separation; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE DETACHED FROM THE NEEDLE DURING USE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE PULL OFF SUTURE NEEDLE. AN OPENED SAMPLE WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE OPENED SAMPLE, THE NEEDLE WAS DETACHED, DUE TO THE SUTURE HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE STRAND WAS NOTED BROKEN IN MULTIPLES SECTIONS. THE FOIL PRESET EXCESSIVE MANIPULATION, DELAMINATION AND MULTIPLES PIN HOLES CAUSED BY WRINKLES. THIS CONDITION CONTRIBUTED TO DEGRADATION OF THE SUTURES. THE PRODUCT CODE W9915 CONTAINS AN ABSORBABLE SUTURE AND AS THE SAMPLES WERE RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURES TO THE ENVIRONMENT COULD NOT BE DETERMINED AND FUNCTIONAL TEST CANNOT BE PERFORMED DUE TO SAMPLE CONDITION. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF PERFORMANCE PULL OFF SUTURE NEEDLE, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8628390&pc=GAM | 2210968-2019-81591 | 2019-03-10 | Malfunction | ETHICON INC. | 2019-05-21 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Degraded; Material Separation; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE DETACHED FROM THE NEEDLE DURING USE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE PULL OFF SUTURE NEEDLE. AN OPENED SAMPLE WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE OPENED SAMPLE, THE NEEDLE WAS DETACHED, DUE TO THE SUTURE HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE STRAND WAS NOTED BROKEN IN MULTIPLES SECTIONS. THE FOIL PRESET EXCESSIVE MANIPULATION, DELAMINATION AND MULTIPLES PIN HOLES CAUSED BY WRINKLES. THIS CONDITION CONTRIBUTED TO DEGRADATION OF THE SUTURES. THE PRODUCT CODE W9915 CONTAINS AN ABSORBABLE SUTURE AND AS THE SAMPLES WERE RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURES TO THE ENVIRONMENT COULD NOT BE DETERMINED AND FUNCTIONAL TEST CANNOT BE PERFORMED DUE TO SAMPLE CONDITION. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF PERFORMANCE PULL OFF SUTURE NEEDLE, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8629497&pc=GAM | 2210968-2019-81613 | 2019-04-29 | Malfunction | ETHICON INC. | 2019-05-21 | GAM | VICRYL RAPID 3/0 70CM 24MM FS1 | Material Frayed | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HEMORRHOIDECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. WHEN THE SURGEON STARTED TO STITCH, THE SUTURE TENDS TO UNRAVEL IN THE MIDDLE OF THE SUTURE IN BETWEEN SWAGE AND IN THE SUTURE. THE SURGEON OPENED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8630370&pc=GAM | 2210968-2019-81616 | 2019-04-29 | Malfunction | ETHICON INC. | 2019-05-21 | GAM | VICRYL RAPID 3/0 70CM 24MM FS1 | Material Frayed | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HEMORRHOIDECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. WHEN THE SURGEON STARTED TO STITCH, THE SUTURE TENDS TO UNRAVEL IN THE MIDDLE OF THE SUTURE IN BETWEEN SWAG AND IN THE SUTURE. THE SURGEON OPENED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8630381&pc=GAM | 2210968-2019-81617 | 2019-04-29 | Malfunction | ETHICON INC. | 2019-05-21 | GAM | VICRYL RAPID 3/0 70CM 24MM FS1 | Material Frayed | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HEMORRHOIDECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. WHEN THE SURGEON STARTED TO STITCH, THE SUTURE TENDS TO UNRAVEL IN THE MIDDLE OF THE SUTURE IN BETWEEN SWAG AND IN THE SUTURE. THE SURGEON OPENED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8630384&pc=GAM | 2210968-2019-81618 | 2019-04-15 | Malfunction | ETHICON INC. | 2019-05-21 | GAM | VICRYL PLUS SUTURE 27"(70CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING SEWING UTERUS THE NEEDLE PULLED OFF THE SUTURE. THE NEEDLE WAS RETRIEVED FROM THE PATIENT. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8624015&pc=GAM | 2210968-2019-81518 | 2019-04-25 | Malfunction | ETHICON INC. | 2019-05-20 | GAM | VICRYL CT UD 30IN USP2-0 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL KNEE REPLACEMENT PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BENT AND BROKE DURING USE. THE PROCEDURE WAS NOT DELAYED AND WAS COMPLETED SUCCESSFULLY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE PRODUCT (OR PHOTO) UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBERS: BATCH LLM083, BATCH MLZ524. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. CONFIRM THE EVENT - DID THE SAME SINGLE REPORTED NEEDLE FIRST BEND AND THEN BREAK? OR 2 NEEDLES INVOLVED WHERE 1 NEEDLE BENT AND OTHER BROKE? CONFIRM THE LOT #'S FOR BOTH NEEDLES? REGARDING THE BROKEN NEEDLE, PLEASE ANSWER THE FOLLOWING: DID THE BROKEN NEEDLE OR NEEDLE PIECE FALL INTO PATIENT? IF YES, WAS IT REMOVED? AND HOW? ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THE EVENT REPORT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8624846&pc=GAM | 2210968-2019-81530 | 2019-04-24 | Malfunction | ETHICON INC. | 2019-05-20 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A MOHS PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE SEPARATED FROM THE SUTURE DURING USE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8625072&pc=GAM | 2210968-2019-81533 | 2019-04-24 | Malfunction | ETHICON INC. | 2019-05-20 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A MOHS PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE SEPARATED FROM THE SUTURE DURING USE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8625434&pc=GAM | 2210968-2019-81537 | 2019-04-25 | Malfunction | ETHICON INC. | 2019-05-20 | GAM | VICRYL VIO 75CM 2/0 SH PLUS | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING USE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MG8CJJB0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8625442&pc=GAM | 2210968-2019-81538 | 2019-04-25 | Malfunction | ETHICON INC. | 2019-05-20 | GAM | VICRYL VIO 75CM 2/0 SH PLUS | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING USE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MG8CJJB0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8626273&pc=GAM | 2210968-2019-81546 | 2019-02-15 | Malfunction | ETHICON INC. | 2019-05-20 | GAM | VICRYL SUTURE 27"(70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE PULLED OFF DURING THE PROCEDURE. THE FALLEN PIECE WAS RETRIEVED FROM THE PATIENT. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WAS NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8626322&pc=GAM | 2210968-2019-81547 | 2019-04-26 | Malfunction | ETHICON INC. | 2019-05-20 | GAM | VICRYL ABS SUR SUT USP SYN | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE SURGERY, THE SUTURE THREAD WAS PULLED OFF FROM THE SWAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8620474&pc=GAM | 2210968-2019-81482 | 2019-04-19 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED BY HEALTH AUTHORITY THAT THE PATIENT UNDERWENT ORTHOPEDIC PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE AND WAS RETRIEVED FROM THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8620762&pc=GAM | 2210968-2019-81486 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL SUTURE 27"(70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE POPPED OFF THE NEEDLE. NEEDLE PULLED OFF THE SUTURE AND WAS RETRIEVED FROM THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MA2717, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8620786&pc=GAM | 2210968-2019-81488 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL SUTURE 27"(70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE POPPED OFF THE NEEDLE. NEEDLE PULLED OFF THE SUTURE AND WAS RETRIEVED FROM THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MA2717, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8620819&pc=GAM | 2210968-2019-81489 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS SUTURE 18"(45CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE POPPED OFF THE NEEDLE. NEEDLE PULLED OFF THE SUTURE AND WAS RETRIEVED FROM THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8620860&pc=GAM | 2210968-2019-81490 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS SUTURE 18"(45CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE POPPED OFF THE NEEDLE. NEEDLE PULLED OFF THE SUTURE AND WAS RETRIEVED FROM THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621286&pc=GAM | 2210968-2019-81492 | 2019-04-24 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A MOHS PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE SEPARATED FROM THE SUTURE WHEN USED ON THE SKIN. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621386&pc=GAM | 2210968-2019-81496 | 2019-04-24 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS VIO BR 18 | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT LAMINECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING INTERRUPTED SUTURING ON THE FASCIA. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621388&pc=GAM | 2210968-2019-81495 | 2019-04-24 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS VIO BR 18 | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT LAMINECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY DURING INTERRUPTED SUTURING ON THE FASCIA. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621488&pc=GAM | 2210968-2019-81497 | 2019-04-25 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE WAS DETACHED FROM THE SUTURE DURING USE ON THE FASCIA. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. RELATED DEVICES REPORTED IN MW 2210968-2019-81497, 2210968-2019-81498, 2210968-2019-81499.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621499&pc=GAM | 2210968-2019-81499 | 2019-04-25 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE WAS DETACHED FROM THE SUTURE DURING USE ON THE FASCIA. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. RELATED DEVICES REPORTED IN MW 2210968-2019-81497, 2210968-2019-81498, 2210968-2019-81499.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621630&pc=GAM | 2210968-2019-81501 | 2019-04-23 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE DETACHED FROM THE SUTURE DURING INTERRUPTED SUTURING. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MMM167 BATCH NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621641&pc=GAM | 2210968-2019-81502 | 2019-04-23 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE DETACHED FROM THE SUTURE DURING INTERRUPTED SUTURING. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MMM167 BATCH NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8623682&pc=GAM | MW5086715 | 2019-05-13 | Malfunction | ETHICON, INC. | 2019-05-17 | GAM | VICRYL SUTURE | Manufacturing, Packaging or Shipping Problem | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: WHEN PACKAGE WAS OPENED THERE WERE NINE NEEDLES INSIDE. PACKAGING READS EIGHT NEEDLES ENCLOSED. FDA SAFETY REPORT ID# (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8624113&pc=GAM | MW5086747 | 2019-04-23 | Malfunction | ETHICON INC. | 2019-05-17 | GAM | VICRYL PLUS | Packaging Problem | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: PACKAGE CONTAINED INCORRECT AMOUNT OF NEEDLES. PACKAGING LISTS 8, BUT THERE WERE 9 IN THE PACKAGE. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8618450&pc=GAM | 2210968-2019-81458 | 2019-04-09 | Malfunction | ETHICON INC. | 2019-05-16 | GAM | VICRYL SUTURE 27"(70CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE BROKE IN THE CAVITY DURING THE SURGICAL PROCEDURE. IT WAS NECESSARY TO USE AN IMAGE INTENSIFIER TO RETRIEVE THE PIECE THAT BROKE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: (B)(4).. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611461&pc=GAM | 2210968-2019-81397 |  | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611672&pc=GAM | 2210968-2019-81401 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL POLYGLACTIN 910 SUTURE - NO DETAILED DESC | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE WAS DETACHED FROM THE NEEDLE DURING USE. IT HAPPENED SEVERAL TIMES DURING PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE ADVISE THE NUMBER OF SUTURES THAT WERE USED AND DETACHED FROM THE NEEDLE DURING THE PROCEDURE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8612206&pc=GAM | 2210968-2019-81403 | 2019-01-21 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL CT VIO 30IN USP2-0 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS WEAK AND EASY TO BREAK. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ATTEMPTS WERE MADE TO OBTAIN THE FOLLOWING INFORMATION, BUT NO FURTHER INFORMATION WAS AVAILABLE. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT WHAT IS THE SPECIFIC NUMBER OF DEVICES THAT FAILED DURING THE 1 PROCEDURE? N/A. THE REPORTED QUANTITY INVOLVED WAS 48. PLEASE CONFIRM THAT 48 DEVICES EXPERIENCED THE QUALITY ISSUE. SUTURES BROKE DURING MANY PROCEDURES. IF 48 DEVICES DID NOT EXPERIENCE THE QUALITY ISSUE, PLEASE CLARIFY THE NUMBER OF DEVICES THAT DID EXPERIENCE THE QUALITY ISSUE, N/A. DID THE EVENT OCCUR DURING ONE OR MULTIPLE PATIENT PROCEDURES? MULTIPLE PROCEDURES. IF THIS EVENT OCCURRED IN MULTIPLE PROCEDURES, PLEASE PROVIDE THE FOLLOWING INFORMATION: WHAT IS THE TOTAL NUMBER OF PROCEDURES? N/A. HAVE ANY OF THESE EVENTS BEEN PREVIOUSLY REPORTED TO ETHICON? IF SO, PROVIDE THE RESPECTIVE REFERENCE NUMBER(S), NO. WHAT ARE THE PROCEDURE NAME(S) AND DATE(S)? N/A. WHAT IS THE QUANTITY INVOLVED PER PROCEDURE? N/A. WAS THERE ANY ADVERSE PATIENT CONSEQUENCE(S) OR SUBSEQUENT MEDICAL/SURGICAL INTERVENTION? NO PATIENT CONSEQUENCE(S). WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? UNKNOWN. WAS PROCEDURE SUCCESSFULLY COMPLETED? UNKNOWN. WERE FRAGMENTS GENERATED? UNKNOWN. IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? UNKNOWN. PATIENT STATUS/ OUTCOME / CONSEQUENCES, NO. WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN. IMPACT PRODUCT. PRODUCT LOT NO. ME6529. QUANTITY OF PRODUCT INVOLVED : 1.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8612230&pc=GAM | 2210968-2019-81402 | 2019-01-21 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL CT VIO 30IN USP2-0 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS WEAK AND EASY TO BREAK. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ATTEMPTS WERE MADE TO OBTAIN THE FOLLOWING INFORMATION, BUT NO FURTHER INFORMATION WAS AVAILABLE. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT THE REPORTED QUANTITY INVOLVED WAS 48. PLEASE CONFIRM THAT 48 DEVICES EXPERIENCED THE QUALITY ISSUE. SUTURES BROKE DURING MANY PROCEDURES. IF 48 DEVICES DID NOT EXPERIENCE THE QUALITY ISSUE, PLEASE CLARIFY THE NUMBER OF DEVICES THAT DID EXPERIENCE THE QUALITY ISSUE. DID THE EVENT OCCUR DURING ONE OR MULTIPLE PATIENT PROCEDURES? MULTIPLE PROCEDURES. IF THIS EVENT OCCURRED IN MULTIPLE PROCEDURES, PLEASE PROVIDE THE FOLLOWING INFORMATION: HAVE ANY OF THESE EVENTS BEEN PREVIOUSLY REPORTED TO ETHICON? IF SO, PROVIDE THE RESPECTIVE REFERENCE NUMBER(S). NO . WAS THERE ANY ADVERSE PATIENT CONSEQUENCE(S) OR SUBSEQUENT MEDICAL/SURGICAL INTERVENTION? NO PATIENT CONSEQUENCE(S). WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? UNKNOWN. WAS PROCEDURE SUCCESSFULLY COMPLETED: UNKNOWN, WERE FRAGMENTS GENERATED? UNKNOWN, IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES: NO, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN, IMPACT PRODUCT: PRODUCT LOT NO. MKM443. QUANTITY OF PRODUCT INVOLVED : 1.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8612433&pc=GAM | 2210968-2019-81406 | 2019-04-18 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PEDIATRIC HERNIA SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF. THE FALLEN PIECE WAS RETRIEVED FROM THE PATIENT. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE IS NO REPORT ON PATIENT'S INJURY. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8612441&pc=GAM | 2210968-2019-81407 | 2019-04-18 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PEDIATRIC HERNIA SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF. THE FALLEN PIECE WAS RETRIEVED FROM THE PATIENT. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE IS NO REPORT ON PATIENT'S INJURY. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8612572&pc=GAM | 2210968-2019-81410 | 2019-04-22 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL PLUS SUTURE 18"(45CM) 0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE TIP WAS BROKEN DURING INTERRUPTED SUTURING. THE BROKEN PIECE OF NEEDLE WAS FOUND AND DISCARDED. ANOTHER PRODUCT WAS USED FOR THE PATIENT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8612776&pc=GAM | 2210968-2019-81419 | 2019-04-19 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT TRANSABDOMINAL PREPERITONEAL APPROACH PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING CONTINUOUS SUTURING ON THE PERITONEUM, THE NEEDLE BROKE AT THE SWAGE PART WHEN IT WAS GRASPED. THE NEEDLE FRAGMENT WAS REMOVED WITH FORCEPS WITHOUT LOSING SIGHT OF THE NEEDLE FRAGMENT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD REVIEW WAS PERFORMED FOR THE FINISHED DEVICE BATCH AND NO NON-CONFORMANCE WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8613081&pc=GAM | 2210968-2019-81422 | 2019-04-19 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL SUTURE 27"(70CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN OPEN COLECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE AT THE MIDDLE PART DURING WOUND CLOSURE BY CONTINUOUS SUTURING. THE BROKEN PIECE OF NEEDLE WAS NOT DROPPED IN THE PATIENT BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: INVESTIGATION SUMMARY: "IT WAS REPORTED PERFORMANCE NEEDLE BREAKAGE. THE ACTUAL DEVICE WAS RETURNED FOR EVALUATION. A FAILED SUTURE NEEDLE, LABELED AS (B)(4), WAS IDENTIFIED AS PRODUCT CODE J317H. A FRACTURE WAS OBSERVED IN THE BODY OF THE NEEDLE. THE NEEDLE WAS RECEIVED IN TWO PIECES, ONLY ONE OF THE MATING FRACTURE SURFACES WAS EXAMINED FOR THIS EVALUATION. A SCANNING ELECTRON MICROSCOPE (SEM) WAS USED TO EXAMINE THE FRACTURE SURFACES AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACES WERE EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION OF (B)(4) REVEALED THE FRACTURE WAS PREDOMINANTLY INTERGRANULAR, WHICH IS EVIDENCE OF A BRITTLE FAILURE. THIS WAS A NON-DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE NEEDLE EXHIBITED AMOUNTS OF INTERGRANULAR FAILURE. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LJ2795, AND NO NONCONFORMANCES WERE IDENTIFIED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8613349&pc=GAM | 2210968-2019-81430 | 2019-03-17 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED BY HEALTH AUTHORITY THAT THE PATIENT UNDERWENT A FRONTAL AVULSION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING SKIN CLOSURE. THE BROKEN PIECE DID NOT FALL INSIDE THE PATIENT. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8613491&pc=GAM | 2210968-2019-81431 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS COMING OFF OF THE NEEDLE. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8613537&pc=GAM | 2210968-2019-81432 |  | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE WAS FOUND DEFECTIVE DURING THE PROCEDURE. THE DEVICE WAS NOT USED ON THE PATIENT. THERE WERE NO REPORTED ADVERSE CONSEQUENCES TO THE PATIENT. DURING EVALUATION, A FRACTURE WAS OBSERVED AT THE TIP OF THE NEEDLE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: REPRESENTATIVE SAMPLE: AN UNOPENED SAMPLE OF PRODUCT CODE J370, LOT MKK815 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLE WAS OPENED AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED; ALSO, THE TIP AND BODY OF THE NEEDLE WERE EXAMINED UNDER MAGNIFICATION AND NO DEFECTS OR DAMAGE WERE FOUND. ALSO, THE SUTURE WAS DISPENSED WITHOUT PROBLEM AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WERE NOTED. THE ACTUAL DEVICE WAS RECEIVED FOR EVALUATION: A FRACTURE WAS OBSERVED AT THE TIP OF THE NEEDLE. ONE SIDE OF THE NEEDLE WAS RECEIVED, THE MATING FRACTURE SURFACE WAS NOT PROVIDED FOR THIS EVALUATION. THE FRACTURE SURFACES WERE MICROSCOPICALLY EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. MECHANICAL DAMAGE OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDES ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE TIP OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. THE NEEDLE BREAKAGE WAS CONFIRMED. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER MKK815-J370, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8613618&pc=GAM | 2210968-2019-81434 | 2019-04-23 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL SUTURE 27"(70CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN OB-GYN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE WAS DETACHED FROM THE SUTURE DURING CONTINUOUS SUTURING OF THE UTERUS. THERE WERE NO REPORTED ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MMK993 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8613620&pc=GAM | 2210968-2019-81433 | 2019-04-23 | Malfunction | ETHICON INC. | 2019-05-15 | GAM | VICRYL SUTURE 27"(70CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN OB-GYN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE WAS DETACHED FROM THE SUTURE DURING CONTINUOUS SUTURING OF THE UTERUS. THERE WERE NO REPORTED ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MMK993 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8607313&pc=GAM | 2210968-2019-81311 |  | Malfunction | ETHICON INC. | 2019-05-14 | GAM | VICRYL SUTURE 18"(45CM) 4-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE OFF QUITE EASILY FROM THE SUTURE. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LEM799 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609112&pc=GAM | 2210968-2019-81382 | 2019-04-17 | Malfunction | ETHICON INC. | 2019-05-14 | GAM | VICRYL SUTURE 18"(45CM) 1 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE TIP OF THE NEEDLE BROKE DURING USE. THE TIP OF THE NEEDLE WAS NOT FOUND. A POST OP X-RAY WAS PERFORMED AND NO FOREIGN OBJECTS WERE FOUND. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609184&pc=GAM | 2210968-2019-81383 | 2019-04-17 | Malfunction | ETHICON INC. | 2019-05-14 | GAM | VICRYL SUTURE 18"(45CM) 1 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE TIP OF THE NEEDLE BROKE WHEN CLOSING THE INCISION. THE TIP OF THE NEEDLE WAS NOT FOUND. A POST OP X-RAY WAS PERFORMED AND UNKNOWN IF FOREIGN OBJECTS WERE FOUND. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609284&pc=GAM | 2210968-2019-81384 | 2019-04-17 | Malfunction | ETHICON INC. | 2019-05-14 | GAM | VICRYL CT UD 18IN USP4-0 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN INGUINAL HERNIA REPAIR PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE TIP OF THE NEEDLE BROKE. IT WAS REPORTED THAT THE NEEDLE MAY HAVE BROKEN WHEN THE NEEDLE WAS PUT BACK INTO THE NEEDLE HOLDER. THE TIP OF THE NEEDLE DID NOT FALL INTO THE PATIENT. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609561&pc=GAM | 2210968-2019-81390 | 2019-04-05 | Malfunction | ETHICON INC. | 2019-05-14 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 UND | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE PRIMARY PACKAGE WAS FOUND DAMAGED AND IT WAS REPORTED THAT STERILITY WAS COMPROMISED. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW WAS THE PACKAGE DAMAGED? UNK. WAS STERILITY COMPROMISED? YES. WAS THERE DAMAGE TO CARTON? NO. WAS THERE DAMAGE TO FOLDER/TRAY? UNK. WAS THERE DAMAGE TO THE PRIMARY PACKAGE? YES. WAS THERE ANY HOLE, TEAR, OR PUNCTURE THAT COMPROMISES STERILITY? NOT SURE. WHAT IS THE LOT NUMBER? UNK.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609727&pc=GAM | 2210968-2019-81359 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-14 | GAM | VICRYL SUTURE 18"(45CM) 4-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT SUBCUTANEOUS PORT PLACEMENT OR SUBCUTANEOUS PORT REMOVAL ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS PREMATURELY POPPING OFF. THE CASE WAS COMPLETED WITH ANOTHER DEVICE FROM A DIFFERENT LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CONFIRM THE SPECIFIC NUMBER OF PATIENT EVENTS/ PROCEDURES. HAVE ANY OF THESE EVENTS/PROCEDURES BEEN PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE COMPLAINT FILE NUMBER/ AFFILIATE NUMBER. FOR EACH PATIENT EVENT PLEASE PROVIDE ADDITIONAL COMPLAINT DETAILS: EVENT DATE, PRODUCT CODE AND LOT NUMBER INVOLVED, PROCEDURE, EVENT DESCRIPTION, QTY INVOLVED? WAS THERE ANY ADVERSE PATIENT OUTCOME? IF YES, WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED, WAS PROCEDURE COMPLETED SUCCESSFULLY AND HOW? ANY SAMPLE RETURN, ETC? EVENT DESCRIPTION, EVENT OCCURRENCE, PRODUCT CODE AND LOT NUMBER, EVENT DATE, PROCEDURE NAME. PATIENT CONSEQUENCES/AE REPORT IF ANY? ANY MEDICAL/SURGICAL INTERVENTION OR TREATMENT PROVIDED TO PATIENT POST-OP? TOTAL QTY OF SUTURES INVOLVED? SAMPLE RETURN, IF ANY?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8603182&pc=GAM | 2210968-2019-81300 | 2019-04-15 | Malfunction | ETHICON INC. | 2019-05-13 | GAM | VICRYL SUTURE 27"(70CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CONTINUOUS AMBULATORY PERITONEAL DIALYSIS PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS BROKEN AT THE SWAGE PART DURING INTERRUPTED SUTURING ON THE FASCIA. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: THE ACTUAL DEVICE WAS RECEIVED FOR EVALUATION. A FRACTURE WAS OBSERVED AT THE SUTURE ATTACHMENT OF THE NEEDLE. THE NEEDLE WAS RECEIVED IN TWO PIECES, ONLY ONE OF THE MATING FRACTURE SURFACES WAS EXAMINED FOR THIS EVALUATION. A MICROSCOPE WAS USED TO EXAMINE THE FRACTURE SURFACES AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACES WERE EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE ATTACHMENT AREA OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8603707&pc=GAM | 2210968-2019-81312 | 2019-03-14 | Malfunction | ETHICON INC. | 2019-05-13 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Break | Unspecified Infection | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. THREE DAYS POSTOPERATIVE, THE PATIENT EXPERIENCED INCISION INFECTION AND SPLIT. ALCOHOL WAS TREATED FOR INCISION AND PATIENT CONDITION CHANGED BETTER. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ME2850, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT DID YOU MEAN BY "SPLIT."? PLEASE CLARIFY- DID THE SUTURE BREAK POST-OP? WAS THERE ANY DEHISCENCE NOTICED? HOW WAS THE PATIENT TREATED FOR INFECTION AND SPLIT? WAS ANY PRESCRIBED MEDICATION / ANTIBIOTICS GIVEN TO PATIENT POST-OP? WAS ANY RE-OPERATION/RE-SUTURING/ REVISION SURGERY DONE ON PATIENT POST-OP INITIAL PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8603987&pc=GAM | 2210968-2019-81316 | 2019-03-26 | Malfunction | ETHICON INC. | 2019-05-13 | GAM | VICRYL* VLT 3-0 70CM (1)TE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC COLPOPROMONTOFIXATION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGICAL THREAD RUPTURED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8604525&pc=GAM | 2210968-2019-81321 | 2019-04-17 | Malfunction | ETHICON INC. | 2019-05-13 | GAM | VICRYL CT VIO 8X18IN 3-0 S/A SH CR | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A STOMA CREATION SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS BROKEN DURING LIGATION ON THE SUBCUTIS BY INTERRUPTED SUTURING. ANOTHER LIKE DEVICE WAS USED FOR THE PATIENT WITH NO PROBLEM. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8604647&pc=GAM | 2210968-2019-81323 | 2019-04-14 | Malfunction | ETHICON INC. | 2019-05-13 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2019 AND SUTURE WAS USED ON THE PERITONEUM. DURING THE PROCEDURE, THE SUTURE BROKE TWO TIMES WHILE SUTURING THE PERITONEUM. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8604733&pc=GAM | 2210968-2019-81327 | 2019-04-17 | Malfunction | ETHICON INC. | 2019-05-13 | GAM | VICRYL VIO 75CM M1.5 4/0 SH-2 PLUS | Break; Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DENTAL CLEARANCE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN SUTURING THROUGH GUM, THE SUTURE BECOME DEBRAIDED AND THE SUTURE SNAPPED WHEN DOING KNOT TYING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PHOTOS HAVE BEEN RECEIVED FOR ANALYSIS. ANALYSIS NOT YET COMPLETE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8605053&pc=GAM | 2210968-2019-81333 | 2019-04-15 | Malfunction | ETHICON INC. | 2019-05-13 | GAM | VICRYL SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE CAME OFF THE END OF THE SUTURE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8600492&pc=GAM | 2210968-2019-81280 | 2019-04-15 | Malfunction | ETHICON INC. | 2019-05-10 | GAM | VICRYL SUTURE 27"(70CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT LAPAROSCOPIC PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. WHEN COMPLETING THE PROCEDURE, THE NEEDLE BROKE APPROXIMATELY 5MM FROM THE SWAGE END. THE NEEDLE WAS RETRIEVED VIA XRAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8601264&pc=GAM | 2210968-2019-81286 | 2019-04-12 | Malfunction | ETHICON INC. | 2019-05-10 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE TIP BROKE DURING INTERRUPTED SUTURING. THE BROKEN NEEDLE TIP DROPPED INTO THE PATIENT. THE PROCEDURE WAS EXTENDED 30 MINUTES TO FIND THE NEEDLE. THE NEEDLE WAS REMOVED FROM THE PATIENT DURING THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). IT WAS REPORTED PERFORMANCE NEEDLE BREAKAGE. A FAILED SUTURE NEEDLE WAS SUBMITTED FOR EVALUATION. A FRACTURE WAS OBSERVED NEAR THE TIP OF THE NEEDLE. THE NEEDLE WAS RECEIVED IN TWO PIECES, ONLY ONE OF THE MATING FRACTURE SURFACES WAS EXAMINED FOR THIS EVALUATION. A MICROSCOPE WAS USED TO EXAMINE THE FRACTURE SURFACES AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACES WERE EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. MECHANICAL DAMAGE AND MACROSCOPIC PLASTIC DEFORMATION OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDE ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS COULDN'T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE TIP OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8601579&pc=GAM | 2210968-2019-81295 | 2019-03-29 | Malfunction | ETHICON INC. | 2019-05-10 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Break | Wound Dehiscence; Unspecified Infection; Impaired Healing | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED INCISION INFECTED AND SPLIT 3 DAYS AFTER C-SECTION. ALCOHOL WAS USED TO TREAT FOR INCISION AND THE PATIENT CONDITION CHANGED TO BETTER. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ME2850, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT DID YOU MEAN BY "SPLIT."? PLEASE CLARIFY- DID THE SUTURE BREAK POST-OP? WAS THERE ANY DEHISCENCE NOTICED? HOW WAS THE PATIENT TREATED FOR INFECTION AND SPLIT? WAS ANY PRESCRIBED MEDICATION / ANTIBIOTICS GIVEN TO PATIENT POST-OP? WAS ANY RE-OPERATION/RE-SUTURING/ REVISION SURGERY DONE ON PATIENT POST-OP INITIAL PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8596577&pc=GAM | 2210968-2019-81224 | 2019-04-11 | Malfunction | ETHICON INC. | 2019-05-09 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT EPISIOTOMY REPAIR ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE WHEN KNOTTING. THE PROCEDURE WAS COMPLETED WITH A DIFFERENT SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). BATCH MANUFACTURING RECORDS OF NW2777/T8038 REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. THE FOLLOWING ADDITIONAL INFORMATION WAS RECEIVED: WERE ALL THE SUTURE STRANDS BREAKING WHILE USE ON THE PATIENT DURING SUTURING? YES. WAS PROCEDURE COMPLETED SUCCESSFULLY? AND HOW? YES. IT WAS EPISIOTOMY REPAIR, DOCTOR COMPLETE THE PROCEDURE WITH CATGUT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8597255&pc=GAM | 2210968-2019-81247 | 2019-04-08 | Malfunction | ETHICON INC. | 2019-05-09 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GYNECOLOGICAL TROCAR PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE DURING USE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MHM161 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8597286&pc=GAM | 2210968-2019-81249 | 2019-04-08 | Malfunction | ETHICON INC. | 2019-05-09 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GYNECOLOGICAL UTERINE FIBROID PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING USE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJ2948, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8591593&pc=GAM | 2210968-2019-81167 |  | Malfunction | ETHICON INC. | 2019-05-08 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DERMATOLOGY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE PULLED OFF WHILE BEING USED BY THE PHYSICIAN TO SUTURE A PATIENT. THE NEEDLE HAD TO BE EXCISED FROM THE PATIENTS SKIN SINCE THE SUTURE BROKE OFF. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS MEANT BY THE NEEDLE WAS EXCISED FROM PATIENT? INITIAL PROCEDURE DATE PROCEDURE NAME WHEN WAS THE NEEDLE EXCISED FROM THE PATIENT (DURING THE INITIAL SURGERY OR WAS THE PATIENT BROUGHT BACK TO OR)? LOCATION THAT THE NEEDLE WAS LOCATED? SIZE OF INCISION TO REMOVE THE NEEDLE? WAS THERE ANY ADVERSE PATIENT OUTCOME? WHAT IS MEANT BY IN ANOTHER INSTANCE THE SUTURE BROKE OFF ? DID THE NEEDLE PULL OFF THE SUTURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8591622&pc=GAM | 2210968-2019-81178 |  | Malfunction | ETHICON INC. | 2019-05-08 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DERMATOLOGY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE PULLED OFF WHILE BEING USED BY THE PHYSICIAN TO SUTURE A PATIENT. THE NEEDLE HAD TO BE EXCISED FROM THE PATIENTS SKIN SINCE THE SUTURE BROKE OFF. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS MEANT BY THE NEEDLE WAS EXCISED FROM PATIENT? INITIAL PROCEDURE DATE PROCEDURE NAME WHEN WAS THE NEEDLE EXCISED FROM THE PATIENT (DURING THE INITIAL SURGERY OR WAS THE PATIENT BROUGHT BACK TO OR)? LOCATION THAT THE NEEDLE WAS LOCATED? SIZE OF INCISION TO REMOVE THE NEEDLE? WAS THERE ANY ADVERSE PATIENT OUTCOME? WHAT IS MEANT BY IN ANOTHER INSTANCE THE SUTURE BROKE OFF ? DID THE NEEDLE PULL OFF THE SUTURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8591884&pc=GAM | 2210968-2019-81182 |  | Malfunction | ETHICON INC. | 2019-05-08 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DERMATOLOGY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE PULLED OFF WHILE BEING USED BY THE PHYSICIAN TO SUTURE A PATIENT. THE NEEDLE MAY HAVE BEEN EXCISED FROM THE PATIENTS SKIN. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS MEANT BY THE NEEDLE WAS EXCISED FROM PATIENT? INITIAL PROCEDURE DATE. PROCEDURE NAME. WHEN WAS THE NEEDLE EXCISED FROM THE PATIENT (DURING THE INITIAL SURGERY OR WAS THE PATIENT BROUGHT BACK TO OR)? LOCATION THAT THE NEEDLE WAS LOCATED? SIZE OF INCISION TO REMOVE THE NEEDLE? WAS THERE ANY ADVERSE PATIENT OUTCOME? WHAT IS MEANT BY IN ANOTHER INSTANCE THE SUTURE BROKE OFF ? DID THE NEEDLE PULL OFF THE SUTURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8592165&pc=GAM | 2210968-2019-81188 | 2019-02-28 | Malfunction | ETHICON INC. | 2019-05-08 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT DENTAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE THREAD BROKE WHEN SUTURING TWICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8592844&pc=GAM | 2210968-2019-81199 | 2019-01-24 | Malfunction | ETHICON INC. | 2019-05-08 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EPISIOTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF DURING SEWING THE PERINEUM. THE NEEDLE WAS REMOVED FROM THE PATIENT DURING THE PROCEDURE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE KG7304, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586470&pc=GAM | 2210968-2019-81119 | 2019-04-09 | Malfunction | ETHICON INC. | 2019-05-07 | GAM | VICRYL SUTURE 27"(70CM) 4-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE WAS BROKEN AT THE SWAGE PART DURING USE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586569&pc=GAM | 2210968-2019-81120 |  | Malfunction | ETHICON INC. | 2019-05-07 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AND AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS DETACHED FROM THE SUTURE SOON AFTER PASSING THE NEEDLE THROUGH THE TISSUE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8587067&pc=GAM | 2210968-2019-81123 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-07 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A MOHS SURGICAL EXCISION OF A NON-MALIGNANT SKIN CANCER IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, UPON CLOSURE OF A SURGICAL EXCISION, THE NEEDLE DETACHED FROM THE SUTURE. THE SURGEON WAS ABLE TO RETRIEVE THE NEEDLE FROM THE PATIENT AFTER THE SUTURE THREAD SEPARATED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THIS OCCUR DURING ONE PATIENT PROCEDURE? - YES. NAME OF PROCEDURE? - MOHS SURGICAL EXCISION OF A NON-MALIGNANT SKIN CANCER. WILL ACTUAL DEVICE BE RETURNED FOR EVALUATION? - YES, WE HAVE 4 EXAMPLES OF WHERE THE THREAD PULLED OFF THE NEEDLE DURING A CLOSURE OF A SURGICAL EXCISION. PLEASE, PROVIDE US WITH THE PROPER ADDRESS AND METHOD OF RETURNING TO YOU THE DEFECTIVE MATERIAL. WERE THERE ADVERSE PATIENT CONSEQUENCES? - FORTUNATELY, THERE WERE NO PATIENT CONSEQUENCES. OUR SURGEON WAS ABLE TO RETRIEVE THE NEEDLE FROM THE PATIENT AFTER THE SUTURE THREAD SEPARATED FROM THE NEEDLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8587070&pc=GAM | 2210968-2019-81122 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-07 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A MOHS SURGICAL EXCISION OF A NON-MALIGNANT SKIN CANCER IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, UPON CLOSURE OF A SURGICAL EXCISION, THE NEEDLE DETACHED FROM THE SUTURE. THE SURGEON WAS ABLE TO RETRIEVE THE NEEDLE FROM THE PATIENT AFTER THE SUTURE THREAD SEPARATED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: AN EMPTY OPENED FOIL, A LABELED WINDING FORMER, A DETACHED NEEDLE OF PRODUCT CODE J497, LOT MJK059 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED, BUT MARKS ON THE EDGE NEEDLE THAT APPEAR TO BE BY USE OF A SURGICAL INSTRUMENT COLD BE OBSERVED AND A SUTURE PIECE STILL WAS ATTACHED TO NEEDLE. ACCORDING TO CONDITION OF DETACHED NEEDLE RECEIVED, THE ASSIGNABLE CAUSE OF THE PERFORMANCE PULL OFF SUTURE NEEDLE WAS AN IMPROPER HANDLING. TO AVOID THIS KIND OF DAMAGE, GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THIS OCCUR DURING ONE PATIENT PROCEDURE? YES. NAME OF PROCEDURE? MOHS SURGICAL EXCISION OF A NON-MALIGNANT SKIN CANCER. WILL ACTUAL DEVICE BE RETURNED FOR EVALUATION? YES, WE HAVE 4 EXAMPLES OF WHERE THE THREAD PULLED OFF THE NEEDLE DURING A CLOSURE OF A SURGICAL EXCISION. PLEASE, PROVIDE US WITH THE PROPER ADDRESS AND METHOD OF RETURNING TO YOU THE DEFECTIVE MATERIAL. WERE THERE ADVERSE PATIENT CONSEQUENCES? FORTUNATELY, THERE WERE NO PATIENT CONSEQUENCES. OUR SURGEON WAS ABLE TO RETRIEVE THE NEEDLE FROM THE PATIENT AFTER THE SUTURE THREAD SEPARATED FROM THE NEEDLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8587071&pc=GAM | 2210968-2019-81124 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-07 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A MOHS SURGICAL EXCISION OF A NON-MALIGNANT SKIN CANCER IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, UPON CLOSURE OF A SURGICAL EXCISION, THE NEEDLE DETACHED FROM THE SUTURE. THE SURGEON WAS ABLE TO RETRIEVE THE NEEDLE FROM THE PATIENT AFTER THE SUTURE THREAD SEPARATED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THIS OCCUR DURING ONE PATIENT PROCEDURE? YES . NAME OF PROCEDURE? MOHS SURGICAL EXCISION OF A NON-MALIGNANT SKIN CANCER. WILL ACTUAL DEVICE BE RETURNED FOR EVALUATION? YES, WE HAVE 4 EXAMPLES OF WHERE THE THREAD PULLED OFF THE NEEDLE DURING A CLOSURE OF A SURGICAL EXCISION. PLEASE, PROVIDE US WITH THE PROPER ADDRESS AND METHOD OF RETURNING TO YOU THE DEFECTIVE MATERIAL. WERE THERE ADVERSE PATIENT CONSEQUENCES? FORTUNATELY, THERE WERE NO PATIENT CONSEQUENCES. OUR SURGEON WAS ABLE TO RETRIEVE THE NEEDLE FROM THE PATIENT AFTER THE SUTURE THREAD SEPARATED FROM THE NEEDLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8587074&pc=GAM | 2210968-2019-81121 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-07 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A MOHS SURGICAL EXCISION OF A NON-MALIGNANT SKIN CANCER IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, UPON CLOSURE OF A SURGICAL EXCISION, THE NEEDLE DETACHED FROM THE SUTURE. THE SURGEON WAS ABLE TO RETRIEVE THE NEEDLE FROM THE PATIENT AFTER THE SUTURE THREAD SEPARATED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THIS OCCUR DURING ONE PATIENT PROCEDURE? YES. NAME OF PROCEDURE? MOHS SURGICAL EXCISION OF A NON-MALIGNANT SKIN CANCER. WILL ACTUAL DEVICE BE RETURNED FOR EVALUATION? YES, WE HAVE 4 EXAMPLES OF WHERE THE THREAD PULLED OFF THE NEEDLE DURING A CLOSURE OF A SURGICAL EXCISION. PLEASE, PROVIDE US WITH THE PROPER ADDRESS AND METHOD OF RETURNING TO YOU THE DEFECTIVE MATERIAL. WERE THERE ADVERSE PATIENT CONSEQUENCES? FORTUNATELY, THERE WERE NO PATIENT CONSEQUENCES. OUR SURGEON WAS ABLE TO RETRIEVE THE NEEDLE FROM THE PATIENT AFTER THE SUTURE THREAD SEPARATED FROM THE NEEDLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8587472&pc=GAM | 2210968-2019-81128 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-05-07 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT DURING A C-SECTION PROCEDURE ON AN UNKNOWN DATE IN (B)(6) 2019 AND SUTURE WAS USED. THE SURGEON WAS SUTURING AN INCISION AND QUESTIONED WHETHER OR NOT A PIECE OF THE NEEDLE BROKE OFF WHEN USING ON THE PATIENT. A TOTAL OF THREE X-RAYS WERE PERFORMED ALONG WITH A METAL DETECTING DEVICE AND ALL CAME BACK NEGATIVE. THERE WAS NO HARM TO THE PATIENT. THE PROCEDURE ENDED A FEW MINUTES LONGER. Manufacturer Narrative: (B)(4). THIS MEDWATCH REPORT IS IN RESPONSE TO RECEIPT OF A VOLUNTARY MEDWATCH REPORT (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8587815&pc=GAM | 2210968-2019-81130 | 2019-03-21 | Malfunction | ETHICON INC. | 2019-05-07 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN (B)(6) 2019 AND SUTURE WAS USED. WHEN PROVIDER WAS SUTURING THE PATIENT¿S INCISION AND RELEASED NEEDLE FROM STRAND THE TIP OF THE NEEDLE REMAINED IN THE NEEDLE HOLDER AND THE REST OF THE NEEDLE FELL TO THE FLOOR. FLAT PLATE WAS TAKEN TO SHOW NOTHING WAS REMAINING IN THE SURGICAL INCISION. THE TWO PIECES OF THE BROKEN NEEDLE WERE RECOVERED AND SENT WITH AN INTACT NEEDLE TO BE SENT TO FURTHER EVALUATE. PER CLINICAL LEADER, POSSIBLE USER ERROR. NO HARM TO PATIENT¿. NO FURTHER INFORMATION MAY BE OBTAINED. NO HARM TO PATIENT. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS MEDWATCH REPORT IS IN RESPONSE TO RECEIPT OF A VOLUNTARY MEDWATCH REPORT #MW5085446. AS NO CONTACT INFORMATION IS AVAILABLE FOR THE INITIAL REPORTER, NO FURTHER INFORMATION MAY BE OBTAINED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8584232&pc=GAM | 2210968-2019-81079 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-05-06 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF AND THE PIECE WAS RETRIEVED FROM THE PATIENT. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MA7086, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8584349&pc=GAM | 2210968-2019-81080 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-05-06 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CRYPTORCHIDISM PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MEM215 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8584385&pc=GAM | 2210968-2019-81081 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-05-06 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORTHOPAEDICS PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF. THE FALLEN PIECE WAS RETRIEVED FROM PATIENT. A LIKE DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ME7093, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8584416&pc=GAM | 2210968-2019-81082 | 2019-03-22 | Malfunction | ETHICON INC. | 2019-05-06 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT HERNIA PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF DURING SEWING SKIN. THE FALLEN PIECE WAS RETRIEVED FROM PATIENT. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MG8040, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8584507&pc=GAM | 2210968-2019-81083 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-06 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORTHOPAEDIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE PULLED OFF DURING SEWING SKIN IN SURGERY. THE FALLEN PIECE WAS RETRIEVED FROM PATIENT. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MGZ050 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8584510&pc=GAM | 2210968-2019-81085 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-06 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORTHOPAEDIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE PULLED OFF DURING SEWING SKIN IN SURGERY. THE FALLEN PIECE WAS RETRIEVED FROM PATIENT. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MGZ050 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8584511&pc=GAM | 2210968-2019-81084 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-06 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORTHOPAEDIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE PULLED OFF DURING SEWING SKIN IN SURGERY. THE FALLEN PIECE WAS RETRIEVED FROM PATIENT. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MGZ050 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8584717&pc=GAM | 2210968-2019-81089 | 2019-04-10 | Malfunction | ETHICON INC. | 2019-05-06 | GAM | VICRYL SUTURE 27"(70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN RECTAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING USE ON THE ANAL BORDER. THE BROKEN PIECE WAS REMOVED DURING THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL3031 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8584731&pc=GAM | 2210968-2019-81090 | 2019-04-10 | Malfunction | ETHICON INC. | 2019-05-06 | GAM | VICRYL SUTURE 27"(70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN RECTAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING USE ON THE ANAL BORDER. THE BROKEN PIECE WAS REMOVED DURING THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL3031 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8580937&pc=GAM | 2210968-2019-81043 | 2019-03-15 | Malfunction | ETHICON INC. | 2019-05-05 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN APPENDICITIS PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE BROKE DURING THE SURGERY. THE BROKEN PIECE WAS RETRIEVED FROM PATIENT. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE WAS NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8578161&pc=GAM | 2210968-2019-81005 |  | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL SUTURE 27"(70CM) 5-0 UND | Component Misassembled | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SURGEON FELT THAT THE LENGTH OF THE SUTURE HAD BEEN SHORTER THAN USUAL DURING USE. THE NEEDLE-SUTURE IN QUESTION WAS USED FOR THE PATIENT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8578490&pc=GAM | 2210968-2019-81008 |  | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS DETACHED FROM THE NEEDLE DURING INTERRUPTED SUTURING ON THE FASCIA. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MM2286, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: IT IS STATED IN EVENT "THIS EVENT OCCURRED IN SOME NEEDLE-SUTURES IN A ROW" CONFIRM THE TOTAL QTY INVOLVED WITH THIS REPORTED ISSUE DURING USE ON SINGLE PATIENT/PROCEDURE? NO FURTHER INFORMATION IS AVAILABLE. WERE ALL THE SUTURE-NEEDLE DEVICES WITH REPORTED ISSUE FROM THE SAME LOT? YES. NO FURTHER INFORMATION WILL BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8578535&pc=GAM | 2210968-2019-81009 |  | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC PROCEDURE ON UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS DETACHED FROM THE NEEDLE DURING INTERRUPTED SUTURING ON THE FASCIA. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MM2286, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: IT IS STATED IN EVENT "THIS EVENT OCCURRED IN SOME NEEDLE-SUTURES IN A ROW" CONFIRM THE TOTAL QTY INVOLVED WITH THIS REPORTED ISSUE DURING USE ON SINGLE PATIENT/PROCEDURE? NO FURTHER INFORMATION IS AVAILABLE. WERE ALL THE SUTURE-NEEDLE DEVICES WITH REPORTED ISSUE FROM THE SAME LOT? YES. NO FURTHER INFORMATION WILL BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8578864&pc=GAM | 2210968-2019-81019 | 2019-04-04 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A FACE LESION EXCISION ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE WAS POPPING OFF THE SUTURE, THIS OCCURRED ON THE FIRST PASS. THE CASE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE, DIFFERENT LOT NUMBER. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY TIMES DID THE NEEDLE POP OFF THE SUTURE? UNKNOWN BUT MORE THAN ONCE. DID THE EVENT OCCUR DURING ONE OR MULTIPLE PATIENT PROCEDURES? ONE PROCEDURE. IF THIS EVENT OCCURRED IN MULTIPLE PROCEDURES, PLEASE PROVIDE THE FOLLOWING INFORMATION: N/A. WHAT IS THE TOTAL NUMBER OF PROCEDURES? N/A. HAVE ANY OF THESE EVENTS BEEN PREVIOUSLY REPORTED TO ETHICON? IF SO, PROVIDE THE RESPECTIVE REFERENCE NUMBER(S). N/A. WHAT ARE THE PROCEDURE NAME(S) AND DATE(S)? N/A. WAS THERE ANY ADVERSE PATIENT CONSEQUENCE(S) OR SUBSEQUENT MEDICAL/SURGICAL INTERVENTION? NO. IS THE DEVICE AVAILABLE TO BE RETURNED FOR EVALUATION? IF SO, PLEASE PROVIDE THE STATUS OF THE RETURN SAMPLE AND ANY AVAILABLE TRACKING INFORMATION. PRODUCT USED IS NOT AVAILABLE. REMAINING BOX TO BE SENT IN FOR EVALUATION. NOTE: EVENTS REPORTED VIA 2210968-2019-81019 AND 2210968-2019-81020.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8578870&pc=GAM | 2210968-2019-81020 | 2019-04-04 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A FACE LESION EXCISION ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE WAS POPPING OFF THE SUTURE, THIS OCCURRED ON THE FIRST PASS. THE CASE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE, DIFFERENT LOT NUMBER. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY TIMES DID THE NEEDLE POP OFF THE SUTURE? UNKNOWN BUT MORE THAN ONCE. DID THE EVENT OCCUR DURING ONE OR MULTIPLE PATIENT PROCEDURES? ONE PROCEDURE. IF THIS EVENT OCCURRED IN MULTIPLE PROCEDURES, PLEASE PROVIDE THE FOLLOWING INFORMATION: N/A. WHAT IS THE TOTAL NUMBER OF PROCEDURES? N/A. HAVE ANY OF THESE EVENTS BEEN PREVIOUSLY REPORTED TO ETHICON? IF SO, PROVIDE THE RESPECTIVE REFERENCE NUMBER(S). N/A. WHAT ARE THE PROCEDURE NAME(S) AND DATE(S)? N/A. WAS THERE ANY ADVERSE PATIENT CONSEQUENCE(S) OR SUBSEQUENT MEDICAL/SURGICAL INTERVENTION? NO. IS THE DEVICE AVAILABLE TO BE RETURNED FOR EVALUATION? IF SO, PLEASE PROVIDE THE STATUS OF THE RETURN SAMPLE AND ANY AVAILABLE TRACKING INFORMATION. PRODUCT USED IS NOT AVAILABLE. REMAINING BOX TO BE SENT IN FOR EVALUATION. NOTE: EVENTS REPORTED VIA 2210968-2019-81019 AND 2210968-2019-81020.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579224&pc=GAM | 2210968-2019-81027 | 2019-04-08 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT WOUND CLOSURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE TIP BROKE OFF IN THE NEEDLE HOLDER. THE NEEDLE TIP AND NEEDLE WERE LOCATED. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MP2304, AND NO NONCONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579288&pc=GAM | 2210968-2019-81029 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL SUTURE 27"(70CM) 0 VIO | Contamination /Decontamination Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE OUTER PACKAGE HAS ONE HAIR. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579491&pc=GAM | 2210968-2019-81031 | 2019-04-08 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL SUTURE 54"(135CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT TOTAL KNEE REPLACEMENT ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PREMATURELY POPPED OFF WHILE SUTURING. THIS OCCURRED AFTER SEVERAL PASSES. THE CASE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579492&pc=GAM | 2210968-2019-81030 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL SUTURE 54"(135CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE PREMATURELY POPPED OFF WHILE SUTURING. THIS OCCURRED AFTER SEVERAL PASSES. THE CASE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579494&pc=GAM | 2210968-2019-81032 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL SUTURE 54"(135CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE PREMATURELY POPPED OFF WHILE SUTURING. THIS OCCURRED AFTER SEVERAL PASSES. THE CASE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579495&pc=GAM | 2210968-2019-81034 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL SUTURE 54"(135CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE PREMATURELY POPPED OFF WHILE SUTURING. THIS OCCURRED AFTER SEVERAL PASSES. THE CASE WAS COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579705&pc=GAM | 2210968-2019-81035 | 2019-03-31 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS REPORTED THAT THERE WAS DELAYED SUTURE ABSORPTION. THE INCISION HEALED NOT WELL IN POSTOPERATIVE 6 DAYS. MEDICATION WAS GIVEN ON INCISION TO ADDRESS IT. THEN PATIENT CONDITION CHANGED TO BETTER. THERE IS NO REPORT ON PATIENT'S INJURY. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT KIND OF MEDICATION / TREATMENT WAS PROVIDED TO PATIENT ON THE INCISION, POST-OP? UNK WAS ANY PRESCRIBED MEDICATION / ANTIBIOTICS GIVEN TO PATIENT POST-OP? UNK. WAS ANY RE-OPERATION/RE-SUTURING/ REVISION SURGERY DONE ON PATIENT POST-OP INITIAL PROCEDURE? UNK.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579911&pc=GAM | 2210968-2019-81038 | 2019-03-23 | Malfunction | ETHICON INC. | 2019-05-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF DURING SEWING THE SKIN. THE FALLEN PIECE WAS RETRIEVED FROM THE PATIENT. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8573565&pc=GAM | 2210968-2019-80857 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-02 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS COMING OFF OF THE NEEDLE. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 BATCH NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. A LABELED WINDING FORMER, A NEEDLE AND A SUTURE PIECE OF PRODUCT CODE J497, LOT MJK059 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE AS EXPECTED AND REMNANT OF THE SUTURE WAS NOTED INTO THE BARREL HOLE AND SLIGHTLY MARKS WERE OBSERVED ON THE EDGE OF THE NEEDLE THAT APPEARS TO BE BY THE USE OF A SURGICAL INSTRUMENT. THE DISPENSED SUTURE PIECE WAS EXAMINED ALONG OF THE STRAND AND BODY FLUIDS WERE NOTED AND THE END OF THE SUTURE WAS CUT APPEARS TO BE BY THE USE OF A SURGICAL INSTRUMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE SAMPLE CONDITION THE ASSIGNABLE OF THE PULL OFF - SUTURE NEEDLE SUGGEST AN IMPROPER HANDLING OF THE SAMPLE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS ALSO REPORTED THAT THIS ISSUE HAS HAPPENED PREVIOUSLY. PLEASE ADVISE: WERE THE PREVIOUS ISSUES REPORTED TO ETHICON? IF YES, DO YOU HAVE A PRODUCT COMPLAINT NUMBER? IF NOT PREVIOUSLY REPORTED TO ETHICON, PLEASE PROVIDE THE FOLLOWING INFORMATION FOR EACH PROCEDURE: PROCEDURE NAME AND DATE, PRODUCT CODE, LOT NUMBER, PLEASE DESCRIBE THE ISSUE THAT OCCURRED WITH THE DEVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8573566&pc=GAM | 2210968-2019-80856 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-02 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS COMING OFF OF THE NEEDLE. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. A LABELED WINDING FORMER, A NEEDLE AND A SUTURE PIECE OF PRODUCT CODE J497, LOT MJK059 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE AS EXPECTED. THE BARREL HOLE WAS EXAMINED UNDER 20X MAGNIFICATION AND REMNANT SUTURE WAS NOTED. THE SUTURE END IS SEVERELY CUT, RESULTING IN A CLIP OFF DEFECT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF THE NEEDLE PULL OFF, IS A CLIP OFF DEFECT, THIS DEFECT IS CAUSED BY EXCESSIVE PRESSURE ON THE SUTURE DURING SWAGING OF THE NEEDLE AND CONSEQUENTLY BREAKING. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS ALSO REPORTED THAT THIS ISSUE HAS HAPPENED PREVIOUSLY. PLEASE ADVISE: WERE THE PREVIOUS ISSUES REPORTED TO ETHICON? IF YES, DO YOU HAVE A PRODUCT COMPLAINT NUMBER? IF NOT PREVIOUSLY REPORTED TO ETHICON, PLEASE PROVIDE THE FOLLOWING INFORMATION FOR EACH PROCEDURE: PROCEDURE NAME AND DATE, PRODUCT CODE, LOT NUMBER, PLEASE DESCRIBE THE ISSUE THAT OCCURRED WITH THE DEVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8573762&pc=GAM | 2210968-2019-80903 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-02 | GAM | VICRYL SUTURE 18"(45CM) 4-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT SKIN EXCISION ON AN UNKNOWN DATE AND SUTURE WAS USED. THE DOCTOR WAS PULLING THE SUTURE THROUGH THE FATTY LAYER OF TISSUE IN AN INTERRUPTED LOOP AND THE NEEDLE SEPARATED FROM THE SUTURE. THE DOCTOR COMPLETED THE PROCEDURE USING A DIFFERENT SIZE OF SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MKZ959 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. IT WAS RECEIVED FOR ANALYSIS AN OPENED BOX WITH EIGHT UNOPENED SAMPLES OF PRODUCT CODE J494, LOT MKZ959. DURING THE VISUAL INSPECTION OF EIGHT SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLES, NO PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8573774&pc=GAM | 2210968-2019-80904 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-02 | GAM | VICRYL SUTURE 18"(45CM) 4-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT SKIN EXCISION ON AN UNKNOWN DATE AND SUTURE WAS USED. THE DOCTOR WAS PULLING THE SUTURE THROUGH THE FATTY LAYER OF TISSUE IN AN INTERRUPTED LOOP AND THE NEEDLE SEPARATED FROM THE SUTURE. THE DOCTOR COMPLETED THE PROCEDURE USING A DIFFERENT SIZE OF SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MKZ959 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. IT WAS RECEIVED FOR ANALYSIS AN OPENED BOX WITH EIGHT UNOPENED SAMPLES OF PRODUCT CODE J494, LOT MKZ959. DURING THE VISUAL INSPECTION OF EIGHT SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLES, NO PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8573777&pc=GAM | 2210968-2019-80905 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-02 | GAM | VICRYL SUTURE 18"(45CM) 4-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT SKIN EXCISION ON AN UNKNOWN DATE AND SUTURE WAS USED. THE DOCTOR WAS PULLING THE SUTURE THROUGH THE FATTY LAYER OF TISSUE IN AN INTERRUPTED LOOP AND THE NEEDLE SEPARATED FROM THE SUTURE. THE DOCTOR COMPLETED THE PROCEDURE USING A DIFFERENT SIZE OF SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MKZ959 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. IT WAS RECEIVED FOR ANALYSIS AN OPENED BOX WITH EIGHT UNOPENED SAMPLES OF PRODUCT CODE J494, LOT MKZ959. DURING THE VISUAL INSPECTION OF EIGHT SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLES, NO PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8575036&pc=GAM | 2210968-2019-80969 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-02 | GAM | VICRYL PLUS SUTURE 27"(70CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SURGEON REPORTED THAT THE SUTURE DID NOT FEEL LIKE A NORMAL SUTURE AND IT KEPT BREAKING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8568663&pc=GAM | 2210968-2019-80858 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION FOR NON-FAVORABLE CERVIX ON (B)(6) 2019 AND SUTURE WAS USED. AT THE END OF THE SUTURE OF THE UTERUS, THE FILAMENTS OF THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL3631 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THIS ISSUE OCCUR WITH 2 PATIENTS OR 3 PATIENTS? DURING EACH PATIENT PROCEDURE, DID 1 SUTURE NEEDLE PULL OFF AND DID 1 SUTURE BREAK WHILE SUTURING THE UTERUS? WILL THE DEVICE INVOLVED IN THE EVENT BE RETURNED FOR EVALUATION? NOTE: BREAKAGE OF SUTURE REPORTED VIA 2210968-2019-80858 AND PULL OFF SUTURE NEEDLE REPORTED VIA 2210968-2019-80859.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8568664&pc=GAM | 2210968-2019-80859 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION FOR NON-FAVORABLE CERVIX ON (B)(6) 2019 AND SUTURE WAS USED. WHEN SUTURING THE FASCIA, THE SUTURE NEEDLE WAS DISASSEMBLED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL3631 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THIS ISSUE OCCUR WITH 2 PATIENTS OR 3 PATIENTS? DURING EACH PATIENT PROCEDURE, DID 1 SUTURE NEEDLE PULL OFF AND DID 1 SUTURE BREAK WHILE SUTURING THE UTERUS? WILL THE DEVICE INVOLVED IN THE EVENT BE RETURNED FOR EVALUATION? NOTE: BREAKAGE OF SUTURE REPORTED VIA 2210968-2019-80858 AND PULL OFF SUTURE NEEDLE REPORTED VIA 2210968-2019-80859.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8568881&pc=GAM | 2210968-2019-80863 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION FOR NON-FAVORABLE CERVIX ON (B)(6) 2019 AND SUTURE WAS USED. AT THE END OF THE SUTURE OF THE UTERUS, THE FILAMENTS OF THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL3631 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THIS ISSUE OCCUR WITH 2 PATIENTS OR 3 PATIENTS? DURING EACH PATIENT PROCEDURE, DID 1 SUTURE NEEDLE PULL OFF AND DID 1 SUTURE BREAK WHILE SUTURING THE UTERUS? WILL THE DEVICE INVOLVED IN THE EVENT BE RETURNED FOR EVALUATION? NOTE: BREAKAGE OF SUTURE REPORTED VIA 2210968-2019-80863 AND PULL OFF SUTURE NEEDLE 2210968-2019-80865.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8568882&pc=GAM | 2210968-2019-80865 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION FOR NON-FAVORABLE CERVIX ON (B)(6) 2019 AND SUTURE WAS USED. WHEN SUTURING THE FASCIA, THE SUTURE NEEDLE WAS DISASSEMBLED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL3631 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THIS ISSUE OCCUR WITH 2 PATIENTS OR 3 PATIENTS? DURING EACH PATIENT PROCEDURE, DID 1 SUTURE NEEDLE PULL OFF AND DID 1 SUTURE BREAK WHILE SUTURING THE UTERUS? WILL THE DEVICE INVOLVED IN THE EVENT BE RETURNED FOR EVALUATION? NOTE: BREAKAGE OF SUTURE REPORTED VIA 2210968-2019-80863 AND PULL OFF SUTURE NEEDLE 2210968-2019-80865.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8568902&pc=GAM | 2210968-2019-80861 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL VIO 90CM M2 USP3/0 SGLE ARMED CP-1 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT A WOUND RECOVERY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8568903&pc=GAM | 2210968-2019-80860 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL VIO 90CM M2 USP3/0 SGLE ARMED CP-1 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OP, THE SUTURE BROKE IN ITS LENGTH AND THE KNOTS WERE INTACT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8569007&pc=GAM | 2210968-2019-80869 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION FOR NON-FAVORABLE CERVIX ON (B)(6) 2019 AND SUTURE WAS USED. WHEN SUTURING THE FASCIA, THE SUTURE NEEDLE WAS DISASSEMBLED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL3631 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THIS ISSUE OCCUR WITH 2 PATIENTS OR 3 PATIENTS? DURING EACH PATIENT PROCEDURE, DID 1 SUTURE NEEDLE PULL OFF AND DID 1 SUTURE BREAK WHILE SUTURING THE UTERUS? WILL THE DEVICE INVOLVED IN THE EVENT BE RETURNED FOR EVALUATION? NOTE: BREAKAGE OF SUTURE REPORTED VIA 2210968-2019-80868 AND PULL OFF SUTURE NEEDLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8569008&pc=GAM | 2210968-2019-80868 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION FOR NON-FAVORABLE CERVIX ON (B)(6) 2019 AND SUTURE WAS USED. AT THE END OF THE SUTURE OF THE UTERUS, THE FILAMENTS OF THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL3631 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THIS ISSUE OCCUR WITH 2 PATIENTS OR 3 PATIENTS? DURING EACH PATIENT PROCEDURE, DID 1 SUTURE NEEDLE PULL OFF AND DID 1 SUTURE BREAK WHILE SUTURING THE UTERUS? WILL THE DEVICE INVOLVED IN THE EVENT BE RETURNED FOR EVALUATION? NOTE: BREAKAGE OF SUTURE REPORTED VIA 2210968-2019-80868 AND PULL OFF SUTURE NEEDLE 2210968-2019-80869.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8569330&pc=GAM | 2210968-2019-80889 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8569676&pc=GAM | 2210968-2019-80891 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE UNKNOWN PROD | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY ON AN UNKNOWN DATE IN 2019 AND ANTIBACTERIAL SUTURE WAS USED. THE SUTURE APPEARS MORE FRAGILE THAN TRADITIONAL SUTURES. THE SUTURE HAS BROKEN/ SNAPPED. NEW SUTURE WAS OPENED, 3 MINUTES DELAY IN THE SURGERY, SAMPLE DISCARDED. NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PRODUCT CODE? DO NOT HAVE INFORMATION. HAVE REQUESTED FOR CUSTOMER TO RECORD GOING FORWARD. LOT #? DO NOT HAVE INFORMATION. HAVE REQUESTED FOR CUSTOMER TO RECORD GOING FORWARD. IT IS STATED EVENT REPORT NOTICED ON NUMEROUS OCCASIONS, IF THAT IS CORRECT. PLEASE CONFIRM THE SPECIFIC NUMBER OF PATIENT EVENTS / PROCEDURES? DO NOT HAVE INFORMATION. HAVE REQUESTED FOR CUSTOMER TO RECORD GOING FORWARD. HAVE ANY OF THESE EVENTS/PROCEDURES BEEN PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE COMPLAINT FILE NUMBER/ AFFILIATE NUMBER? NO. FOR EACH PATIENT EVENT PLEASE PROVIDE ADDITIONAL COMPLAINT DETAILS: EVENT DATE, PRODUCT CODE AND LOT NUMBER INVOLVED, PROCEDURE, EVENT DESCRIPTION, QTY INVOLVED? WAS THERE ANY ADVERSE PATIENT OUTCOME? IF YES, WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED, WAS PROCEDURE COMPLETED SUCCESSFULLY AND HOW? ANY SAMPLE RETURN..ETC? DO NOT HAVE INFORMATION. HAVE REQUESTED FOR CUSTOMER TO RECORD GOING FORWARD. "TRADITIONAL SUTURES" CAN YOU PROVIDE EXACTLY WHICH SUTURE THE SURGEON IS REFERRING TO- PRODUCT CODE? MATERIAL ¿PLUS MONOCRYL, VICRYL AND PDS. HOW WAS THESE SUTURES STORES IN THE HOSPITAL? THE SAME WAY THE "TRADITIONAL SUTURES" WERE STORED? STORED IN THE SAME WAY AS ALL OTHER SUTURES. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8570198&pc=GAM | 2210968-2019-80906 | 2019-04-04 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL SUTURE 12"(30CM) 6-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EYE PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE JUST BROKE IN HALF AS IF SOMEONE CUT IT. CASE COMPLETED WITH ANOTHER DEVICE OF THE SAME PRODUCT CODE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8570333&pc=GAM | 2210968-2019-80907 | 2019-04-04 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL SUTURE 12"(30CM) 6-0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EYE PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN TYING THE SUTURE ON THE EYE MUSCLE, THE SUTURE POPPED OFF THE NEEDLE. THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8570378&pc=GAM | 2210968-2019-80914 | 2019-04-04 | Malfunction | ETHICON INC. | 2019-05-01 | GAM | VICRYL SUTURE 12"(30CM) 6-0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EYE PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN TYING THE SUTURE ON THE EYE MUSCLE, THE SUTURE POPPED OFF THE NEEDLE. THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8565670&pc=GAM | 2210968-2019-80840 | 2019-03-31 | Malfunction | ETHICON INC. | 2019-04-30 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CONIZATION OF CERVIX ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, WHILE SEWING CERVIX, THE NEEDLE BROKE. THE BROKEN PIECE WAS RETRIEVED FROM PATIENT. CHANGED TO ANOTHER ONE TO COMPLETE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8565715&pc=GAM | 2210968-2019-80842 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-04-30 | GAM | VICRYL PLUS SUTURE 18"(45CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A COLORECTAL SURGERY ON (B)(6) 2019 AND A BARBED SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THE BROKEN PIECE WAS RETRIEVED FROM PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8566789&pc=GAM | 2210968-2019-80803 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-04-30 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: AN EMPTY OPENED FOIL, A DETACHED NEEDLE AND A DISPENSED USED SUTURE OF PRODUCT CODE VCP433, LOT # MGK372 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF A DETACHED NEEDLE, THE SWAGE AND ATTACHMENT PRESENT MARKS AND THE BARREL HOLE WAS DAMAGED (FLAT) THAT APPEARS TO BE BY USE OF THE SURGICAL INSTRUMENTS COULD BE OBSERVED. THE SUTURE WAS OBSERVED DETACHED DO THE STRESS APPLIED TO THE STRAND, PRESENT DAMAGED ON THE SURFACE AND STRESS AT THE END. PER THE NEEDLES CONDITION THE ASSIGNABLE CAUSE OF THE PERFORMANCE - PULL OFF SUTURE NEEDLE, WAS AN IMPROPER HANDLING. TO AVOID THIS KIND OF DAMAGE THE PACKAGES, GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. IN HANDLING THIS OR ANY OTHER SUTURE MATERIAL, CARE SHOULD BE TAKEN TO AVOID DAMAGE FROM HANDLING. AVOID CRUSHING OR CRIMPING DAMAGE DUE TO APPLICATION OF SURGICAL INSTRUMENTS SUCH AS FORCEPS OR NEEDLE HOLDERS. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MGK372 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 04/30/3019. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8566814&pc=GAM | 2210968-2019-80848 | 2019-03-05 | Malfunction | ETHICON INC. | 2019-04-30 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE DURING SEWING UTERUS. CHANGED ANOTHER ONE TO COMPLETE SURGERY. NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AND NO NONCONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8561393&pc=GAM | 2210968-2019-80789 | 2019-04-02 | Malfunction | ETHICON INC. | 2019-04-29 | GAM | VICRYL PLUS SUTURE 18"(45CM) 5-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD CUT AT THE LEVEL OF THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENT STATES ''4 THREADS CUT AT THE LEVEL OF THE NEEDLE''. DID THE SUTURE BREAK OR THE NEEDLE PULLED OFF THE SUTURE THREAD? DID THE ISSUE OCCUR DURING ACTUAL USE ON THE PATIENT? QUANTITY INVOLVED REPORTED AS 1; HOW MANY DEVICES HAD ISSUE IN THIS PROCEDURE?1 OR 4? HOW WAS CASE COMPLETED? IS THE LOT NUMBER PBBBPWS0 CORRECT? ANY PATIENT CONSEQUENCES AND WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM? IS THE PRODUCT AVAILABLE FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8561398&pc=GAM | 2210968-2019-80787 | 2019-04-02 | Malfunction | ETHICON INC. | 2019-04-29 | GAM | VICRYL PLUS SUTURE 18"(45CM) 5-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD CUT AT THE LEVEL OF THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). IT WAS REPORTED THAT THIS DEVICE IS NOT MALFUNCTION REPORTABLE. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. THE FOLLOWING ADDITIONAL INFORMATION WAS OBTAINED: THE NEEDLE PULLED OFF FROM THE THREAD DURING THE USE ON THE PATIENT. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE. NO PATIENT CONSEQUENCE. 2 DEVICES HAS BEEN USED PER TWO PROCEDURES. DEVICES NOT AVAILABLE FOR ANALYSIS. NOTE: EVENTS FOR FIRST PROCEDURE REPORTED VIA 2210968-2019-80789, 2210968-2019-80788. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENT STATES '4 THREADS CUT AT THE LEVEL OF THE NEEDLE'; DID THE SUTURE BREAK OR THE NEEDLE PULLED OFF THE SUTURE THREAD? DID THE ISSUE OCCUR DURING ACTUAL USE ON THE PATIENT? QUANTITY INVOLVED REPORTED AS 1; HOW MANY DEVICES HAD ISSUE IN THIS PROCEDURE?1 OR 4? HOW WAS CASE COMPLETED? IS THE LOT NUMBER PBBBPWS0 CORRECT? ANY PATIENT CONSEQUENCES AND WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM? IS THE PRODUCT AVAILABLE FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8561401&pc=GAM | 2210968-2019-80788 | 2019-04-02 | Malfunction | ETHICON INC. | 2019-04-29 | GAM | VICRYL PLUS SUTURE 18"(45CM) 5-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD CUT AT THE LEVEL OF THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENT STATES '4 THREADS CUT AT THE LEVEL OF THE NEEDLE'; DID THE SUTURE BREAK OR THE NEEDLE PULLED OFF THE SUTURE THREAD? DID THE ISSUE OCCUR DURING ACTUAL USE ON THE PATIENT? QUANTITY INVOLVED REPORTED AS 1; HOW MANY DEVICES HAD ISSUE IN THIS PROCEDURE?1 OR 4? HOW WAS CASE COMPLETED? IS THE LOT NUMBER PBBBPWS0 CORRECT? ANY PATIENT CONSEQUENCES AND WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM? IS THE PRODUCT AVAILABLE FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8561423&pc=GAM | 2210968-2019-80786 | 2019-04-02 | Malfunction | ETHICON INC. | 2019-04-29 | GAM | VICRYL PLUS SUTURE 18"(45CM) 5-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD CUT AT THE LEVEL OF THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENT STATES '4 THREADS CUT AT THE LEVEL OF THE NEEDLE'; DID THE SUTURE BREAK OR THE NEEDLE PULLED OFF THE SUTURE THREAD? DID THE ISSUE OCCUR DURING ACTUAL USE ON THE PATIENT? QUANTITY INVOLVED REPORTED AS 1; HOW MANY DEVICES HAD ISSUE IN THIS PROCEDURE?1 OR 4? HOW WAS CASE COMPLETED? IS THE LOT NUMBER PBBBPWS0 CORRECT? ANY PATIENT CONSEQUENCES AND WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM? IS THE PRODUCT AVAILABLE FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8561704&pc=GAM | 2210968-2019-80791 |  | Malfunction | ETHICON INC. | 2019-04-29 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 2-0 UND | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS DETACHED FROM THE SUTURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8561943&pc=GAM | 2210968-2019-80798 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-04-29 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS BROKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJ2877, AND NO NONCONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8561971&pc=GAM | 2210968-2019-80799 | 2019-03-28 | Malfunction | ETHICON INC. | 2019-04-29 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS DETACHED FROM THE SUTURE WHEN PASSING THE NEEDLE THROUGH THE TISSUE. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8555940&pc=GAM | 2210968-2019-80746 | 2019-03-28 | Malfunction | ETHICON INC. | 2019-04-26 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC OVARIAN CYST PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING USE. CHANGED ANOTHER ONE TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ML2674, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8548863&pc=GAM | 2210968-2019-80699 | 2019-03-04 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT BREAST AUGMENTATION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OUT FROM THE SUTURE HEAD AND BROKE OFF AT THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MK2325, AND NO NONCONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. NEEDLE PULLING OUT FROM SUTURE HEAD. DID 6 DEVICES HAD ISSUE IN ONE PROCEDURE? IF YES, DID ALL 6 HAD ISSUE DURING ACTUAL SUTURING? IF NO, FOR EACH DEVICE STATE WHEN DID THE ISSUE OCCUR? ALL ISSUES WERE DURING SUTURING IN ONE BREAST LIFT/AUGMENT PROCEDURE. -IF NO, PLEASE PROVIDE SPECIFIC NUMBER OF PROCEDURES INVOLVED. FOR EACH PROCEDURE A SEPARATE PC SHOULD BE OPENED WITH FOLLOWING INFORMATION- EVENT DATE? (B)(6). PROCEDURE NAME? - MASTOPEXI/AUGMENTATION QUANTITY INVOLVED? - SIX LOT NUMBER IF DIFFERENT? SAME. EVENT DESCRIPTION STATING WHEN DURING THE PROCEDURE DID EACH DEVICE HAD THE ISSUE. ANY ADVERSE PATIENT CONSEQUENCES? IF YES, EXPLAIN AND STATE WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM. NO ADVERSE EVENT TO PATIENT. CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. SUTURES BROKE AT NEEDLE. NOTE: MEDWATCH FOR REPORTED EVENT SUBMITTED VIA 2210968-2019-80698, 2210968-2019-80700, 2210968-2019-80701, 2210968-2019-80702, 2210968-2019-80703.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8548864&pc=GAM | 2210968-2019-80700 | 2019-03-04 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT BREAST AUGMENTATION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OUT FROM THE SUTURE HEAD AND BROKE OFF AT THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MK2325, AND NO NONCONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. -NEEDLE PULLING OUT FROM SUTURE HEAD. DID 6 DEVICES HAD ISSUE IN ONE PROCEDURE? IF YES, DID ALL 6 HAD ISSUE DURING ACTUAL SUTURING? IF NO, FOR EACH DEVICE STATE WHEN DID THE ISSUE OCCUR? ALL ISSUES WERE DURING SUTURING IN ONE BREAST LIFT/AUGMENT PROCEDURE. IF NO, PLEASE PROVIDE SPECIFIC NUMBER OF PROCEDURES INVOLVED. FOR EACH PROCEDURE A SEPARATE PC SHOULD BE OPENED WITH FOLLOWING INFORMATION- EVENT DATE? (B)(6). PROCEDURE NAME? MASTOPEXI/AUGMENTATION. QUANTITY INVOLVED? SIX. LOT NUMBER IF DIFFERENT? SAME. EVENT DESCRIPTION STATING WHEN DURING THE PROCEDURE DID EACH DEVICE HAD THE ISSUE. ANY ADVERSE PATIENT CONSEQUENCES? IF YES, EXPLAIN AND STATE WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM. NO ADVERSE EVENT TO PATIENT. CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. SUTURES BROKE AT NEEDLE. NOTE: MEDWATCH FOR REPORTED EVENT SUBMITTED VIA 2210968-2019-80698, 2210968-2019-80699, 2210968-2019-80701, 2210968-2019-80702, 2210968-2019-80703.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8548865&pc=GAM | 2210968-2019-80698 | 2019-03-04 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT BREAST AUGMENTATION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OUT FROM THE SUTURE HEAD AND BROKE OFF AT THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MK2325, AND NO NONCONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. NEEDLE PULLING OUT FROM SUTURE HEAD. DID 6 DEVICES HAD ISSUE IN ONE PROCEDURE? IF YES, DID ALL 6 HAD ISSUE DURING ACTUAL SUTURING? IF NO, FOR EACH DEVICE STATE WHEN DID THE ISSUE OCCUR? ALL ISSUES WERE DURING SUTURING IN ONE BREAST LIFT/AUGMENT PROCEDURE. IF NO, PLEASE PROVIDE SPECIFIC NUMBER OF PROCEDURES INVOLVED. FOR EACH PROCEDURE A SEPARATE PC SHOULD BE OPENED WITH FOLLOWING INFORMATION- EVENT DATE? (B)(6). PROCEDURE NAME? MASTOPEXI/AUGMENTATION QUANTITY INVOLVED? SIX. LOT NUMBER IF DIFFERENT? SAME. EVENT DESCRIPTION STATING WHEN DURING THE PROCEDURE DID EACH DEVICE HAD THE ISSUE. ANY ADVERSE PATIENT CONSEQUENCES? IF YES, EXPLAIN AND STATE WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM. NO ADVERSE EVENT TO PATIENT. CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. SUTURES BROKE AT NEEDLE. NOTE: MEDWATCH FOR REPORTED EVENT SUBMITTED VIA 2210968-2019-80699, 2210968-2019-80700, 2210968-2019-80701, 2210968-2019-80702, 2210968-2019-80703.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8548867&pc=GAM | 2210968-2019-80701 | 2019-03-04 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT BREAST AUGMENTATION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OUT FROM THE SUTURE HEAD AND BROKE OFF AT THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MK2325, AND NO NONCONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. NEEDLE PULLING OUT FROM SUTURE HEAD. DID 6 DEVICES HAD ISSUE IN ONE PROCEDURE? IF YES, DID ALL 6 HAD ISSUE DURING ACTUAL SUTURING? IF NO, FOR EACH DEVICE STATE WHEN DID THE ISSUE OCCUR? ALL ISSUES WERE DURING SUTURING IN ONE BREAST LIFT/AUGMENT PROCEDURE. IF NO, PLEASE PROVIDE SPECIFIC NUMBER OF PROCEDURES INVOLVED. FOR EACH PROCEDURE A SEPARATE PC SHOULD BE OPENED WITH FOLLOWING INFORMATION- EVENT DATE? (B)(6). PROCEDURE NAME? MASTOPEXI/AUGMENTATION. QUANTITY INVOLVED? SIX. LOT NUMBER IF DIFFERENT? SAME. EVENT DESCRIPTION STATING WHEN DURING THE PROCEDURE DID EACH DEVICE HAD THE ISSUE. ANY ADVERSE PATIENT CONSEQUENCES? IF YES, EXPLAIN AND STATE WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM. NO ADVERSE EVENT TO PATIENT. CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. SUTURES BROKE AT NEEDLE. NOTE: MEDWATCH FOR REPORTED EVENT SUBMITTED VIA 2210968-2019-80698, 2210968-2019-80699, 2210968-2019-80700, 2210968-2019-80702, 2210968-2019-80703.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8548868&pc=GAM | 2210968-2019-80702 | 2019-03-04 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT BREAST AUGMENTATION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OUT FROM THE SUTURE HEAD AND BROKE OFF AT THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MK2325, AND NO NONCONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. NEEDLE PULLING OUT FROM SUTURE HEAD. DID 6 DEVICES HAD ISSUE IN ONE PROCEDURE? IF YES, DID ALL 6 HAD ISSUE DURING ACTUAL SUTURING? IF NO, FOR EACH DEVICE STATE WHEN DID THE ISSUE OCCUR? ALL ISSUES WERE DURING SUTURING IN ONE BREAST LIFT/AUGMENT PROCEDURE. IF NO, PLEASE PROVIDE SPECIFIC NUMBER OF PROCEDURES INVOLVED. FOR EACH PROCEDURE A SEPARATE PC SHOULD BE OPENED WITH FOLLOWING INFORMATION- EVENT DATE? (B)(6). PROCEDURE NAME? MASTOPEXI/AUGMENTATION. QUANTITY INVOLVED? SIX. LOT NUMBER IF DIFFERENT? SAME. EVENT DESCRIPTION STATING WHEN DURING THE PROCEDURE DID EACH DEVICE HAD THE ISSUE. ANY ADVERSE PATIENT CONSEQUENCES? IF YES, EXPLAIN AND STATE WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM. NO ADVERSE EVENT TO PATIENT. CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. SUTURES BROKE AT NEEDLE. NOTE: MEDWATCH FOR REPORTED EVENT SUBMITTED VIA 2210968-2019-80698, 2210968-2019-80699, 2210968-2019-80700, 2210968-2019-80701, 2210968-2019-80703.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8548870&pc=GAM | 2210968-2019-80703 | 2019-03-04 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT BREAST AUGMENTATION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OUT FROM THE SUTURE HEAD AND BROKE OFF AT THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MK2325, AND NO NONCONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. NEEDLE PULLING OUT FROM SUTURE HEAD. DID 6 DEVICES HAD ISSUE IN ONE PROCEDURE? -IF YES, DID ALL 6 HAD ISSUE DURING ACTUAL SUTURING? IF NO, FOR EACH DEVICE STATE WHEN DID THE ISSUE OCCUR? ALL ISSUES WERE DURING SUTURING IN ONE BREAST LIFT/AUGMENT PROCEDURE. IF NO, PLEASE PROVIDE SPECIFIC NUMBER OF PROCEDURES INVOLVED. FOR EACH PROCEDURE A SEPARATE PC SHOULD BE OPENED WITH FOLLOWING INFORMATION- EVENT DATE? (B)(6). PROCEDURE NAME? MASTOPEXI/AUGMENTATION. QUANTITY INVOLVED? SIX. LOT NUMBER IF DIFFERENT? SAME. EVENT DESCRIPTION STATING WHEN DURING THE PROCEDURE DID EACH DEVICE HAD THE ISSUE. ANY ADVERSE PATIENT CONSEQUENCES? IF YES, EXPLAIN AND STATE WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THEM. NO ADVERSE EVENT TO PATIENT. CLARIFY THE ISSUE-DID THE SUTURES BREAK OR THE NEEDLE PULLING OUT FROM SUTURE THREAD. SUTURES BROKE AT NEEDLE. NOTE: MEDWATCH FOR REPORTED EVENT SUBMITTED VIA 2210968-2019-80698, 2210968-2019-80699, 2210968-2019-80700, 2210968-2019-80701, 2210968-2019-80702.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8550328&pc=GAM | 2210968-2019-80713 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CESAREAN SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE THREAD PULLED OFF THE NEEDLE WHILE IN USE DURING CLOSURE. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM ANOTHER LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MM2688, AND NO NONCONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8550329&pc=GAM | 2210968-2019-80712 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CESAREAN SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE THREAD PULLED OFF THE NEEDLE WHILE IN USE DURING CLOSURE. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM ANOTHER LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MM2688, AND NO NONCONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8550747&pc=GAM | 2210968-2019-80715 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT HYSTERECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGEON WAS PUTTING CONTINUOUS SUTURING AND TOWARDS THE END WHILE TRYING TO PUT KNOTS THE SUTURE BROKE UP. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ONE OPENED COMPLAINT SAMPLE FOIL ONLY AS A COMPLAINT SAMPLE FOR CODE NW2346 LOT T8010 WAS RECEIVED. SUTURES AND NEEDLE MATERIAL WERE NOT RETURNED ALONG WITH COMPLAINT SAMPLE FOIL FOR INVESTIGATION. THE FOIL PACK WAS VISUALLY INSPECTED UNDER A 10 X MAGNIFICATION FOR ANY DAMAGE AND SHOWED A MULTITUDE OF WRINKLES OVER THE WHOLE FOIL ALONG WITH SEVERAL PINHOLES AT TOP LABEL SIDE AS WELL AS BOTTOM CAVITY SIDE OF THE PACKS WERE OBSERVED. FIVE RETAIN SAMPLES OF INCIDENT CODES AND LOT NUMBER WERE RETRIEVED FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS LIKE PSEUDO SEAL, PIN HOLE, PRODUCT IN SEAL, PRESS MARKS ON PACK BUT NO SUCH DEFECTS WERE OBSERVED. THE PRIMARY PACKS WERE OPENED, AND SUTURES AND NEEDLES WERE FOUND INTACT. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE KINKS, WEAK SPOTS, BROKEN PIECE, FRAY, BRITTLENESS, DETACHED NEEDLES BUT NO SUCH DEFECTS WERE OBSERVED. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE BEND, CRACKED BARREL, FINS BUT NO SUCH DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WERE TESTED FOR KNOT PULL TENSILE STRENGTH TEST AND FOUND TO MEET THE SPECIFICATION. AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED FOR ANY PROCESS DEVIATION, BUT NO DEVIATION WAS OBSERVED. FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION. FROM THE ABOVE ANALYSIS, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. AS THE COMPLAINT IS RELATED TO BREAKAGE SUTURE, KNOT PULL TENSILE STRENGTH TEST IS APPLICABLE & MEASURABLE PARAMETER. KNOT PULL TENSILE STRENGTH TEST WAS PERFORMED OVER FIVE RETAIN SAMPLES TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. ALL THE INDIVIDUAL AS WELL AS AVERAGE KNOT PULL VALUES WERE FOUND TO MEET THE USP SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8550751&pc=GAM | 2210968-2019-80714 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT HYSTERECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGEON WAS PUTTING CONTINUOUS SUTURING AND TOWARDS THE END WHILE TRYING TO PUT KNOTS THE SUTURE BROKE UP. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8551760&pc=GAM | 2210968-2019-80719 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE WHEN SUTURING AT THE SKIN. THE PIECES WERE RETRIEVED AND THERE WAS NO PATIENT IMPACT. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8551813&pc=GAM | 2210968-2019-80720 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GASTRIC BYPASS PROCEDURE ON 3/1/2019 AND SUTURE WAS USED. THE SUTURE DISENGAGED FROM THE NEEDLE DURING THE PROCEDURE. ANOTHER DEVICE WAS USED TO SUCCESSFULLY COMPLETE THE PROCEDURE WITH 3 MINUTE DELAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8551815&pc=GAM | 2210968-2019-80721 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A GASTRIC BYPASS PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE DISENGAGED FROM THE NEEDLE DURING THE PROCEDURE. ANOTHER DEVICE WAS USED TO SUCCESSFULLY COMPLETE THE PROCEDURE WITH 3 MINUTE DELAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8551860&pc=GAM | 2210968-2019-80722 | 2019-03-28 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL PLUS SUTURE 36"(90CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT LAPAROSCOPIC OVARIAN CYST PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MD2660, AND NO NONCONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552113&pc=GAM | 2210968-2019-80726 | 2019-03-10 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Degraded; Material Separation; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE DETACHED FROM THE NEEDLE DURING USE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.¿
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552216&pc=GAM | 2210968-2019-80728 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-25 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE THE NEEDLE PULLED OFF THE SUTURE. THE FALLEN PIECE WAS RETRIEVED FROM THE PATIENT. ANOTHER SUTURE WAS USED TO COMPLETE THE SURGERY. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT MEASURES WERE TAKEN TO RETRIEVE THE BROKEN PIECE? UNK. WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? UNK. AT WHAT LOCATION WAS THE NEEDLE FOUND? UNK. WAS THE PATIENT BROUGHT BACK TO OR TO HAVE NEEDLE REMOVED OR WAS THE NEEDLE REMOVED DURING THE INITIAL PROCEDURE? UNK.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8544963&pc=GAM | 2210968-2019-80680 | 2019-03-28 | Malfunction | ETHICON INC. | 2019-04-24 | GAM | VICRYL RAP UND 70CM M1::USP5/0 SGLE ARMED RB-1 PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE WHEN MAKING THE KNOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8545009&pc=GAM | 2210968-2019-80681 | 2019-03-28 | Malfunction | ETHICON INC. | 2019-04-24 | GAM | VICRYL RAP UND 70CM M1::USP5/0 SGLE ARMED RB-1 PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE WHEN MAKING THE KNOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8545179&pc=GAM | 2210968-2019-80684 | 2019-03-21 | Malfunction | ETHICON INC. | 2019-04-24 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED BY HEALTH AUTHORITY THAT THE PATIENT UNDERWENT A MYOMECTOMY AND EXCISION OF RIGHT OVARIAN CYST PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING USE. THE BROKEN PIECE WAS RETRIEVED FROM THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MD2358, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8546740&pc=GAM | 2210968-2019-80691 | 2019-04-01 | Malfunction | ETHICON INC. | 2019-04-24 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT OSTEOTOMY ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS FOUND THAT THE NEEDLE TIP HAD BEEN BROKEN WHEN HOLDING THE NEEDLE BY NEEDLE-HOLDER. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MPZ884 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8542043&pc=GAM | 2210968-2019-80650 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-04-23 | GAM | VICRYL SUTURE 27"(70CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE TIP BROKE OFF. THE TIP WAS FOUND OUTSIDE THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8537346&pc=GAM | 2210968-2019-80579 | 2019-03-26 | Malfunction | ETHICON INC. | 2019-04-22 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN OB-GYN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE SEPARATED FROM THE SUTURE DURING USE. IT WAS NOT A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK725 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8533217&pc=GAM | 2210968-2019-80554 | 2019-03-16 | Malfunction | ETHICON INC. | 2019-04-19 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF WHILE SEWING SKIN. THE FALLEN PIECE WAS RETRIEVED FROM PATIENT. A LIKE DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MG7087, AND NO NONCONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8533249&pc=GAM | 2210968-2019-80555 | 2019-03-13 | Malfunction | ETHICON INC. | 2019-04-19 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF WHILE SEWING THE SKIN. THE FALLEN PIECE WAS RETRIEVED FROM PATIENT. A LIKE DEVICE WAS USED TO COMPLETE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MG7087, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8533407&pc=GAM | 2210968-2019-80547 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-19 | GAM | VICRYL SUTURE 27"(70CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING USE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NONCONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8527071&pc=GAM | 2210968-2019-80486 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-04-18 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAP GASTRIC BYPASS ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE DISENGAGED FROM THE NEEDLE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. NO ADVERSE PATIENT OUTCOME REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). EVALUATION SUMMARY: THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE WAS EXAMINED UNDER 20X MAGNIFICATION FOR SUTURE REMNANT AND NONE WAS NOTED. ADDITIONALLY, THERE ARE SLIGHTLY MARKS ON BODY THAT APPEARS TO BE BY USE OF SURGICAL INSTRUMENT. THE END OF SUTURE PIECE WAS EXAMINED AND THERE ISN'T SUFFICIENT IMPRESSION AT THE SWAGE END, RESULTING IN THE NEEDLE PULL OFF. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE IS A LIGHT SWAGE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? -NO. WAS PROCEDURE SUCCESSFULLY COMPLETED? YES. WERE FRAGMENTS GENERATED? UNKNOWN. IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? UNKNOWN. PATIENT STATUS/ OUTCOME / CONSEQUENCES NO. WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8528245&pc=GAM | 2210968-2019-80491 | 2019-03-25 | Malfunction | ETHICON INC. | 2019-04-18 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT POSTERIOR LUMBAR INTERBODY FUSION ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS DETACHED FROM THE NEEDLE EASILY DURING INTERRUPTED SUTURING ON THE FASCIA. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MM2310, AND NO NONCONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8528963&pc=FTL | 2210968-2019-80501 |  | Malfunction | ETHICON INC. | 2019-04-18 | FTL | VICRYL MESH 3X2CM | Peeled/Delaminated | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DENTAL PROCEDURE ON AN UNKNOWN DATE IN 2019 AND THE MESH WAS IMPLANTED. WHEN REMOVING THE DEVICE FROM THE STERILE BAG THE MEMBRANE DISINTEGRATED. A REPLACEMENT DEVICE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. "WHEN REMOVING THE DEVICE FROM THE STERILE BAG" IS THE STERILE BAG YOU ARE REFERRING TO THE PRIMARY PACKAGING? WAS THE PRIMARY PACKAGING INTACT? WHERE THERE ANY HOLES IN THE PACKAGING? PLEASE CONFIRM THE PRODUCT CODE. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8529184&pc=GAM | 2210968-2019-80506 | 2019-03-21 | Malfunction | ETHICON INC. | 2019-04-18 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN OVARIECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING USE, THE NEEDLE BROKE. THE BROKEN PIECE WAS RETRIEVED FROM THE PATIENT. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE IS NO REPORT ON PATIENT'S INJURY. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8530151&pc=GAM | 2210968-2019-80530 | 2019-03-22 | Malfunction | ETHICON INC. | 2019-04-18 | GAM | VICRYL CTD VIO 75CM M4 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN THE NEEDLE WAS PASSED THROUGH THE UTERINE WALL, THE SUTURE CAME OUT OF THE NEEDLE WITH THE FIRST BITE ON THE UTERUS. THE NEEDLE WAS CHECKED THOROUGHLY, AND CONFIRMED THAT NO NEEDLE PIECE LEFT INSIDE THE CAVITY. THE UTERUS WAS CLOSED WITH ANOTHER SUTURE AND THE PROCEDURE CONCLUDED UNEVENTFULLY. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8522722&pc=GAM | 2210968-2019-80445 |  | Malfunction | ETHICON, INC. | 2019-04-17 | GAM | VICRYL SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 638 | Event Description: THIS REPORT SUMMARIZES <NOE> 638 </NOE> MALFUNCTION EVENTS, INCLUDING 589 INITIAL EVENTS - 209 SUTURE BREAKAGE, 68 NEEDLE BREAKAGE AND 312 NEEDLE PULL-OFF THAT OCCURRED INTRAOPERATIVELY. IT ALSO SUMMARIZES 49 FOLLOW UP MALFUNCTION EVENTS, INCLUDING 24 SUTURE BREAKAGE, 5 NEEDLE BREAKAGE AND 20 NEEDLE PULL-OFF THAT OCCURRED INTRAOPERATIVELY. Manufacturer Narrative: SUTURE BREAKAGE MALFUNCTION EVENTS: 233, SUTURE BREAKAGE DEVICES RETURNED: 405, MANUFACTURING DEFICIENCY: 1. NEEDLE BREAKAGE MALFUNCTION EVENTS: 73, NEEDLE BREAKAGE DEVICES RETURNED: 122, MANUFACTURING DEFICIENCY: 6. NEEDLE PULL OFF MALFUNCTION EVENTS: 332, NEEDLE PULL OFF DEVICES RETURNED: 583, MANUFACTURING DEFICIENCY: 30. ALL DEVICES LABELED FOR SINGLE USE; 0 DEVICES REPROCESSED AND RE-USED 0 REMEDIAL ACTION . A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED. COMPLAINT INFORMATION WILL BE INCLUDED IN COMPLAINT TRENDING THAT IS REVIEWED ON A REGULAR BASIS TO DETERMINE IF FURTHER ACTION IS NECESSARY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8523559&pc=GAM | 2210968-2019-80454 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-04-17 | GAM | VICRYL VIOLET SUTURE FS-1 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OVARIECTOMY AND SUTURE WAS USED. THE SURGEON USED SUTURE TO PERFORM OVARIAN PEDICLE LIGATURES, MUSCULAR WALL SUTURES, SUBCUTANEOUS SUTURES AND CUTANEOUS SUTURES. NO BANDAGE COVERS THE WOUND. GREATER THAN 3 DAYS LATER, THE SUTURE BROKE. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518718&pc=GAM | 2210968-2019-80396 | 2019-02-14 | Malfunction | ETHICON INC. | 2019-04-16 | GAM | VICRYL RAPID UND 4-0 75 PS-2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DENTAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE RIPPED. UPON EVALUATION, THE SUTURE HAD BEGUN THE PROCESS OF DEGRADATION. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE DEVICE HAD BREAKAGE SUTURE ISSUES. A LABELED WINDING FORMER AND A NEEDLE-SUTURE PIECE WAS RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. A PIECE OF SUTURE WAS EXAMINED, AND THE END WAS OBSERVED TO BE DAMAGED (BUSTED). HOWEVER, DURING THE FUNCTIONAL TEST THE SUTURE WAS BROKEN; SINCE HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE TEST COULD NOT BE PERFORMED. THE PRODUCT CODE CONTAINS AN ABSORBABLE SUTURE AND DUE TO THE SAMPLE WAS RECEIVED OPENED, THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED AND NO FUNCTIONAL TEST CAN BE PERFORMED DUE TO SAMPLE CONDITION. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT. SIX UNOPENED SAMPLES WERE RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SIX UNOPENED SAMPLES, NO DEFECTS WERE OBSERVED ON THE PACKETS. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE EXAMINED ALONG OF STRAND AND NO DEFECTS OR DAMAGED WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED ON THE SAMPLES AND THE TENSILE FORCE MET THE REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE REPRESENTATIVE SAMPLES, NO PERFORMANCE BREAKAGE SUTURE WAS OBSERVED, AND THE TESTED SAMPLE MEET THE FINISHED GOODS REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519141&pc=GAM | 2210968-2019-80407 | 2015-06-25 | Malfunction | ETHICON INC. | 2019-04-16 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE EASILY RELEASED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE 539986 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519143&pc=GAM | 2210968-2019-80406 | 2015-06-25 | Malfunction | ETHICON INC. | 2019-04-16 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE EASILY RELEASED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE 539986 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519145&pc=GAM | 2210968-2019-80405 | 2015-06-25 | Malfunction | ETHICON INC. | 2019-04-16 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2015 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE EASILY RELEASED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE 539986 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8520344&pc=GAM | 2210968-2019-80417 | 2019-03-21 | Malfunction | ETHICON INC. | 2019-04-16 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE DID NOT FIX, IT BROKE DURING USE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL DEVICE REPORTED IN MW 2210968-2019-80415.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8520359&pc=GAM | 2210968-2019-80415 | 2019-03-21 | Malfunction | ETHICON INC. | 2019-04-16 | GAM | VICRYL SUTURE 36"(90CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE SEPARATED FROM THE SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL DEVICE REPORTED IN MW 2210968-2019-80417.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8513975&pc=GAM | 2210968-2019-80359 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-04-15 | GAM | VICRYL 3-0 1X75CM FS-1 VIOLET | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT A FEMALE DOG UNDERWENT MASTECTOMY ON UNKNOWN DATE AND SUTURE WAS USED. DURING SUTURING OF THE ABDOMINAL WALL OF THE SUBCUTANEOUS TISSUE, THE NEEDLE PULLED OFF THE THREAD AFTER SOME TISSUE CROSSINGS, WITHOUT ANY FORCE APPLIED. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, JV452 MMBCRHR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8515334&pc=GAM | 2210968-2019-80368 | 2019-03-20 | Malfunction | ETHICON INC. | 2019-04-15 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BROKEN. THE SURGERY WAS NOT DELAYED. IT WAS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8515764&pc=GAM | 2210968-2019-80373 | 2019-04-02 | Malfunction | ETHICON INC. | 2019-04-15 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 3-0 UND | Break; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS BROKEN WHEN TYING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE VR416, LJ9DKHE0 NUMBER, AND NO NON-CONFORMANCES RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8515766&pc=GAM | 2210968-2019-80372 | 2019-04-02 | Malfunction | ETHICON INC. | 2019-04-15 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS BROKEN WHEN TYING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE VR416, LJ9DKHE0 NUMBER, AND NO NON-CONFORMANCES RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8515769&pc=GAM | 2210968-2019-80374 | 2019-03-20 | Malfunction | ETHICON INC. | 2019-04-15 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE TIP OF THE NEEDLE BROKE DURING USE. IT WAS NOT REPORTED IF THE TIP WAS RETRIEVED. IT WAS NOT REPORTED HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8516790&pc=GAM | 2210968-2019-80384 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-15 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A FACE SKIN AND BACK CANCER REMOVAL ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED AWAY FROM THE THREAD. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL EVALUATION SUMMARY ¿ THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED AND REMNANT OF THE SUTURE WAS NOTED INTO THE BARREL HOLE OF NEEDLES. ALSO, SLIGHT MARKS WERE OBSERVED ON THE EDGE OF THE NEEDLES THAT APPEARS TO BE BY THE USE OF A SURGICAL INSTRUMENT. THE SUTURES WERE NOT RETURNED FOR ANALYSIS. MARKS WERE OBSERVED ON THE EDGE OF THE NEEDLE. IN ORDER TO AVOID THIS KIND OF DAMAGE THE PACKAGES INSERT (IFU) CAUTIONS: GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. PER THE SAMPLES CONDITION, THE ASSIGNABLE CAUSE SUGGESTS AN IMPROPER HANDLING OF THE SAMPLES. REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING A INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITIONS OF THE SAMPLES RECEIVED, THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS .
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8516799&pc=GAM | 2210968-2019-80385 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-15 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A FACE SKIN AND BACK CANCER REMOVAL ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED AWAY FROM THE THREAD. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL EVALUATION SUMMARY ¿ THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED AND REMNANT OF THE SUTURE WAS NOTED INTO THE BARREL HOLE OF NEEDLES. ALSO, SLIGHT MARKS WERE OBSERVED ON THE EDGE OF THE NEEDLES THAT APPEARS TO BE BY THE USE OF A SURGICAL INSTRUMENT. THE SUTURES WERE NOT RETURNED FOR ANALYSIS. MARKS WERE OBSERVED ON THE EDGE OF THE NEEDLE. IN ORDER TO AVOID THIS KIND OF DAMAGE THE PACKAGES INSERT (IFU) CAUTIONS: GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. PER THE SAMPLES CONDITION, THE ASSIGNABLE CAUSE SUGGESTS AN IMPROPER HANDLING OF THE SAMPLES. REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING A INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITIONS OF THE SAMPLES RECEIVED, THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS .
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8516804&pc=GAM | 2210968-2019-80386 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-15 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A FACE SKIN AND BACK CANCER REMOVAL ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED AWAY FROM THE THREAD. THERE WERE NO PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL EVALUATION SUMMARY ¿ REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING A INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITIONS OF THE SAMPLES RECEIVED, THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8509475&pc=GAM | 2210968-2019-80328 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-12 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON AN (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING USE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AND NO NONCONFORMANCE WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE SPECIFIC NUMBER OF DEVICES THAT FAILED DURING THIS ONE PROCEDURE? DEVICE RETURN STATUS/ FOLLOW UP.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8509530&pc=GAM | 2210968-2019-80329 |  | Malfunction | ETHICON INC. | 2019-04-12 | GAM | VICRYL VIO 75CM M3 USP2/0 SGLE ARMED FSL | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE AT THE LEVEL OF THE CRIMPING DURING USE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AND NO NONCONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8509846&pc=GAM | 2210968-2019-80330 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-12 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN OB-GYN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS EASILY DETACHED FROM THE NEEDLE DURING INTRAPERITONEAL CONTINUOUS SUTURING. IT WAS NOT REPORTED WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8510969&pc=GAM | 2210968-2019-80341 |  | Malfunction | ETHICON INC. | 2019-04-12 | GAM | VICRYL 3-0 1X75CM FS-1 VIOLET | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT A MASTECTOMY PROCEDURE ON AN UNKNOWN DATE IN (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE THREAD DURING USE. THE PROCEDURE WAS COMPLETED WITH AN ALTERNATE LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AND NO NONCONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511061&pc=GAM | 2210968-2019-80342 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-04-12 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE FOR FACE SKIN CANCER REMOVAL ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING CLOSING UP WHERE THE DEFECTS ARE, THE NEEDLE PULLED AWAY FROM THE THREAD. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511176&pc=GAM | 2210968-2019-80344 |  | Malfunction | ETHICON INC. | 2019-04-12 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE HAD BEEN DETACHING FROM THE SUTURE THREAD. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. ATTEMPTS ARE BEING MADE TO OBTAIN ADDITIONAL INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. SPECIFIC NUMBER OF PROCEDURES IN WHICH THE ISSUE OCCURRED? UNKNOWN. EVENT DATE? PROCEDURE NAME? EVENT DESCRIPTION INDICATING WHEN DID THE ISSUE OCCUR? QUANTITY INVOLVED? HOW WAS THE CASE COMPLETED? UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511393&pc=GAM | 2210968-2019-80347 | 2019-03-20 | Malfunction | ETHICON INC. | 2019-04-12 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN INGUINAL HERNIA REPAIR PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS BROKEN AT THE BODY PART DURING INTERRUPTED SUTURING. THE WAY OF SUTURING WAS AS USUAL. NO BROKEN PIECE OF NEEDLE WAS DROPPED INSIDE THE PATIENT¿S BODY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511815&pc=GAM | 2210968-2019-80349 | 2019-03-20 | Malfunction | ETHICON INC. | 2019-04-12 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE IS BROKEN. THE SURGERY WAS NOT DELAYED. IT WAS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO SPECIFIC DETAILS REGARDING THE ADDITIONAL EVENTS. ATTEMPTS ARE BEING MADE TO OBTAIN ADDITIONAL INFORMATION. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8505637&pc=GAM | 2210968-2019-80272 | 2019-03-18 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A MEDIPORT REMOVAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. AFTER THE MEDIPORT WAS REMOVED, THE SUTURE BROKE AT THE SWAGE WHILE CLOSING THE INCISION. THE PROCEDURE WAS COMPLETED WITH A SECOND LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8505950&pc=GAM | 2210968-2019-80280 | 2019-03-06 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A LAPAROSCOPIC CHOLECYSTECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE, AND THE BROKEN PIECE FELL ON THE GROUND. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8505999&pc=GAM | 2210968-2019-80282 | 2019-03-11 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A MEDIPORT REMOVAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE AT THE SWAGE WHILE CLOSING THE INCISION. THE PROCEDURE WAS COMPLETED WITH A SECOND LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8506018&pc=GAM | 2210968-2019-80283 | 2019-03-11 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A MEDIPORT REMOVAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE AT THE SWAGE WHILE CLOSING THE INCISION. THE PROCEDURE WAS COMPLETED WITH A SECOND LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8506082&pc=GAM | 2210968-2019-80286 |  | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL SUTURE 27"(70CM) 0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE BROKE IN HALF WHEN CLOSING THE INCISION. THE BROKEN PIECE WAS FOUND, AND THE PROCEDURE WAS NOT DELAYED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE NEEDLE BREAKAGE. A FAILED SUTURE NEEDLE WAS SUBMITTED FOR FRACTOGRAPHIC EVALUATION. A FRACTURE WAS OBSERVED IN THE BODY OF THE NEEDLE. THE NEEDLE WAS RECEIVED IN TWO PIECES, ONLY ONE OF THE MATING FRACTURE SURFACES WAS EXAMINED FOR THIS EVALUATION. A MICROSCOPE WAS USED TO EXAMINE THE FRACTURE SURFACES AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACES WERE EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. MECHANICAL DAMAGE AND A CUP-AND-CONE SHAPED FRACTURE OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDE ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS COULDN'T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED FROM MECHANICAL DEFORMATION, INDENTS AND SCRATCHES, DUE TO GRIPPING, BENDING AND OVERSTRESSING OF THE NEEDLE. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW THAT WOULD CAUSE PREMATURE FAILURE. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8506565&pc=GAM | 2210968-2019-80289 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT ABDOMINAL HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE THE NEEDLE TIP BROKE OFF. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM A DIFFERENT LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, XNV360G, MJ8CGJQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. A PHOTO OF THE PRODUCT HAS BEEN RECEIVED AND IS PENDING EVALUATION. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: ANY PATIENT CONSEQUENCE/AE OUTCOME AS A RESULT OF THE EVENT REPORT? TIME-DELAY DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON PATIENT? OR BEFORE USE ON PATIENT? 6 SUTURES FROM SAME BOX WERE OPENED, ALL GIVING SAME ISSUE. NOT ALL OF THEM WERE RETURNED. DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON SINGLE PATIENT EVENT/PROCEDURE OR MULTIPLE PATIENT EVENTS/PROCEDURES? SINGLE PATIENT, BOX WAS REMOVED FROM THEATRE BY UNIT MANAGER. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? UNKNOWN. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN. UNKNOWN. WHAT INSTRUMENTS WERE USED TO GRASP THE NEEDLE? NEEDLE HOLDER. WERE X-RAYS PERFORMED TO LOCALIZE THE NEEDLE FRAGMENT? NO. WERE THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? YES. DOES THE PIECE/DEVICE REMAIN RETAINED IN THE PATIENT'S TISSUE? NO. IF THE PIECE(S) WERE RETAINED, WHERE ARE THE LOCATED/IN WHAT STRUCTURE? NA. IF RETAINED, WERE THERE ANY PATIENT CONSEQUENCES? NA. NOTE: ADDITIONAL MEDWATCH REPORTS SUBMITTED VIA 2210968-2019-80294, 2210968-2019-80293, 2210968-2019-80292, 2210968-2019-80291, 2210968-2019-80290.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8506566&pc=GAM | 2210968-2019-80290 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT ABDOMINAL HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE THE NEEDLE TIP BROKE OFF. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM A DIFFERENT LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, XNV360G, MJ8CGJQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. A PHOTO OF THE PRODUCT HAS BEEN RECEIVED AND IS PENDING EVALUATION. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: ANY PATIENT CONSEQUENCE/AE OUTCOME AS A RESULT OF THE EVENT REPORT? TIME-DELAY. DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON PATIENT? OR BEFORE USE ON. PATIENT? 6 SUTURES FROM SAME BOX WERE OPENED, ALL GIVING SAME ISSUE. NOT ALL OF THEM WERE RETURNED. DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON SINGLE PATIENT. EVENT/PROCEDURE OR MULTIPLE PATIENT EVENTS/PROCEDURES? SINGLE PATIENT, BOX WAS REMOVED FROM THEATRE BY UNIT MANAGER. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? UNKNOWN. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN. UNKNOWN. WHAT INSTRUMENTS WERE USED TO GRASP THE NEEDLE? NEEDLE HOLDER. WERE X-RAYS PERFORMED TO LOCALIZE THE NEEDLE FRAGMENT? NO. WERE THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? YES. DOES THE PIECE/DEVICE REMAIN RETAINED IN THE PATIENT'S TISSUE? NO. IF THE PIECE(S) WERE RETAINED, WHERE ARE THE LOCATED/IN WHAT STRUCTURE? NA. IF RETAINED, WERE THERE ANY PATIENT CONSEQUENCES? NA. NOTE: ADDITIONAL MEDWATCH REPORTS SUBMITTED VIA 2210968-2019-80294, 2210968-2019-80293, 2210968-2019-80292, 2210968-2019-80291, 2210968-2019-80289.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8506568&pc=GAM | 2210968-2019-80291 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT ABDOMINAL HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE THE NEEDLE TIP BROKE OFF. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM A DIFFERENT LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, XNV360G, MJ8CGJQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. A PHOTO OF THE PRODUCT HAS BEEN RECEIVED AND IS PENDING EVALUATION. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: ANY PATIENT CONSEQUENCE/AE OUTCOME AS A RESULT OF THE EVENT REPORT? TIME-DELAY. DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON PATIENT? OR BEFORE USE ON. PATIENT? 6 SUTURES FROM SAME BOX WERE OPENED, ALL GIVING SAME ISSUE. NOT ALL OF THEM WERE RETURNED. DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON SINGLE PATIENT EVENT/PROCEDURE OR MULTIPLE PATIENT EVENTS/PROCEDURES? SINGLE PATIENT, BOX WAS REMOVED FROM THEATRE BY UNIT MANAGER. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? UNKNOWN. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN. UNKNOWN. WHAT INSTRUMENTS WERE USED TO GRASP THE NEEDLE? NEEDLE HOLDER. WERE X-RAYS PERFORMED TO LOCALIZE THE NEEDLE FRAGMENT? NO. WERE THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? YES. DOES THE PIECE/DEVICE REMAIN RETAINED IN THE PATIENT'S TISSUE? NO. IF THE PIECE(S) WERE RETAINED, WHERE ARE THE LOCATED/IN WHAT STRUCTURE? NA. IF RETAINED, WERE THERE ANY PATIENT CONSEQUENCES? NA. NOTE: ADDITIONAL MEDWATCH REPORTS SUBMITTED VIA 2210968-2019-80294, 2210968-2019-80293, 2210968-2019-80292, 2210968-2019-80290, 2210968-2019-80289.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8506572&pc=GAM | 2210968-2019-80292 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT ABDOMINAL HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE THE NEEDLE TIP BROKE OFF. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM A DIFFERENT LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, XNV360G, MJ8CGJQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. A PHOTO OF THE PRODUCT HAS BEEN RECEIVED AND IS PENDING EVALUATION. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: ANY PATIENT CONSEQUENCE/AE OUTCOME AS A RESULT OF THE EVENT REPORT? TIME-DELAY. DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON PATIENT? OR BEFORE USE ON. PATIENT? 6 SUTURES FROM SAME BOX WERE OPENED, ALL GIVING SAME ISSUE. NOT ALL OF THEM WERE RETURNED . DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON SINGLE PATIENT? EVENT/PROCEDURE OR MULTIPLE PATIENT EVENTS/PROCEDURES? SINGLE PATIENT, BOX WAS REMOVED FROM THEATRE BY UNIT MANAGER. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? UNKNOWN. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN. UNKNOWN. WHAT INSTRUMENTS WERE USED TO GRASP THE NEEDLE? NEEDLE HOLDER. WERE X-RAYS PERFORMED TO LOCALIZE THE NEEDLE FRAGMENT? NO. WERE THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? YES. DOES THE PIECE/DEVICE REMAIN RETAINED IN THE PATIENT'S TISSUE? NO. IF THE PIECE(S) WERE RETAINED, WHERE ARE THE LOCATED/IN WHAT STRUCTURE? NA. IF RETAINED, WERE THERE ANY PATIENT CONSEQUENCES? NA. NOTE: ADDITIONAL MEDWATCH REPORTS SUBMITTED VIA 2210968-2019-80294, 2210968-2019-80293, 2210968-2019-80291, 2210968-2019-80290, 2210968-2019-80289.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8506573&pc=GAM | 2210968-2019-80293 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT ABDOMINAL HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE THE NEEDLE TIP BROKE OFF. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM A DIFFERENT LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, XNV360G, MJ8CGJQ0 NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. A PHOTO OF THE PRODUCT HAS BEEN RECEIVED AND IS PENDING EVALUATION. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: ANY PATIENT CONSEQUENCE/AE OUTCOME AS A RESULT OF THE EVENT REPORT? TIME-DELAY. DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON PATIENT? OR BEFORE USE ON PATIENT? 6 SUTURES FROM SAME BOX WERE OPENED, ALL GIVING SAME ISSUE. NOT ALL OF THEM WERE RETURNED. DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON SINGLE PATIENT EVENT/PROCEDURE OR MULTIPLE PATIENT EVENTS/PROCEDURES? SINGLE PATIENT, BOX WAS REMOVED FROM THEATRE BY UNIT MANAGER. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? UNKNOWN. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN. UNKNOWN. WHAT INSTRUMENTS WERE USED TO GRASP THE NEEDLE? NEEDLE HOLDER. WERE X-RAYS PERFORMED TO LOCALIZE THE NEEDLE FRAGMENT? NO. WERE THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? YES. DOES THE PIECE/DEVICE REMAIN RETAINED IN THE PATIENT'S TISSUE? NO. IF THE PIECE(S) WERE RETAINED, WHERE ARE THEY LOCATED/IN WHAT STRUCTURE? NA. IF RETAINED, WERE THERE ANY PATIENT CONSEQUENCES? NA. NOTE: ADDITIONAL MEDWATCH REPORTS SUBMITTED VIA 2210968-2019-80294, 2210968-2019-80293, 2210968-2019-80292, 2210968-2019-80291, 2210968-2019-80290, 2210968-2019-80289.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8506574&pc=GAM | 2210968-2019-80294 | 2019-03-19 | Malfunction | ETHICON INC. | 2019-04-11 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT ABDOMINAL HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE THE NEEDLE TIP BROKE OFF. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE FROM A DIFFERENT LOT NUMBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, XNV360G, MJ8CGJQ0 NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. A PHOTO OF THE PRODUCT HAS BEEN RECEIVED AND IS PENDING EVALUATION. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: ANY PATIENT CONSEQUENCE/AE OUTCOME AS A RESULT OF THE EVENT REPORT? TIME-DELAY DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON PATIENT? OR BEFORE USE ON PATIENT? 6 SUTURES FROM SAME BOX WERE OPENED, ALL GIVING SAME ISSUE. NOT ALL OF THEM WERE RETURNED. DID ALL REPORTED QTY OF 6 NOTICED WITH ISSUE DURING USE ON SINGLE PATIENT? EVENT/PROCEDURE OR MULTIPLE PATIENT EVENTS/PROCEDURES? SINGLE PATIENT, BOX WAS REMOVED FROM THEATRE BY UNIT MANAGER. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? UNKNOWN. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN. UNKNOWN. WHAT INSTRUMENTS WERE USED TO GRASP THE NEEDLE? NEEDLE HOLDER. WERE X-RAYS PERFORMED TO LOCALIZE THE NEEDLE FRAGMENT? NO. WERE THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? YES, DOES THE PIECE/DEVICE REMAIN RETAINED IN THE PATIENT'S TISSUE? NO, IF THE PIECE(S) WERE RETAINED, WHERE ARE THE LOCATED/IN WHAT STRUCTURE? NA. IF RETAINED, WERE THERE ANY PATIENT CONSEQUENCES? NA, NOTE: ADDITIONAL MEDWATCH REPORTS SUBMITTED VIA 2210968-2019-80294, 2210968-2019-80293, 2210968-2019-80292, 2210968-2019-80291, 2210968-2019-80290, 2210968-2019-80289.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8499634&pc=GAM | 2210968-2019-80198 | 2019-03-13 | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF. THE FALLEN PIECE WAS RETRIEVED FROM THE PATIENT. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE IS NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8499640&pc=GAM | 2210968-2019-80200 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE PHYSICIAN REPORTED THAT DURING THE PROCEDURE IN THE OFFICE, THE DEVICE DID NOT FEEL LIKE MULITIPACK NEEDLES. IT WAS REPORTED THAT NONE OF THE NEEDLES WERE NECESSARILY DULL, JUST DIDN'T FEEL SHARP THAN USUAL. IT WAS ALSO REPORTED THERE WAS AN ISSUE WITH SUTURE BREAKAGE WITH THE SUTURE MATERIAL AND NEEDLES POPPING OFF. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. FOUR UNOPENED SAMPLES OF PRODUCT CODE (B)(4), LOT # MJK059 WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE AND TENSILE STRENGTH WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLE RECEIVED, NO PULL OFF SUTURE NEEDLE, BREAKAGE SUTURE AND SURFACE RELATED DEFECT - NEEDLE WERE FOUND, AND THE TESTED SAMPLE MET THE FINISHED GOODS REQUIREMENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY SUTURE BROKE, PULLED OFF DURING THIS ONE PATIENT PROCEDURE? - ALL EXCEPT ONE. HOW MANY WERE NOTED TO BE DULL DURING THIS ONE PATIENT PROCEDURE? - NONE NECESSARY DULL, JUST DIDN'T FEEL SHARP AND CUSTOMER ADVISED THAT ISSUED OCCURRED WITH 2 OR 3 PACKS, NOT BOXES. NOTE: MEDWATCH FOR ADDITIONAL DEVICES SUBMITTED VIA 2210968-2019-80199, 2210968-2019-79442.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8499644&pc=GAM | 2210968-2019-80199 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE PHYSICIAN REPORTED THAT DURING THE PROCEDURE IN THE OFFICE, THE DEVICE DID NOT FEEL LIKE MULITPACK NEEDLES. IT WAS REPORTED THAT NONE OF THE NEEDLES WERE NECESSARILY DULL, JUST DIDN¿T FEEL SHARP THAN USUAL. IT WAS ALSO REPORTED THERE WAS AN ISSUE WITH SUTURE BREAKAGE WITH THE SUTURE MATERIAL AND NEEDLES POPPING OFF. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. (B)(6) FOUR UNOPENED SAMPLES OF PRODUCT CODE J497, LOT # MJK059 WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE AND TENSILE STRENGTH WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLE RECEIVED, NO PULL OFF SUTURE NEEDLE, BREAKAGE SUTURE AND SURFACE RELATED DEFECT - NEEDLE WERE FOUND, AND THE TESTED SAMPLE MET THE FINISHED GOODS REQUIREMENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY SUTURE BROKE, PULLED OFF DURING THIS ONE PATIENT PROCEDURE? ALL EXCEPT ONE. HOW MANY WERE NOTED TO BE DULL DURING THIS ONE PATIENT PROCEDURE? NONE NECESSARY DULL, JUST DIDN¿T FEEL SHARP AND CUSTOMER ADVISED THAT ISSUED OCCURRED WITH 2 OR 3 PACKS, NOT BOXES. NOTE: MEDWATCH FOR ADDITIONAL DEVICES SUBMITTED VIA 2210968-2019-79442, 2210968-2019-80200.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8500807&pc=GAM | 2210968-2019-80218 | 2019-02-28 | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. NO PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. EVALUATION: AN UNOPENED SAMPLE OF PRODUCT WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE, NO DEFECTS WERE FOUND ON THE PACKAGE. THE SAMPLE WAS OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURE WAS DISPENSED WITHOUT PROBLEM AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE REPRESENTATIVE SAMPLE, NO PERFORMANCE - BREAKAGE SUTURE WAS FOUND AND THE TESTED SAMPLE MET THE FINISHED GOODS REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8500826&pc=GAM | 2210968-2019-80223 | 2019-02-28 | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. NO PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8500829&pc=GAM | 2210968-2019-80224 | 2019-02-28 | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. NO PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8500835&pc=GAM | 2210968-2019-80225 | 2019-02-28 | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. NO PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8500838&pc=GAM | 2210968-2019-80226 | 2019-02-28 | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. NO PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8500855&pc=GAM | 2210968-2019-80227 | 2019-03-13 | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE IS NO REPORT ON PATIENT'S INJURY. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8501212&pc=GAM | 2210968-2019-80239 |  | Malfunction | ETHICON INC. | 2019-04-10 | GAM | VICRYL SUTURE 18"(45CM) 5-0 UND | Degraded; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE WHEN THE PACKAGE WAS OPENED, THE SUTURE WAS FOUND BROKEN. THE PRODUCT WAS NOT USED ON A PATIENT. DURING THE ANALYSIS, THE SUTURE COULD NOT BE DISPENSED SINCE IT HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: THE ACTUAL DEVICE WAS RECEIVED AS AN EMPTY OPENED FOIL, A PAPER LID AND A SUTURE INSIDE OF WINDING FORMER OF PRODUCT CODE J493, LOT LGM371. DURING THE VISUAL INSPECTION OF THE SAMPLE, THE SWAGE AND ATTACHMENT AREA THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURE COULD NOT BE DISPENSED SINCE IT HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED. THE FOIL WAS EXAMINED AND SHOWED MULTIPLES WRINKLES AND HOLES IN CAVITY ON BOTTOM FOIL PACKAGE DUE TO EXCESSIVE MANIPULATION. PER THE CONDITION OF THE SAMPLE, THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE PRE-OP, IS SUTURE DEGRADATION; DUE TO THE IMPROPER HANDLING OF PACKAGE. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LGM371 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8495996&pc=GAM | 2210968-2019-80150 | 2019-03-15 | Malfunction | ETHICON INC. | 2019-04-09 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON 3/15/2019 AND SUTURE WAS USED. THE SUTURE BROKE WHILE TYING THE KNOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8496025&pc=GAM | 2210968-2019-80151 | 2019-03-15 | Malfunction | ETHICON INC. | 2019-04-09 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE WHILE TYING THE KNOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8496300&pc=GAM | 2210968-2019-80160 | 2019-03-11 | Malfunction | ETHICON INC. | 2019-04-09 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. SUTURE BROKEN DURING USE. NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED. EVALUATION: WE RECEIVED 01 COMPLAINT SAMPLE FOIL ALONG WITH SUTURE, NEEDLE, ZIPPER TRAY AND LID FOR CODE (B)(4), LOT T5006. SUTURE RECEIVED IN PARTIAL TO COMPLETE DISINTEGRATED CONDITION. AS THE COMPLAINT SAMPLE RECEIVED IN OPEN CONDITION FURTHER INVESTIGATION ON COMPLAINT SAMPLE CANNOT BE PERFORMED EXCEPT VISUAL INSPECTION. THE FOIL PACK WAS INSPECTED UNDER MAGNIFICATION FOR ANY DAMAGE AND SHOWED A MULTITUDE OF WRINKLES OVER THE WHOLE FOIL ALONG WITH SEVERAL PINHOLES AT TOP LABEL SIDE AS WELL AS BOTTOM CAVITY SIDE OF THE PACKS WERE OBSERVED. RETURNED NEEDLE WAS INSPECTED UNDER MAGNIFICATION AND FOUND SATISFACTORY. FIVE RETAIN SAMPLES OF INCIDENT CODES AND LOT NUMBER WERE RETRIEVED FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS, NO DEFECTS WERE OBSERVED. THE PRIMARY PACKS WERE OPENED, AND SUTURES AND NEEDLES WERE FOUND INTACT. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS. NO DEFECTS WERE OBSERVED. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTE DEFECTS. NO DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WERE TESTED FOR KNOT PULL TENSILE STRENGTH TEST AND FOUND TO MEET THE SPECIFICATION. AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED FOR ANY PROCESS DEVIATION, BUT NO DEVIATION WAS OBSERVED. FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION. FROM THE ABOVE ANALYSIS, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. AS THE COMPLAINT IS RELATED TO PERFORMANCE-BREAKAGE SUTURE, KNOT PULL TENSILE STRENGTH TEST IS APPLICABLE & MEASURABLE PARAMETER. KNOT PULL TENSILE STRENGTH TEST WAS PERFORMED OVER FIVE RETAIN SAMPLES TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. THE AVERAGE KNOT PULL TENSILE STRENGTH VALUE OF THE RETAIN SAMPLE WAS FOUND TO MEET THE USP AVERAGE KNOT PULL TENSILE STRENGTH REQUIREMENT. ALL THE INDIVIDUAL KNOT PULL TENSILE STRENGTH VALUES WERE FOUND TO MEET THE USP SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS. IF THE ANY FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8496732&pc=GAM | 2210968-2019-80169 | 2019-02-22 | Malfunction | ETHICON INC. | 2019-04-09 | GAM | VICRYL SUTURE 27"(70CM) 0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE NEEDLE BROKE OFF INSIDE THE PATIENT DURING SURGERY. IT IS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE BROKEN NEEDLE PIECE REMOVED DURING THE INITIAL SURGERY? WAS THERE ANY ADDITIONAL MEDICAL OR SURGICAL INTERVENTION REQUIRED? IS THE BROKEN NEEDLE AVAILABLE TO BE RETURNED FOR EVALUATION? THIS MEDWATCH REPORT IS IN RESPONSE TO RECEIPT OF SUS VOLUNTARY EVENT REPORT MW5084549.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8491780&pc=GAM | 2210968-2019-80095 | 2019-01-16 | Malfunction | ETHICON INC. | 2019-04-08 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND A SUTURE WAS USED. DURING SURGERY, THE SUTURE BROKE FIVE TIMES. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492118&pc=GAM | 2210968-2019-80098 | 2019-03-12 | Malfunction | ETHICON INC. | 2019-04-08 | GAM | VICRYL PLUS SUTURE 36"(90CM) 3-0 UND | Component Incompatible | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS OPENED AND PASSED OFF TO THE FIELD AND THE ACTUAL SUTURE THAT WAS INSIDE THE PACKAGE WAS NOT THE SAME. THE FOIL WRAP SAID ONE PRODUCT CODE AND THE ACTUAL SUTURE WAS DIFFERENT. ANOTHER LIKE DEVICE WITH THE CORRECT PRODUCT CODE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492583&pc=GAM | 2210968-2019-80114 | 2019-03-14 | Malfunction | ETHICON INC. | 2019-04-08 | GAM | VICRYL ABS SUR SUT USP SYN 4-0 70CM | Material Separation; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN THE DOCTOR WAS ABOUT TO SUTURE THE SKIN LAYER, THE NEEDLE AND SUTURE THREAD GOT SEPARATED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE WERE IDENTIFIED. INVESTIGATION SUMMARY: WE RECEIVED 01 OPENED COMPLAINT SAMPLE FOIL ALONG WITH ZIPPER TRAY, LID. SUTURE AND NEEDLE WAS NOT RETURNED WITH COMPLAINT SAMPLE. THE FOIL PACK WAS INSPECTED UNDER A 10 X MAGNIFICATION FOR ANY DAMAGE AND SHOWED A MULTITUDE OF WRINKLES OVER THE WHOLE FOIL ALONG WITH SEVERAL PINHOLES AT TOP LABEL SIDE AS WELL AS BOTTOM CAVITY SIDE OF THE PACKS WERE OBSERVED. FIVE RETAIN SAMPLES OF INCIDENT CODES AND LOT NUMBER WERE RETRIEVED FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS LIKE PSEUDO SEAL, PIN HOLE, PRODUCT IN SEAL, PRESS MARKS ON PACK BUT NO SUCH DEFECTS WERE OBSERVED. THE PRIMARY PACKS WERE OPENED, AND SUTURES AND NEEDLES WERE FOUND INTACT. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE KINKS, WEAK SPOTS, BROKEN PIECE, FRAY, BRITTLENESS, DETACHED NEEDLES BUT NO SUCH DEFECTS WERE OBSERVED. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE BEND, CRACKED BARREL, FINS BUT NO SUCH DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WERE TESTED FOR NEEDLE PULL TEST AND FOUND TO MEET THE SPECIFICATION. AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED FOR ANY PROCESS DEVIATION, BUT NO DEVIATION WAS OBSERVED. FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION. FROM THE ABOVE ANALYSIS, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. AS THE COMPLAINT IS RELATED TO PERFORMANCE-PULL OFF SUTURE NEEDLE, NEEDLE PULL OFF TEST IS APPLICABLE & MEASURABLE PARAMETER. ALL THE INDIVIDUAL AS WELL AS AVERAGE NEEDLE PULL-OFF VALUES WERE FOUND TO MEET THE USP SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492585&pc=GAM | 2210968-2019-80115 | 2019-03-14 | Malfunction | ETHICON INC. | 2019-04-08 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN THE DOCTOR WAS ABOUT TO SUTURE THE SKIN LAYER, THE NEEDLE AND SUTURE THREAD GOT SEPARATED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492843&pc=GAM | 2210968-2019-80123 |  | Malfunction | ETHICON INC. | 2019-04-08 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE KEPT BREAKING. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8494230&pc=GAM | 2210968-2019-80141 |  | Malfunction | ETHICON INC. | 2019-04-08 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE THE SUTURE WAS BREAKING. IT IS UNKNOWN IF ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8487836&pc=GAM | 2210968-2019-80059 | 2019-03-13 | Malfunction | ETHICON INC. | 2019-04-05 | GAM | VICRYL ABS SUR SUT USP SYN | Break; Material Separation; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE SURGERY, THE SUTURE WAS BROKEN AND THE NEEDLE SEPARATED FROM THE SUTURE. NO ADVERSE PATIENT OUTCOME REPORTED. Manufacturer Narrative: (B)(4). EVALUATION: ONE OPENED COMPLAINT SAMPLE FOIL PACK ALONG WITH ZIPPER TRAY LID WAS RECEIVED. SUTURE MATERIAL WAS INSPECTED UNDER MAGNIFICATION AND SUTURE FOUND IN PARTIAL TO COMPLETE DISINTEGRATED CONDITION AS THE COMPLAINT SAMPLE RECEIVED IN OPEN CONDITION FURTHER INVESTIGATION ON COMPLAINT SAMPLE CANNOT PERFORMED EXCEPT VISUAL INSPECTION. THE FOIL PACK WAS VISUALLY INSPECTED UNDER MAGNIFICATION FOR ANY DAMAGE AND SHOWED A MULTITUDE OF WRINKLES OVER THE WHOLE FOIL ALONG WITH SEVERAL PINHOLES AT THE TOP LABEL SIDE AS WELL AS ON BOTTOM CAVITY SIDE OF THE PACKS WERE OBSERVED. NEEDLE WAS NOT RETURNED FOR INVESTIGATION. FIVE RETAIN SAMPLES OF INCIDENT CODES AND LOT NUMBER WAS RETRIEVED FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS LIKE PSEUDO SEAL, PIN HOLE, PRODUCT IN SEAL, PRESS MARKS ON PACK BUT NO SUCH DEFECTS WERE OBSERVED. THE PRIMARY PACKS WERE OPENED, AND SUTURES AND NEEDLES WERE FOUND INTACT. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE KINKS, WEAK SPOTS, BROKEN PIECE, FRAY, BRITTLENESS, DETACHED NEEDLES BUT NO SUCH DEFECTS WERE OBSERVED. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE BEND, CRACKED BARREL, FINS BUT NO SUCH DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WERE TESTED FOR KNOT PULL TENSILE STRENGTH & NEEDLE PULL TEST AND FOUND TO MEET THE SPECIFICATION. AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED FOR ANY PROCESS DEVIATION, BUT NO DEVIATION WAS OBSERVED. FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION REQUIREMENT. FROM THE ABOVE ANALYSIS, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. AS THE COMPLAINT IS RELATED TO PERFORMANCE ¿ BREAKAGE SUTURE AS WELL AS PERFORMANCE-PULL OFF SUTURE NEEDLE, KNOT PULL TENSILE STRENGTH TEST AS WELL AS NEEDLE PULL TEST IS APPLICABLE & MEASURABLE PARAMETER. ALL THE INDIVIDUAL NEEDLE PULL-OFF AND KNOT PULL TENSILE STRENGTH VALUES WERE FOUND TO MEET THE USP SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8489048&pc=GAM | 2210968-2019-80079 | 2019-03-11 | Malfunction | ETHICON INC. | 2019-04-05 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING USE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJ2901, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8484674&pc=GAM | 2210968-2019-80028 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-04-04 | GAM | VICRYL SUTURE 18"(45CM) 6-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. AS THE VETERINARY SURGEON WENT TO TIE A KNOT IN THE SUTURE, THE SUTURE MATERIAL DISINTEGRATED. IT WAS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8484745&pc=GAM | 2210968-2019-80029 | 2019-03-08 | Malfunction | ETHICON INC. | 2019-04-04 | GAM | VICRYL PLUS SUTURE 27"(70CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS REPORTED THAT SUTURE BREAKAGE OCCURRED DURING SUTURE PROCESS. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE WAS NO ADVERSE PATIENT OUTCOME REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8484898&pc=GAM | 2210968-2019-80030 |  | Malfunction | ETHICON INC. | 2019-04-04 | GAM | VICRYL VIO 75CM M3.5::USP0 SGLE ARMED CT | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS EASY TO BREAK. IT WAS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8478921&pc=GAM | 2210968-2019-79965 | 2019-03-08 | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL PLUS SUTURE 18"(45CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING USE, THE NEEDLE WAS BROKEN. THE LOT NUMBER IS UNKNOWN. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8479566&pc=GAM | 2210968-2019-79981 | 2019-03-02 | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A CESAREAN BIRTH ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULL OFF DURING USE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE IS NO REPORTED PATIENT CONSEQUENCE. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8479625&pc=GAM | 2210968-2019-79978 |  | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE SNAPPED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBER: BATCH LP8BLJP0. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE VERIFY THE LOT NUMBER AS IT MAY BE INVALID. COULD IT BE LP8BLJP0 INSTEAD OF LP8BLJPO? YES, REGARDING THE LOT NUMBER. I AM PICKING UP THE SUTURES OF THE SAME LOT SO WILL BE ABLE TO CONFIRM THIS MYSELF ALSO ON THAT DAY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8479630&pc=GAM | 2210968-2019-79977 |  | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE SNAPPED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE INTERNATIONAL AFFILIATE REPORTS THE FOLLOWING POSSIBLE BATCH NUMBER: BATCH LP8BLJP0. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE VERIFY THE LOT NUMBER AS IT MAY BE INVALID. COULD IT BE LP8BLJP0 INSTEAD OF LP8BLJPO? YES, REGARDING THE LOT NUMBER. I AM PICKING UP THE SUTURES OF THE SAME LOT SO WILL BE ABLE TO CONFIRM THIS MYSELF ALSO ON THAT DAY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8479948&pc=GAM | 2210968-2019-79994 | 2019-03-10 | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL PLUS SUTURE 18"(45CM) 0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN OB-GYN PROCEDURE LAPAROSCOPIC ADNEXECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE AT THE SWAGE PART DURING CLOSING WOUND OF THE FASCIA. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8480352&pc=GAM | 2210968-2019-80000 | 2019-03-11 | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL HIP REPLACEMENT PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE SWAGE END CAME AWAY FROM THE NEEDLE DURING USE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJ7651, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8480377&pc=GAM | 2210968-2019-80001 | 2019-03-11 | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL HIP REPLACEMENT PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE SWAGE END CAME AWAY FROM THE NEEDLE DURING USE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJ7651, AND NO NON-CONFORMANCE WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8480400&pc=GAM | 2210968-2019-80002 | 2019-03-11 | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL HIP REPLACEMENT PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE SWAGE END CAME AWAY FROM THE NEEDLE DURING USE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJ7651, AND NO NON-CONFORMANCES WERE IDENTIFIED
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8481167&pc=GAM | 2210968-2019-80005 | 2018-12-29 | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 UND | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. IT WAS REPORTED THAT NEEDLE BREAKAGE OCCURRED DURING THE SUTURE TIME. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE WAS NO ADVERSE PATIENT OUTCOME REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8481185&pc=GAM | 2210968-2019-80006 | 2019-03-12 | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS FOUND THAT THE NEEDLE HAD BEEN DETACHED FROM THE SUTURE. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8481200&pc=GAM | 2210968-2019-80008 | 2019-02-20 | Malfunction | ETHICON INC. | 2019-04-03 | GAM | VICRYL PLUS SUTURE 36"(90CM) 0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE BENT AND SUTURE BREAKAGE OCCURRED DURING THE SUTURE OF UTERUS. CHANGED ANOTHER ONE TO COMPLETE SURGERY. THERE WAS NO ADVERSE PATIENT OUTCOME. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED LOT AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475616&pc=GAM | 2210968-2019-79913 | 2019-02-13 | Malfunction | ETHICON INC. | 2019-04-02 | GAM | VICRYL RAP UND 75CM M1 USP5/0 SGLE ARMED C-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A DENTAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED AWAY FROM THE THREAD AND STAYED IN THE GINGIVA OF THE PATIENT. IT WAS DIFFICULT TO RETRIEVE THE NEEDLE BECAUSE IT WAS SMALL. THE NEEDLE WAS RETRIEVED. PATIENT CONSEQUENCE NOT DETAILED. ADDITIONAL INFORMATION REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS AND OR THE DEVICE ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE DESCRIBE WHAT MINOR CONSEQUENCES DID THE PATIENT SUFFER? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THE CONSEQUENCE? WAS THE NEEDLE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? WAS THERE ANY UNEXPECTED OUTCOME OR COMPLICATION DUE TO THIS ISSUE RESULTING INTO ALTERATION IN TREATMENT PLAN/POST-OP CARE? PATIENT'S PRESENT CONDITION
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8476603&pc=GAM | 2210968-2019-79933 |  | Malfunction | ETHICON INC. | 2019-04-02 | GAM | VICRYL SUTURE 8-27" (70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NEPHROURETERECTOMY PROCEDURE IN (B)(6) 2019 AND SUTURE WAS USED. WHEN RETURNING THE NEEDLE-HOLDER TO A NURSE AFTER SUTURING ON THE PERITONEUM, IT WAS FOUND THAT THE NEEDLE HAD BEEN MISSING. THE MISSING NEEDLE WAS FOUND, BUT THERE WAS NOT NEEDLE TIP. THE MISSING NEEDLE TIP WAS SEARCHED USING MRI, BUT IT COULD NOT BE FOUND, SO THE WOUND WAS CLEANED AND CLOSED. THE PATIENT'S CONDITION IS STABLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IS THE NEEDLE TIP RETAINED IN THE PATIENT¿S TISSUE? UNKNOWN. THE NEEDLE TIP WAS NOT FOUND. IF RETAINED, ARE THERE ANY PLANS TO REMOVE THE NEEDLE TIP? NO FURTHER INFORMATION WILL BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8476610&pc=GAM | 2210968-2019-79934 | 2019-03-07 | Malfunction | ETHICON INC. | 2019-04-02 | GAM | VICRYL SUTURE 6-30" (75CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN CHEST SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS BROKEN DURING LIGATION ON THE RIGHT PULMONARY ARTERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8468978&pc=GAM | 2210968-2019-79835 | 2019-03-04 | Malfunction | ETHICON INC. | 2019-04-01 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE SPLIT DURING THE SURGERY. USED ANOTHER SUTURE TO COMPLETE. THERE WAS NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8469299&pc=GAM | 2210968-2019-79838 | 2019-02-09 | Malfunction | ETHICON INC. | 2019-04-01 | GAM | VICRYL PLUS SUTURE 27"(70CM) 5-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF DURING THE PROCEDURE. CHANGED ANOTHER ONE TO COMPLETE SURGERY. NO ADDITIONAL INFORMATION IS AVAILABLE. NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8472005&pc=GAM | 2210968-2019-79883 | 2018-02-13 | Malfunction | ETHICON INC. | 2019-04-01 | GAM | VICRYL RAP UND 75CM M1 USP5/0 SGLE ARMED C-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DENTAL SURGERY ON (B)(6)2018 AND SUTURE WAS USED. DURING USE, THE NEEDLE PULLED AWAY FROM THE THREAD AND STAYED INTO THE GINGIVA OF THE PATIENT. IT WAS DIFFICULT TO RETRIEVE THE NEEDLE BECAUSE IT WAS SMALL. THERE WERE MINOR PATIENT CONSEQUENCES BECAUSE THE NEEDLE WAS RETRIEVED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE DESCRIBE WHAT MINOR CONSEQUENCES DID THE PATIENT SUFFER? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THE CONSEQUENCE? WAS THE NEEDLE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? WAS THERE ANY UNEXPECTED OUTCOME OR COMPLICATION DUE TO THIS ISSUE RESULTING INTO ALTERATION IN TREATMENT PLAN/POST-OP CARE? PATIENT'S PRESENT CONDITION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8464187&pc=GAM | 2210968-2019-79801 | 2019-03-07 | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HARDWARE REMOVAL TENDON TRANSFER PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE DETACHED DURING SUTURING. THE PROCEDURE WAS COMPLETED WITH THIS SAME SUTURE MATERIAL USING MUCH SLOWER PASS THROUGH AND PULLING OF SUTURE. IT TOOK LONGER TO CLOSE INCISION DUE TO EXTRA PRECAUTION USED TO KEEP NEEDLE FROM DETACHING. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8464346&pc=GAM | 2210968-2019-79802 | 2019-03-07 | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HARDWARE REMOVAL TENDON TRANSFER PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE DETACHED DURING SUTURING. THE PROCEDURE WAS COMPLETED WITH THIS SAME SUTURE MATERIAL USING MUCH SLOWER PASS THROUGH AND PULLING OF SUTURE. IT TOOK LONGER TO CLOSE INCISION DUE TO EXTRA PRECAUTION USED TO KEEP NEEDLE FROM DETACHING. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8465008&pc=GAM | 2210968-2019-79806 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT MOHS REPAIR SURGERY & EXCISION IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE OFF FROM THE SUTURE. THE PROCEDURE WAS COMPLETED USING A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 BATCH NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8465014&pc=GAM | 2210968-2019-79807 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT MOHS REPAIR SURGERY & EXCISION IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE OFF FROM THE SUTURE. THE PROCEDURE WAS COMPLETED USING A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 BATCH NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8465015&pc=GAM | 2210968-2019-79808 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT MOHS REPAIR SURGERY & EXCISION IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE OFF FROM THE SUTURE. THE PROCEDURE WAS COMPLETED USING A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 BATCH NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8465416&pc=GAM | 2210968-2019-79810 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL VIO 70CM M3 USP2/0 SGLE ARMED CT-1 PLUS | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE SEPARATED FROM SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8465419&pc=GAM | 2210968-2019-79811 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL VIO 70CM M3 USP2/0 SGLE ARMED CT-1 PLUS | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE SEPARATED FROM SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8465420&pc=GAM | 2210968-2019-79812 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL VIO 70CM M3 USP2/0 SGLE ARMED CT-1 PLUS | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE SEPARATED FROM SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8465708&pc=GAM | 2210968-2019-79816 |  | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL PLUS SUTURE 18"(45CM) 1 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS FOUND THAT THE NEEDLE HAD BEEN DETACHED FROM THE SUTURE. IT WAS A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8465992&pc=GAM | 2210968-2019-79820 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-03-29 | GAM | VICRYL*RAPID UND 4-0 70CM 1PS2 | Break; Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A BREAST MAMMOPLASTY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE IS DISASSEMBLED FROM THE NEEDLE. THE NEEDLE WAS BENT AND WHEN KNOTTING BROKE. THERE WAS NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW WAS CASE COMPLETED? LOT NUMBER PRODUCT LOT NUMBER: AH5311. NAME OF THE SURGEON: DR. (B)(6). NAME OF THE HOSPITAL: MARLY CLINIC. NO SAMPLES AVAILABLE, IT WAS DISCARDED AT THE INSTITUTION. THE PATIENT DID NOT HAVE ANY INJURY. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE VERIFY THE EVENT- DID 1 DEVICE HAD ALL 3 ISSUES- NEEDLE BENDING, NEEDLE PULL OFF AND SUTURE BREAKAGE ISSUE? IF NO, PLEASE EXPLAIN QUANTITY INVOLVED AND SPECIFIC ISSUE WITH EACH DEVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8459991&pc=GAM | 2210968-2019-79760 |  | Malfunction | ETHICON INC. | 2019-03-28 | GAM | VICRYL SUTURE 27"(70CM) 4-0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE CUSTOMER EXPERIENCED THE SUTURE FRAYING AND COMING AWAY FROM THE SWAGE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. TYPE OF PROCEDURE. AT WHAT STAGE OF THE PROCEDURE DID THE ISSUE OCCUR. HOW MANY SUTURES INVOLVED. WHAT STEPS WERE TAKEN TO RECTIFY THE SITUATION (I.E. OPENED A NEW SUTURE PACKET ETC..). ANY IMPACT ON THE PATIENT BECAUSE OF THE ISSUE RAISED. IS THE IMPACTED PRODUCT AVAILABLE FOR COLLECTION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8461168&pc=GAM | 2210968-2019-79768 |  | Malfunction | ETHICON INC. | 2019-03-28 | GAM | VICRYL PLUS SUTURE 18"(45CM) 4-0 UND | Break; Material Frayed | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT PACEMAKER PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE AND FRAYED DURING USE. THE STRANDS UNRAVELED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MG6771 BATCH NUMBER, AND NO NON-CONFORMANCE'S RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8462812&pc=GAM | 2210968-2019-79787 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-03-28 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE DISENGAGED FROM THE NEEDLE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. NO ADVERSE PATIENT OUTCOME REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AND NO NON-CONFORMANCE'S RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? --> NO. WAS PROCEDURE SUCCESSFULLY COMPLETED? --> YES. WERE FRAGMENTS GENERATED? --> UNKNOWN. IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? --> UNKNOWN. PATIENT STATUS/ OUTCOME / CONSEQUENCES --> NO. WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC. REVISION) REQUIRED: --> UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8462823&pc=GAM | 2210968-2019-79788 | 2019-03-01 | Malfunction | ETHICON INC. | 2019-03-28 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAP GASTRIC BYPASS ON 3/1/2019 AND SUTURE WAS USED. THE SUTURE DISENGAGED FROM THE NEEDLE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. NO ADVERSE PATIENT OUTCOME REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AND NO NON-CONFORMANCE RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? --> NO. WAS PROCEDURE SUCCESSFULLY COMPLETED? --> YES. WERE FRAGMENTS GENERATED? --> UNKNOWN. IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? --> UNKNOWN. PATIENT STATUS/ OUTCOME / CONSEQUENCES --> NO. WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC. REVISION) REQUIRED: --> UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8462881&pc=GAM | 2210968-2019-79791 |  | Malfunction | ETHICON INC. | 2019-03-28 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HARDWARE REMOVAL TENDON TRANSFER PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE DETACHED DURING SUTURING. THE PROCEDURE WAS COMPLETED WITH THIS SAME SUTURE MATERIAL USING MUCH SLOWER PASS THROUGH AND PULLING OF SUTURE. IT TOOK LONGER TO CLOSE INCISION DUE TO EXTRA PRECAUTION USED TO KEEP NEEDLE FROM DETACHING. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED LOT DEVICE AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL EVALUATION SUMMARY: REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITIONS OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF SUTURE NEEDLE WERE FOUND, AND THE TESTED SAMPLE MET THE FINISHED GOODS REQUIREMENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: EVENT DATE? ON (B)(6) 2019. PROCEDURE NAME: HARDWARE REMOVAL TENDON TRANSFER. QUANTITY INVOLVED? 1 . WHEN DID THE ISSUE OCCUR? NEEDLE DETACHED DURING SUTURING. HOW WAS CASE COMPLETED? COMPLETED WITH THIS SAME SUTURE MATERIAL USING MUCH SLOWER PASS THROUGH AND PULLING OF SUTURE. ANY ADVERSE PATIENT CONSEQUENCES? NO. TOOK LONGER TO CLOSE INCISION DUE TO EXTRA PRECAUTION USED TO KEEP NEEDLE FROM DETACHING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8456494&pc=GAM | 2210968-2019-79725 | 2019-02-27 | Malfunction | ETHICON INC. | 2019-03-27 | GAM | VICRYL PLUS SUTURE 27"(70CM) 0 VIO | Break; Degraded; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN OPEN HYSTERECTOMY SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE BREAKAGE SUTURE. AN EMPTY OPENED FOIL, TWO SUTURE PIECES AND A WINDING FORMER WITH TWO RE-PARKED USED NEEDLES AND SIX NEEDLE-SUTURE COMBINATIONS WERE RETURNED FOR ANALYSIS.. THE PRODUCT CODE IS CONTROL RELEASE AND REQUIRED EIGHT STRANDS PER PACKET. DURING THE VISUAL INSPECTION OF TWO SUTURE PIECES, BODY FLUIDS WERE OBSERVED ALONG OF STRANDS AND THE ENDS OF SUTURE PIECE WERE BUSTED DUE TO THE SUTURES HAVE BEGUN WITH THE DEGRADATION PROCESS AND NO FUNCTIONAL TEST COULD BE PERFORMED. SIX NEEDLE-SUTURES WERE VISUALLY INSPECTED, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS; HOWEVER, THE SUTURES HAVE BEGUN WITH DEGRADATION PROCESS AND NO FUNCTIONAL TEST COULD BE PERFORMED. THE FOIL WAS EXAMINED AND SHOWED MULTIPLES WRINKLES AND HOLES IN CAVITY ON BOTTOM FOIL PACKAGE DUE TO EXCESSIVE MANIPULATION. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE, THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE, IS SUTURE DEGRADATION; DUE TO THE IMPROPER HANDLING OF PACKAGE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8457922&pc=GAM | 2210968-2019-79739 | 2019-02-26 | Malfunction | ETHICON INC. | 2019-03-27 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT HYSTERECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING SEWING. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WILL THE PRODUCT BE RETURNED FOR EVALUATION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8457999&pc=GAM | 2210968-2019-79746 | 2019-03-04 | Malfunction | ETHICON INC. | 2019-03-27 | GAM | VICRYL SUTURE 27"(70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TRANSABDOMINAL PREPERITONEAL HERNIA PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE WAS BROKEN DURING CONTINUOUS SUTURING ON THE PERITONEUM. THE NEEDLE FRAGMENT DROPPED ON THE PERITONEUM AND WAS REMOVED WITH FORCEPS DURING THE PROCEDURE. THE PATIENT'S CONDITION IS STABLE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8458733&pc=GAM | 2210968-2019-79753 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-03-27 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. DURING THE CLOSURE THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION SUMMARY: AN OPENED BOX WITH THIRTY-TWO UNOPENED SAMPLES OF PRODUCT CODE VCP519, LOT MMBBQBS0 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THIRTEEN SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 3 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-79753, 2210968-2019-79754 AND 2210968-2019-79755. ANALYSIS RESULTS HAVE BEEN INCLUDED IN FOLLOW-UP REPORTS FOR EACH. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8458737&pc=GAM | 2210968-2019-79754 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-03-27 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. DURING THE CLOSURE THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8458744&pc=GAM | 2210968-2019-79755 | 2019-02-01 | Malfunction | ETHICON INC. | 2019-03-27 | GAM | VICRYL PLUS SUTURE 36"(90CM) 1 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE AND SUTURE WAS USED. DURING THE CLOSURE THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8451570&pc=GAM | 2210968-2019-79658 | 2007-01-01 | Malfunction | ETHICON INC. | 2019-03-26 | GAM | VICRYL SUTURE 27"(70CM) 5-0 UND | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT LAPAROSCOPIC PYLEOPLASTY PROCEDURE IN 2007 AND SUTURE WAS USED. DURING THE PROCEDURE, A NEEDLE BROKE AND IT WAS LEFT IN THE PATIENT. IT WAS REPORTED THAT AFTER COMPLETING THE PYELOPLASTY ANASTOMOSIS AND DURING THE NEEDLE RETRIEVAL, THE NEEDLE BOUNCED OFF THE DRIVER AND WAS LOST. IT WAS REPORTED THAT THE NEEDLE PIECE WAS RETAINED IN THE RENAL FOSSA CLOSE TO THE LEFT KIDNEY. PATIENT HAS BEEN SCHEDULED FOR AN MRI, NOT TO LOCATE OR RETRIEVE A NEEDLE, FOR AN UNKNOWN REASON. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THIS PATIENT EVENT FROM 2007 PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE COMPLAINT FILE NUMBER/ AFFILIATE NUMBER? NO IT WAS NOT REPORTED. IF NOT, PROVIDE FOLLOWING: ADDITIONAL COMPLAINT DETAILS: PROCEDURE NAME? PYLEOPLASTY PROCEDURE, PRODUCT CODE AND LOT NUMBER INVOLVED? UNKNOWN. WOULD YOU BE ABLE TO PROVIDE AN EVENT DESCRIPTION? AFTER COMPLETING THE PYELOPLASTY ANASTOMOSIS AND DURING THE NEEDLE RETRIEVAL, THE NEEDLE BOUNCED OFF THE DRIVER AND WAS LOST. QTY INVOLVED? ONLY ONE NEEDLE. WHAT TISSUE WAS THE NEEDLE PIECE LEFT IN PATIENT BODY? IN THE RENAL FOSSA CLOSE TO THE LEFT KIDNEY. WAS THERE ANY ADVERSE PATIENT OUTCOME? NO. WAS PROCEDURE COMPLETED SUCCESSFULLY AND HOW? YES, LAPAROSCOPICALLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8452462&pc=GAM | 2210968-2019-79670 |  | Malfunction | ETHICON INC. | 2019-03-26 | GAM | VICRYL VIO 70CM M3 USP2/0 SGLE ARMED CT-1 PLUS | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE SEPARATED FROM SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8453289&pc=GAM | 2210968-2019-79679 | 2019-03-13 | Malfunction | ETHICON INC. | 2019-03-26 | GAM | VICRYL*RAPIDE*UND 3-0 75CM V-4 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DENTAL EXTRACTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF. THE NEEDLE FELL INTO THE MOUTH AND WAS RETRIEVED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PROCEDURE NAME? DENTAL EXTRACTION. DID THE NEEDLE FALL INTO PATIENT? YES, IN THE MOUTH. IF YES, WAS IT REMOVED? YES. AND HOW? UNK. WAS PROCEDURE COMPLETED SUCCESSFULLY? AND HOW? UNK. ANY PATIENT CONSEQUENCES / AE OUTCOME AS A RESULT OF EVENT REPORT? NO PATIENT CONSEQUENCE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8428967&pc=GAM | 2210968-2019-79545 | 2019-01-30 | Malfunction | ETHICON INC. | 2019-03-18 | GAM | VICRYL PLUS SUTURE 27"(70CM) 4-0 VIO | Break | No Consequences Or Impact To Patient; Not Applicable | K032420 |  | 1 | Event Description: THIS CASE IS FROM HEALTH AUTHORITY. IT WAS REPORTED PATIENT ACCEPTED TONSILLECTOMY ON 30TH JAN. THE NEEDLE BROKE IN PATIENT DURING PROCEDURE. PROFESSIONAL WAS INVITED TO REMOVE BROKEN PIECE. FINALLY PROCEDURE WAS COMPLETED SMOOTHLY. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MK2985, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TISSUE WAS THE VICRYL PLUS SUTURE USED ON? WHAT TISSUE WAS THE NEEDLE PIECE RETAINED IN? WAS THE ENTIRE NEEDLE REMOVED DURING THE SAME PROCEDURE? DID THE SURGEON REMOVE THE NEEDLE PIECES DURING THE SAME PROCEDURE ON 30 JAN? WHAT IS THE SURGEON OPINION AS TO THE POTENTIAL CAUSE OF THE NEEDLE BREAKAGE? DO YOU HAVE THE STATUS OF SUTURE FOR EVALUATION? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8420478&pc=GAM | 2210968-2019-79469 | 2019-02-14 | Malfunction | ETHICON INC. | 2019-03-14 | GAM | VICRYL VIO 70CM M2::USP3/0 SGLE ARMED SH-1 PLUS | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AND UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THREADS CANNOT BE ABSORBED. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IS NOT AVAILABLE: ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THIS EVENT REPORT? WHICH TISSUE LAYER, AT WHICH LOCATION WAS THE SUTURE, WAS USED TO CLOSE? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION"? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? THERE IS A QTY INVOLVED AND TO BE RETURNED. DID 6 SUTURES NOT ABSORB? WHAT DOES 1/3 MEAN? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8420502&pc=GAM | 2210968-2019-79466 | 2019-02-14 | Malfunction | ETHICON INC. | 2019-03-14 | GAM | VICRYL VIO 70CM M2::USP3/0 SGLE ARMED SH-1 PLUS | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AND UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THREADS CANNOT BE ABSORBED. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ADDITIONAL EVALUATION SUMMARY: AN UNOPENED SAMPLE WAS RETURNED FOR ANALYSIS. THE REPRESENTATIVE SAMPLE OF PRODUCT CODE V311 WAS EXAMINED FOR VISUAL DEFECTS: APPEARANCE, COLOR PACKAGES WRINKLES IN THE SEAL AREA, CONTINUOUS SEALS, SEAL MARGINS, OVER-SEALING, DAMAGE (TORN, PUNCTURED). THE PACKET WAS OPENED AND DURING THE VISUAL INSPECTION THE NEEDLES/SUTURES WAS CORRECTLY PLACED ON THE WINDING FORMER. THE SUTURE WAS DISPENSED WITHOUT PROBLEM AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. ACCORDING TO THE SAMPLE CONDITION THE ASSIGNABLE CAUSE CANNOT BE DETERMINED SINCE NO DEFECTS WERE OBSERVED ON THE PACKET OR THE SUTURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IS NOT AVAILABLE: ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THIS EVENT REPORT? WHICH TISSUE LAYER, AT WHICH LOCATION WAS THE SUTURE, WAS USED TO CLOSE? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION"? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? THERE IS A QTY INVOLVED AND TO BE RETURNED. DID 6 SUTURES NOT ABSORB? WHAT DOES 1/3 MEAN? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8420503&pc=GAM | 2210968-2019-79468 | 2019-02-14 | Malfunction | ETHICON INC. | 2019-03-14 | GAM | VICRYL VIO 70CM M2::USP3/0 SGLE ARMED SH-1 PLUS | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AND UNKNOWN PROCEDURE ON (B)(4) 2019, AND SUTURE WAS USED. THREADS CANNOT BE ABSORBED. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ADDITIONAL EVALUATION SUMMARY: AN UNOPENED SAMPLE WAS RETURNED FOR ANALYSIS. THE REPRESENTATIVE SAMPLE OF PRODUCT CODE V311 WAS EXAMINED FOR VISUAL DEFECTS: APPEARANCE, COLOR PACKAGES ¿ WRINKLES IN THE SEAL AREA, CONTINUOUS SEALS, SEAL MARGINS, OVER-SEALING ¿ DAMAGE (TORN, PUNCTURED) . THE PACKET WAS OPENED AND DURING THE VISUAL INSPECTION THE NEEDLES/SUTURES WAS CORRECTLY PLACED ON THE WINDING FORMER. THE SUTURE WAS DISPENSED WITHOUT PROBLEM AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. ACCORDING TO THE SAMPLE CONDITION THE ASSIGNABLE CAUSE CANNOT BE DETERMINED SINCE NO DEFECTS WERE OBSERVED ON THE PACKET OR THE SUTURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IS NOT AVAILABLE: ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THIS EVENT REPORT? WHICH TISSUE LAYER, AT WHICH LOCATION WAS THE SUTURE, WAS USED TO CLOSE? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION"? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? THERE IS A QTY INVOLVED AND TO BE RETURNED. DID 6 SUTURES NOT ABSORB? WHAT DOES 1/3 MEAN? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8420504&pc=GAM | 2210968-2019-79467 | 2019-02-14 | Malfunction | ETHICON INC. | 2019-03-14 | GAM | VICRYL VIO 70CM M2::USP3/0 SGLE ARMED SH-1 PLUS | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AND UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THREADS CANNOT BE ABSORBED. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IS NOT AVAILABLE: ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THIS EVENT REPORT? WHICH TISSUE LAYER, AT WHICH LOCATION WAS THE SUTURE, WAS USED TO CLOSE? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION"? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? THERE IS A QTY INVOLVED AND TO BE RETURNED. DID 6 SUTURES NOT ABSORB? WHAT DOES 1/3 MEAN? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8420507&pc=GAM | 2210968-2019-79465 | 2019-02-14 | Malfunction | ETHICON INC. | 2019-03-14 | GAM | VICRYL VIO 70CM M2::USP3/0 SGLE ARMED SH-1 PLUS | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AND UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THREADS CANNOT BE ABSORBED. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IS NOT AVAILABLE: ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THIS EVENT REPORT? WHICH TISSUE LAYER, AT WHICH LOCATION WAS THE SUTURE, WAS USED TO CLOSE? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION"? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? THERE IS A QTY INVOLVED AND TO BE RETURNED. DID 6 SUTURES NOT ABSORB? WHAT DOES 1/3 MEAN? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8421882&pc=GAM | 2210968-2019-79442 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-03-14 | GAM | VICRYL SUTURE 18"(45CM) 3-0 UND | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGICAL PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE PHYSICIAN REPORTED THAT DURING THE PROCEDURE IN THE OFFICE, THE DEVICE DID NOT FEEL LIKE MULITPACK NEEDLES. ONLY ONE NEEDLE OUT OF THE ENTIRE BOX FELT LIKE THE RIGHT ONE ACCORDING TO THE GENERAL SURGEON. IT WAS ALSO REPORTED THERE WAS AN ISSUE WITH SUTURE BREAKAGE WITH THE SUTURE MATERIAL AND NEEDLES POPPING OFF. THE SURGEON COMMENTED THAT NEEDLES FELT MORE DULL COMPARED TO USUAL MULTI-PASS COATED NEEDLES. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). FOUR UNOPENED SAMPLES OF PRODUCT CODE J497, LOT # MJK059 WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE AND TENSILE STRENGTH WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLE RECEIVED, NO PULL OFF SUTURE NEEDLE, BEAKAGE SUTURE AND SURFACE RELATED DEFECT - NEEDLE WERE FOUND, AND THE TESTED SAMPLE MET THE FINISHED GOODS REQUIREMENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY SUTURE BROKE, PULLED OFF DURING THIS ONE PATIENT PROCEDURE?: ALL EXCEPT ONE. HOW MANY WERE NOTED TO BE DULL DURING THIS ONE PATIENT PROCEDURE?: NONE NECESSARY DULL, JUST DIDN¿T FEEL SHARP. CUSTOMER ADVISED THAT ISSUED OCCURRED WITH 2 OR 3 PACKS, NOT BOXES. NOTE: MEDWATCH FOR ADDITIONAL DEVICES SUBMITTED VIA 2210968-2019-80199, 2210968-2019-80200. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJK059 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT DO YOU MEAN BY "DON'T FEEL LIKE MULITPACK NEEDLES. ONLY ONE NEEDLE OUT OF THE ENTIRE BOX FELT LIKE THE RIGHT ONE ":? -SURGEON SAID NEEDLE QUALITY WAS ONLY NORMAL WITH ONE OF THE SUTURES. WAS THERE INCORRECT QTY OF NEEDLE-SUTURES IN THE PACKAGE? - NO 12 SUTURE IN PACKAGE. WERE ALL THE OTHER NEEDLE COMPONENTS FOUND TO BE INCORRECT OTHER THAN THE ONE? - YES ISSUES WITH BREAKAGE. IF YES, HOW SO? - BREAKAGE ISSUE WITH THE SUTURE MATERIAL + NEEDLES POPPING OFF, NEEDLES FELT MORE DULL COMPARED TO USUAL MULTI-PASS COATED NEEDLES PER SURGEON COMMENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW MANY SUTURE BROKE, PULLED OFF DURING THIS ONE PATIENT PROCEDURE? HOW MANY WERE NOTED TO BE DULL DURING THIS ONE PATIENT PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8416166&pc=GAM | 2210968-2019-79430 | 2019-02-14 | Malfunction | ETHICON INC. | 2019-03-13 | GAM | VICRYL VIO 70CM M2::USP3/0 SGLE ARMED SH-1 PLUS | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AND UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THREADS CANNOT BE ABSORBED. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE FOLLOWING ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IS NOT AVAILABLE: ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THIS EVENT REPORT? WHICH TISSUE LAYER, AT WHICH LOCATION WAS THE SUTURE, WAS USED TO CLOSE? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION"? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? THERE IS A QTY INVOLVED AND TO BE RETURNED. DID 6 SUTURES NOT ABSORB? WHAT DOES 1/3 MEAN? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8412824&pc=GEA | 2210968-2019-79395 |  | Malfunction | ETHICON INC. | 2019-03-12 | GEA | VICRYL*VLT 0 45CM - ENDOLOOP | Break | No Consequences Or Impact To Patient | K925914 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC APPENDECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE LOOP SNAPPED BEHIND THE INTESTINES, WHICH COULD HAVE BEEN POTENTIALLY DAMAGING. THERE WAS NO DAMAGE TO THE PATIENT¿S BODY INTERNALLY AND NO ADVERSE PATIENT CONSEQUENCES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: LOT # AK9898. CONFIRM THE TOTAL QTY INVOLVED FOR THE REPORTED SINGLE PATIENT EVENT ? CONFIRMED. WHY IS A TOTAL OF 18 QTY BEING REPORTED? BECAUSE HOSPITAL SENDS THEM ALL BACK. IT IS STATED "VICRYL LOOP SNAPPED BEHIND THE INTESTINES, POTENTIALLY DAMAGING" WHAT DOES THIS MEAN? THERE COULD BE A RISK. WAS THERE ANY DAMAGE TO THE PATIENT BODY INTERNALLY? NO. ANY PATIENT CONSEQUENCES/ AE OUTCOME AS A RESULT OF THE EVENT REPORT? NO. WAS THE PROCEDURE COMPLETED SUCCESSFULLY? YES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8412825&pc=GEA | 2210968-2019-79396 |  | Malfunction | ETHICON INC. | 2019-03-12 | GEA | VICRYL*VLT 0 45CM - ENDOLOOP | Break | No Consequences Or Impact To Patient | K925914 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC APPENDECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE LOOP SNAPPED BEHIND THE INTESTINES, WHICH COULD HAVE BEEN POTENTIALLY DAMAGING. THERE WAS NO DAMAGE TO THE PATIENT¿S BODY INTERNALLY AND NO ADVERSE PATIENT CONSEQUENCES. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: LOT # AK9898. CONFIRM THE TOTAL QTY INVOLVED FOR THE REPORTED SINGLE PATIENT EVENT ? CONFIRMED WHY IS A TOTAL OF (B)(4) QTY BEING REPORTED? BECAUSE HOSPITAL SENDS THEM ALL BACK IT IS STATED "VICRYL LOOP SNAPPED BEHIND THE INTESTINES, POTENTIALLY DAMAGING" WHAT DOES THIS MEAN? THERE COULD BE A RISK. WAS THERE ANY DAMAGE TO THE PATIENT BODY INTERNALLY? NO. ANY PATIENT CONSEQUENCES/ AE OUTCOME AS A RESULT OF THE EVENT REPORT? NO. WAS THE PROCEDURE COMPLETED SUCCESSFULLY? YES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8394912&pc=GAM | 2210968-2019-79305 |  | Malfunction | ETHICON INC. | 2019-03-06 | GAM | VICRYL SUTURE 18"(45CM) 1 VIO | Break; Material Puncture/Hole | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN OB-GYN PROCEDURE ON AN UNKNOWN DATE AND A SUTURE WAS USED. DURING SURGERY, IT WAS FOUND THAT THE SUTURE HAD BEEN BROKEN INTO PIECES WHEN OPENING THE PACKAGE. UPON EVALUATION OF THE RETURNED SAMPLE IT WAS FOUND THAT MULTIPLES WRINKLES AND HOLES WERE IN CAVITY OF FOIL PACKAGE. THE SUTURE HAD BEGUN THE DEGRADATION PROCESS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. EVALUATION: AN EMPTY OPENED FOIL, TWO DETACHED NEEDLES, SIX PARKED NEEDLE/ SUTURE AND SEVERAL SUTURE PIECES OF PRODUCT WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, THE SWAGE AND ATTACHMENT AREA OF THE EIGHT NEEDLES WERE AS EXPECTED. THE SUTURES COULD NOT BE DISPENSED SINCE; IT HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED. ALSO, THE FOIL WAS EXAMINED FOR VISUAL INSPECTION AND SHOWED MULTIPLES WRINKLES AND HOLES IN CAVITY ON THE TOP AND BOTTOM FOIL PACKAGE (FROM OUTSIDE TO INSIDE) DUE TO EXCESSIVE MANIPULATION. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF SUTURE BREAKAGE PRE-OP IS A SUTURE DEGRADATION; WHICH IS A RESULT OF IMPROPER HANDLING OF PACKAGE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8371967&pc=GAM | 2210968-2019-79083 | 2019-01-16 | Malfunction | ETHICON INC. | 2019-02-26 | GAM | VICRYL PLUS SUTURE 36"(90CM) 2-0 VIO | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND A SUTURE WAS USED. DURING SURGERY, THE SUTURE BROKE FIVE TIMES. A REPLACEMENT DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8311982&pc=GAM | MW5083786 | 2017-02-07 | Malfunction | ETHICON, LLC. | 2019-02-05 | GAM | VICRYL RAPIDE SUTURE | Adverse Event Without Identified Device or Use Problem | Purulent Discharge; Pain |  |  | 1 | Event Description: VICRYL SUTURE WAS USED BY MY DERMATOLOGIST AFTER A WIDE EXCISION FOR BASAL CELL ON MY LEFT UPPER ARM ON (B)(6) 2017. ON (B)(6) 2017 THE INCISION SPONTANEOUSLY SPLIT OPEN AND OOZED BLACK-RED PUS. AFTER 2 YEARS AND 2 COURSES OF ANTIBIOTICS AND SURFACE SCAR BIOPSY, THE ENTIRE LENGTH OF THE ORIGINAL INCISION IS PAINFUL AND THE DELTOID MUSCLE IS SORE DIRECTLY UNDER THE SCAR.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8289417&pc=GAM | 2210968-2019-78625 | 2018-09-04 | Malfunction | ETHICON, INC. | 2019-01-30 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K022269 |  | 458 | Event Description: THIS REPORT SUMMARIZES <NOE> 458 </NOE> INITIAL MALFUNCTION EVENTS, INCLUDING 188 SUTURE BREAKAGE, 81 NEEDLE BREAKAGE AND 189 NEEDLE PULL-OFF THAT OCCURRED INTRAOPERATIVELY. IT ALSO SUMMARIZES 45 FOLLOW UP MALFUNCTION EVENTS, INCLUDING 34 SUTURE BREAKAGE, 8 NEEDLE BREAKAGE AND 3 NEEDLE PULL-OFF THAT OCCURRED INTRAOPERATIVELY. Manufacturer Narrative: SUTURE BREAKAGE MALFUNCTION EVENTS: 222. SUTURE BREAKAGE DEVICES RETURNED: 385. MANUFACTURING DEFICIENCY: 12. NEEDLE BREAKAGE MALFUNCTION EVENTS: 89. NEEDLE BREAKAGE DEVICES RETURNED: 155. MANUFACTURING DEFICIENCY - 8. NEEDLE PULL OFF MALFUNCTION EVENTS: 192. NEEDLE PULL OFF DEVICES RETURNED: 297. MANUFACTURING DEFICIENCY - 6. ALL DEVICES LABELED FOR SINGLE USE; 0 DEVICES REPROCESSED AND RE-USED. 0 REMEDIAL ACTION. THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. COMPLAINT INFORMATION WILL BE INCLUDED IN COMPLAINT TRENDING THAT IS REVIEWED ON A REGULAR BASIS TO DETERMINE IF FURTHER ACTION IS NECESSARY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8249744&pc=GAM | 2210968-2019-78421 | 2018-12-18 | Malfunction | ETHICON INC. | 2019-01-15 | GAM | VICRYL SUTURE 27"(70CM) 5-0 UND | Component Incompatible | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UROLOGICAL PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. IT WAS FOUND THAT WHEN OPENING THE PACKAGE, THERE HAD BEEN A DOUBLE-ARMED SUTURE THOUGH THE CODE IS A SINGLE-ARMED SUTURE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES FOR THE PATIENT. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8237855&pc=GAM | 2210968-2019-78328 |  | Malfunction | ETHICON INC. | 2019-01-10 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8231573&pc=GAM | 2210968-2019-78289 | 2018-12-06 | Malfunction | ETHICON INC. | 2019-01-09 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Break; Component Incompatible; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. IT WAS NOTED TO HAVE POOR TENSILE QUALITY AND A 2-0 SUTURE FOIL IN INCORRECT PACKAGE. NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY WHAT IS MEANT BY "2-0 SUTURE FOIL IN VP 2421 PACKAGING" 2-0 SIZE OF SUTURE WAS FOUND IN ONE OF THE FOILS OF VP2421 WHICH IS A NUMBER 1. WHAT IS THE LOT NUMBER? LOT-T7021. PLEASE CONNECT WITH SALES REPRESENTATIVE AND ASK ADDITIONAL INFORMATION REGARDING PLEASE PROVIDE LOT NUMBERS OF BOTH CODES. VP2421-T7021, VP2382-LOT NUMBER NOT AVAILABLE ON PRODUCT SLIP. WAS THE EVENT IDENTIFIED AFTER OPENING OF SEALED (DOZEN) CECO BOX./OR THE BOX WAS PROVIDED IN OPEN CONDITION BY PHARMACY? IDENTIFIED AFTER A FOIL WAS OPENED. WAS FOILS OF VP2382 OBSERVED IN THE CECO BOX OF VP2421? THE SUTURE OF SPECIFICATION VP2382 WAS FOUND INSIDE THE FOIL OF VP2421. HOW MANY FOILS/SUTURES OBSERVED IN VP2421? ONE FOIL. WAS THE PRIMARY PACKING (ZIPPER TRAYS) OBSERVED AFTER OPENING OF FOILS? ZIPPER TRAY WAS OF VP2382 INSIDE THE FOIL OF VP2421. HOW MANY SUTURES GOT AFFECTED/BROKEN? ONE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8226311&pc=GAM | 2210968-2019-78253 |  | Malfunction | ETHICON INC. | 2019-01-08 | GAM | VICRYL RAP UND 45CM 4/0 PS-2 PRIME MP | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORAL FRENECTOMY PROCEDURE IN 2018 AND A SUTURE WAS USED. THE SURGEON FOUND THAT THE SUTURE APPEARED TO BE COMPLETELY ABSORBED ONLY FIVE DAYS AFTER THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY ANTIBIOTICS-MEDICATION /TREATMENT OR MEDICAL/RE-SUTURING/SURGICAL INTERVENTION PROVIDED TO THE PATIENT POST-OP AS A RESULT OF THE EVENT REPORTED? NO ANTIBIOTICS / TREATMENT OR SURGICAL INTERVENTION WAS REQUIRED. IT APPEARED THE PATIENTS¿ TISSUE HAD ACHIEVED ADEQUATE SELF SUPPORT. PATIENT CURRENT CONDITION IF KNOWN? PATIENT IS WELL ¿ NO COMPLICATIONS PATIENT¿S INFORMATION SUCH AS, DOES THE PATIENT HAVE ANY OTHER CO-MORBIDITIES? LIFESTYLE E.G. ALCOHOL INTAKE? NO UNUSUAL PATIENT FACTORS / CO MORBIDITIES / LIFESTYLE CHOICES TO NOTE. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8226314&pc=GAM | 2210968-2019-78256 |  | Malfunction | ETHICON INC. | 2019-01-08 | GAM | VICRYL SUTURE 18"(45CM) 4-0 UND | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORAL FRENECTOMY PROCEDURE IN 2018 AND A SUTURE WAS USED. THE SURGEON FOUND THAT THE SUTURE APPEARED TO BE COMPLETELY ABSORBED ONLY FIVE DAYS AFTER THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY ANTIBIOTICS-MEDICATION /TREATMENT OR MEDICAL/RE-SUTURING/SURGICAL INTERVENTION PROVIDED TO THE PATIENT POST-OP AS A RESULT OF THE EVENT REPORTED? NO ANTIBIOTICS / TREATMENT OR SURGICAL INTERVENTION WAS REQUIRED. IT APPEARED THE PATIENTS¿ TISSUE HAD ACHIEVED ADEQUATE SELF SUPPORT. PATIENT CURRENT CONDITION IF KNOWN? PATIENT IS WELL ¿ NO COMPLICATIONS PATIENT¿S INFORMATION SUCH AS, DOES THE PATIENT HAVE ANY OTHER CO-MORBIDITIES? LIFESTYLE E.G. ALCOHOL INTAKE? NO UNUSUAL PATIENT FACTORS / CO MORBIDITIES / LIFESTYLE CHOICES TO NOTE. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8212202&pc=GAM | 2210968-2019-78133 |  | Malfunction | ETHICON INC. | 2019-01-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2018 AND A SUTURE WAS USED. PRIOR TO SURGERY, IT WAS FOUND THAT THE PACKAGING WAS PERFORATED BY THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WAS MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8215869&pc=GAM | 2210968-2019-78171 | 2018-11-01 | Malfunction | ETHICON INC. | 2019-01-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 UND | Appropriate Term/Code Not Available | Local Reaction | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN (B)(6) 2018 AND SUTURE WAS USED. POST-OPERATIVELY, THE PATIENT DEVELOPED REDNESS, AND THE SUTURE DID NOT ABSORB. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME. TISSUE WHERE SUTURE WAS USED. PROCEDURE DATE. EVENT DATE. HOW MAY DAYS POST-OP DID THE PATIENT RETURNED WITH THE ISSUE? LOCATION AND SIZE OF REACTION. ANY PHOTOS SHOWING THE REACTION? WAS ANY DEFICIENCY/NON-CONFORMANCE NOTED WITH THE DEVICE BEFORE OR USE OR DURING POST-OP EXAMINATION? IF YES, PLEASE DESCRIBE. WAS THE DEVICE SUBJECTED TO ANY OTHER TREATMENT BEFORE USE? WAS ANTIBIOTICS PROVIDED PROPHYLACTICALLY TO THE PATIENT WHO IS ALLERGIC TO ANTIBIOTICS, WHICH MAY HAVE CAUSED ADVERSE REACTION? KNOWN ALLERGIES OR ANY ALLERGY TESTING DONE ON THE OTHER TWO PATIENTS? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THE REACTION? PLEASE INCLUDE MEDICINE NAME, STRENGTH AND DOSAGE, IF PRESCRIBED. WERE THE SUTURES REMOVED? PATIENT PRESENT CONDITION. LOT NUMBER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8215872&pc=GAM | 2210968-2019-78173 | 2018-11-01 | Malfunction | ETHICON INC. | 2019-01-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 UND | Appropriate Term/Code Not Available | Reaction | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN (B)(6) 2018 AND SUTURE WAS USED. POST-OPERATIVELY, THE PATIENT DEVELOPED REDNESS, AND THE SUTURE DID NOT ABSORB. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME? TISSUE WHERE SUTURE WAS USED? PROCEDURE DATE? EVENT DATE? HOW MAY DAYS POST-OP DID THE PATIENT RETURNED WITH THE ISSUE? LOCATION AND SIZE OF REACTION ANY PHOTOS SHOWING THE REACTION? WAS ANY DEFICIENCY/NON-CONFORMANCE NOTED WITH THE DEVICE BEFORE OR USE OR DURING POST-OP EXAMINATION? IF YES, PLEASE DESCRIBE WAS THE DEVICE SUBJECTED TO ANY OTHER TREATMENT BEFORE USE? WAS ANTIBIOTICS PROVIDED PROPHYLACTICALLY TO THE PATIENT WHO IS ALLERGIC TO ANTIBIOTICS, WHICH MAY HAVE CAUSED ADVERSE REACTION? KNOWN ALLERGIES OR ANY ALLERGY TESTING DONE ON THE OTHER TWO PATIENTS? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THE REACTION? PLEASE INCLUDE MEDICINE NAME, STRENGTH AND DOSAGE, IF PRESCRIBED. WERE THE SUTURES REMOVED? PATIENT PRESENT CONDITION? LOT NUMBER?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8215880&pc=GAM | 2210968-2019-78174 | 2018-11-01 | Malfunction | ETHICON INC. | 2019-01-03 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 UND | Appropriate Term/Code Not Available | Local Reaction | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN (B)(6) 2018 AND SUTURE WAS USED. POST-OPERATIVELY, THE PATIENT DEVELOPED REDNESS, AND THE SUTURE DID NOT ABSORB. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME. TISSUE WHERE SUTURE WAS USED. PROCEDURE DATE. EVENT DATE. HOW MAY DAYS POST-OP DID THE PATIENT RETURNED WITH THE ISSUE? LOCATION AND SIZE OF REACTION? ANY PHOTOS SHOWING THE REACTION? WAS ANY DEFICIENCY/NON-CONFORMANCE NOTED WITH THE DEVICE BEFORE OR USE OR DURING POST-OP EXAMINATION? IF YES, PLEASE DESCRIBE. WAS THE DEVICE SUBJECTED TO ANY OTHER TREATMENT BEFORE USE? WAS ANTIBIOTICS PROVIDED PROPHYLACTICALLY TO THE PATIENT WHO IS ALLERGIC TO ANTIBIOTICS, WHICH MAY HAVE CAUSED ADVERSE REACTION? KNOWN ALLERGIES OR ANY ALLERGY TESTING DONE ON THE OTHER TWO PATIENTS? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED TO MANAGE THE REACTION? PLEASE INCLUDE MEDICINE NAME, STRENGTH AND DOSAGE, IF PRESCRIBED. WERE THE SUTURES REMOVED? PATIENT PRESENT CONDITION? LOT NUMBER?
The MAUDE database houses medical device reports submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers. Each year | the FDA receives several hundred thousand medical device reports (MDRs) of suspected device-associated deaths | serious injuries and malfunctions. The FDA uses MDRs to monitor device performance | detect potential device-related safety issues | and contribute to benefit-risk assessments of these products. The MAUDE database houses MDRs submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers.
Although MDRs are a valuable source of information | this passive surveillance system has limitations | including the potential submission of incomplete | inaccurate | untimely | unverified | or biased data. In addition | the incidence or prevalence of an event cannot be determined from this reporting system alone due to potential under-reporting of events and lack of information about frequency of device use. Because of this | MDRs comprise only one of the FDA's several important postmarket surveillance data sources.
MDR data alone cannot be used to establish rates of events | evaluate a change in event rates over time or compare event rates between devices. The number of reports cannot be interpreted or used in isolation to reach conclusions about the existence | severity | or frequency of problems associated with devices. Confirming whether a device actually caused a specific event can be difficult based solely on information provided in a given report. Establishing a cause-and-effect relationship is especially difficult if circumstances surrounding the event have not been verified or if the device in question has not been directly evaluated.
MAUDE data does not represent all known safety information for a reported medical device and should be interpreted in the context of other available information when making device-related or treatment decisions. Variations in trade | product | and company names affect search results. Searches only retrieve records that contain the search term(s) provided by the requester.Submission of a medical device report and the FDA's release of that information is not necessarily an admission that a product | user facility | importer | distributor | manufacturer | or medical personnel caused or contributed to the event. Certain types of report information are protected from public disclosure under the Freedom of Information Act (FOIA).
If a report contains trade secret or confidential business information | that text is replaced by "(b)(4)". If a report contains personnel or medical files information | that text is replaced by "(b)(6)". The designations "(b)(4)" and "(b)(6)" refer to the exemptions in the FOIA. For example | "(b)(4)" may be found in place of the product's composition and "(b)(6)" may be found in place of a patient's age. MAUDE is updated monthly and the search page reflects the date of the most recent update. The FDA seeks to include all reports received prior to the update but the inclusion of some reports may be delayed.
Manufacturers and importers must submit reports when they become aware of information that reasonably suggests that one of their marketed devices may have caused or contributed to a death or serious injury or has malfunctioned and the malfunction of the device or a similar device that they market would be likely to cause or contribute to a death or serious injury if the malfunction were to recur. Manufacturers must send reports of such deaths | serious injuries and malfunctions to the FDA. Importers must send reports of deaths and serious injuries to the FDA and the manufacturer | and reports of malfunctions to the manufacturer.
Device user facilities include hospitals | outpatient diagnostic or treatment facilities | nursing homes and ambulatory surgical facilities. Device user facilities must submit reports when they become aware of information that reasonably suggests that a device may have caused or contributed to a death or serious injury of a patient in their facility. Death reports must be sent to the FDA and the manufacturer | if known. Serious injury reports must be sent to the manufacturer or to the FDA | if the manufacturer is not known.