## Sheet1
| Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9530187&pc=GAM | 2210968-2019-91380 | 2019-12-09 | Malfunction | ETHICON INC. | 2019-12-30 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE OFF THE NEEDLE WITHOUT APPLYING TOO MUCH TENSION. ANOTHER LIKE DEVICE FROM THE SAME LOT WAS USED AND IT WAS FINE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THE 3 SUTURES FAIL IN THE SINGLE PROCEDURE? DID THE NEEDLE PULL OFF THE SUTURE THREAD OR THE NEEDLE BROKE? DID THE EVENT OCCUR DURING SUTURING (USE ON PATIENT)? ARE ACTUAL SAMPLES BEING RETURNED OR REPRESENTATIVE? DEVICE RETURN STATUS/ FOLLOW UP NOTE: EVENTS REPORTED IN 2210968-2019-91381, AND 2210968-2019-91382.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9530190&pc=GAM | 2210968-2019-91381 | 2019-12-09 | Malfunction | ETHICON INC. | 2019-12-30 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE OFF THE NEEDLE WITHOUT APPLYING TOO MUCH TENSION. ANOTHER LIKE DEVICE FROM THE SAME LOT WAS USED AND IT WAS FINE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4) ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THE 3 SUTURES FAIL IN THE SINGLE PROCEDURE? DID THE NEEDLE PULL OFF THE SUTURE THREAD OR THE NEEDLE BROKE? DID THE EVENT OCCUR DURING SUTURING (USE ON PATIENT)? ARE ACTUAL SAMPLES BEING RETURNED OR REPRESENTATIVE? DEVICE RETURN STATUS/ FOLLOW UP. NOTE: EVENTS REPORTED IN 2210968-2019-91380, 2210968-2019-91381, AND 2210968-2019-91382.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9530192&pc=GAM | 2210968-2019-91382 | 2019-12-09 | Malfunction | ETHICON INC. | 2019-12-30 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE OFF THE NEEDLE WITHOUT APPLYING TOO MUCH TENSION. ANOTHER LIKE DEVICE FROM THE SAME LOT WAS USED AND IT WAS FINE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THE 3 SUTURES FAIL IN THE SINGLE PROCEDURE? DID THE NEEDLE PULL OFF THE SUTURE THREAD OR THE NEEDLE BROKE? DID THE EVENT OCCUR DURING SUTURING (USE ON PATIENT)? ARE ACTUAL SAMPLES BEING RETURNED OR REPRESENTATIVE? DEVICE RETURN STATUS/ FOLLOW UP. NOTE: EVENTS REPORTED IN 2210968-2019-91380, 2210968-2019-91381, AND 2210968-2019-91382.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9531407&pc=GAM | 2210968-2019-91397 | 2019-10-15 | Malfunction | ETHICON INC. | 2019-12-30 | GAM | VICRYL CT UD 45CM M1.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE IN TWO PARTS WHEN PERFORMING THE SUTURE OF THE LAST STITCH. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V494H; PDBDCST0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE NEEDLE BREAKAGE. A FAILED SUTURE NEEDLES WAS SUBMITTED FOR FRACTOGRAPHIC EVALUATION. A FRACTURE WAS OBSERVED AT THE SUTURE ATTACHMENT OF SAMPLE B. ONE SIDE OF SAMPLE B WAS RECEIVED, THE MATING FRACTURE SURFACE WAS STILL ATTACHED TO THE BODY OF THE SPECIMEN. A MICROSCOPE WAS USED TO EXAMINE THE FRACTURE SURFACE AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACE WAS EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE ATTACHMENT AREA OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9521574&pc=GAM | 2210968-2019-91332 | 2019-01-04 | Malfunction | ETHICON INC. | 2019-12-26 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE CAME OFF THE STRAND. IT HAPPENED TO SEVERAL SUTURES FROM THE SAME BOX. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CONFIRM THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY ACTUALLY INVOLVED WITH REPORTED ISSUE) THAT FAILED DURING THIS REPORTED EVENT / PROCEDURE? 3. THE FOLLOWING INFORMATION WAS REQUESTED BUT UNAVAILABLE: LOT #? NOTE: THE EVENTS HAVE BEEN REPORTED IN MW # 2210968-2019-91331, 2210968-2019-91332, 2210968-2019-91333.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9521575&pc=GAM | 2210968-2019-91333 | 2019-01-04 | Malfunction | ETHICON INC. | 2019-12-26 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE CAME OFF THE STRAND. IT HAPPENED TO SEVERAL SUTURES FROM THE SAME BOX. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CONFIRM THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY ACTUALLY INVOLVED WITH REPORTED ISSUE) THAT FAILED DURING THIS REPORTED EVENT / PROCEDURE? 3 THE FOLLOWING INFORMATION WAS REQUESTED BUT UNAVAILABLE: LOT #? NOTE: THE EVENTS HAVE BEEN REPORTED IN MW # 2210968-2019-91331, 2210968-2019-91332, 2210968-2019-91333.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9521579&pc=GAM | 2210968-2019-91331 | 2019-01-04 | Malfunction | ETHICON INC. | 2019-12-26 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE CAME OFF THE STRAND. IT HAPPENED TO SEVERAL SUTURES FROM THE SAME BOX. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: CONFIRM THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY ACTUALLY INVOLVED WITH REPORTED ISSUE) THAT FAILED DURING THIS REPORTED EVENT / PROCEDURE? 3. THE FOLLOWING INFORMATION WAS REQUESTED BUT UNAVAILABLE: LOT #? NOTE: THE EVENTS HAVE BEEN REPORTED IN MW # 2210968-2019-91331, 2210968-2019-91332, 2210968-2019-91333.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516373&pc=GAM | 2210968-2019-91276 | 2019-12-05 | Malfunction | ETHICON INC. | 2019-12-24 | GAM | VICRYL RAP UND 70CM 3/0 PS-1 PRIME MP | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN OPEN APPENDECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE AND SWAGE CAME APART. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PDBDRRP0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516375&pc=GAM | 2210968-2019-91275 | 2019-12-05 | Malfunction | ETHICON INC. | 2019-12-24 | GAM | VICRYL RAP UND 70CM 3/0 PS-1 PRIME MP | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN OPEN APPENDECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE AND SWAGE CAME APART. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PDBDRRP0, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9501809&pc=GAM | 2210968-2019-91145 | 2019-11-24 | Malfunction | ETHICON INC. | 2019-12-20 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE PROBLEM OF PULLING OFF SUTURE NEEDLE HAPPENED WHEN SEWING DURING PROCEDURE. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9501850&pc=GAM | 2210968-2019-91146 | 2019-11-12 | Malfunction | ETHICON INC. | 2019-12-20 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A NATURAL BIRTH PROCEDURE ON (B)(6) AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE PROBLEM OF PULLING OFF SUTURE NEEDLE HAPPENED WHEN SEWING PERINEAL MUCOSA. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9505688&pc=GAM | 2210968-2019-91182 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-12-20 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND SUTURE WAS USED. DURING THE SUTURES, THE STRANDS PULLED OFF THE NEEDLE. NO CHANGE REGARDING THE OPERATIVE ROOM. NO PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV453; PEBDZKR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL ANIMAL EVENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9506862&pc=GAM | 2210968-2019-91201 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-12-20 | GAM | VICRYL CT VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT A HYSTERECTOMY PROCEDURE IN 2019 AND SUTURE WAS USED. DURING THE LIGATION OF THE OVARIAN PEDICLE THE THREAD BREAKS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9506938&pc=GAM | 2210968-2019-91202 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-12-20 | GAM | VICRYL VIO 75CM M2 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT A HYSTERECTOMY PROCEDURE IN 2019 AND SUTURE WAS USED. DURING THE LIGATION OF THE OVARIAN PEDICLE THE THREAD BREAKS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DEVICE RETURN STATUS/FOLLOW UP? PGVCGVRO / PGVCGVR0 INVALID LOTS. CAN YOU CLARIFY THE LOT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9500024&pc=GAM | 2210968-2019-91131 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-12-19 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT A HYSTERECTOMY PROCEDURE IN 2019 AND SUTURE WAS USED. DURING THE LIGATION OF THE OVARIAN PEDICLE THE THREAD BREAKS. A DIFFERENT SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DEVICE RETURN STATUS/FOLLOW UP? CHABCHGRO / CHABCHGR0 INVALID LOTS. CAN YOU PROVIDE THE CORRECT LOT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9485548&pc=GAM | 2210968-2019-91034 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-12-18 | GAM | VICRYL CT UD 70CM M3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING USE THE NEEDLE PULLED OFF. IT WAS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). REPRESENTATIVE SAMPLES WERE RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLE RECEIVED, NO PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLE MET THE FINISHED GOODS REQUIREMENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9489982&pc=GAM | 2210968-2019-91074 | 2019-11-25 | Malfunction | ETHICON INC. | 2019-12-18 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE SURGERY, THE SUTURE BROKE OFF. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9474508&pc=GAM | 2210968-2019-90944 |  | Malfunction | ETHICON INC. | 2019-12-16 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE AND THE THREAD SEEMED TO COME LOOSE AND ENDED UP SEPARATED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY: ¿THE THREAD SEEMS TO BE COMING LOOSE. ¿ WHEN DID THE ISSUE OCCUR; IN THE PACKAGE/ DURING DISPENSING (BEFORE USE ON PATIENT/DURING ACTUAL SUTURING (USE ON PATIENT)? NAME OF THE PROCEDURE? DATE THE EVENT OCCURRED? DEVICE RETURN STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9475914&pc=GAM | 2210968-2019-90960 | 2019-11-01 | Malfunction | ETHICON INC. | 2019-12-16 | GAM | VICRYL RAP UND 75CM M1.5 USP4/0 SGLE ARMED PC-3 PRIME | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A HEMORRHOIDECTOMY PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. THE NEEDLE DISCONNECTED FROM THE SUTURE WHILE SUTURING DURING SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. *HOW MANY DEVICES FAILED DURING THIS SINGLE PROCEDURE? *DEVICE RETURN STATUS/FOLLOW UP?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9475927&pc=GAM | 2210968-2019-90961 | 2019-11-01 | Malfunction | ETHICON INC. | 2019-12-16 | GAM | VICRYL RAP UND 75CM M1.5 USP4/0 SGLE ARMED PC-3 PRIME | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A HEMORRHOIDECTOMY PROCEDURE ON UNKNOWN DATE AND THE SUTURE WAS USED. THE NEEDLE DISCONNECTED FROM THE SUTURE WHILE SUTURING DURING SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9470259&pc=GAM | 2210968-2019-90917 | 2019-11-22 | Malfunction | ETHICON INC. | 2019-12-13 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN APPENDECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE IN THE FASCIA OF A PATIENT USING IT ON A HISSAN DURING THE PROCEDURE. THE NEEDLE PIECES WERE REMOVED FROM THE PATIENT DURING THE PROCEDURE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9457362&pc=GAM | 2210968-2019-90733 | 2019-11-21 | Malfunction | ETHICON INC. | 2019-12-12 | GAM | VICRYL RAP UND 75CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A DENTAL SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE THREAD DURING SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A PHOTO OF THE DEVICE HAS BEEN RECEIVED. A SUPPLEMENTAL MEDWATCH WILL BE SENT WITH EVALUATION RESULTS. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9459353&pc=GAM | 2210968-2019-90809 | 2019-10-20 | Malfunction | ETHICON INC. | 2019-12-12 | GAM | VICRYL CT VIO 90CM M4 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN OB-GYN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. AT THE CLOSING OF THE UTERUS, THE NEEDLE PULLED OFF THE STRAND. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9461379&pc=GAM | 2210968-2019-90852 | 2019-11-13 | Malfunction | ETHICON INC. | 2019-12-12 | GAM | VICRYL VIO 75CM M1 | Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE GOT SPLIT AFTER SEWING THE SKIN FOR THE FIRST STITCH DURING THE PROCEDURE. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. NO ADDITIONAL INFORMATION COULD BE PROVIDED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT CODE W9442 LOT LL7BGCM, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9450740&pc=GAM | 2210968-2019-90745 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-12-11 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS NOTED THAT THE PACKAGING READ NOT FOR HUMAN USE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. UPON EVALUATION OF THE RETURNED BOX, IT WAS OBSERVED AN INAPPROPRIATE INFORMATION AT THE CARDBOARD. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NON CONFORMANCE, WHICH WAS RAISED TO ADDRESS THE EVENT REPORTED. THE NECESSARY ACTIONS HAVE BEEN TAKEN TO ADDRESS THE ISSUE. INVESTIGATION SUMMARY: IT WAS RECEIVED FOR ANALYSIS AN OPENED BOX WITH THREE UNOPENED SAMPLES OF PRODUCT CODE V372, LOT PHBBCJQ0. DURING THE VISUAL INSPECTION OF THE BOX, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES INSIDE THE BOX WERE REVIEWED AND ALL THE INFORMATION WAS ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9443968&pc=GAM | 2210968-2019-90694 | 2018-07-01 | Malfunction | ETHICON INC. | 2019-12-10 | GAM | VICRYL CT VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TAHBSO IN (B)(6) 2018 AND SUTURE WAS USED FOR CLOSURE OF THE ABDOMINAL WALL FASCIA. DURING THE PROCEDURE, THE NEEDLE BROKE AT THE NEEDLE TIP WITH NO TRACES OF MATERIAL LEFT INSIDE THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT LE8CLRB0, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE NEEDLE BREAKAGE. A FAILED SUTURE NEEDLE WAS SUBMITTED FOR FRACTOGRAPHIC EVALUATION. THE COMPONENTS WERE IDENTIFIED AS PRODUCT CODE W9431. A FRACTURE WAS OBSERVED AT THE TIP BODY OF SAMPLE B. ONE SIDE OF THE NEEDLE WAS RECEIVED, THE MATING FRACTURE SURFACE WAS NOT PROVIDED FOR THIS EVALUATION. A MICROSCOPE WAS USED TO EXAMINE THE FRACTURE SURFACE AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACE WAS EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE,WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. INTERGRANULAR FRACTURE WAS ALSO IDENTIFIED IN THIS REGION AND IN SMALL AREAS ACROSS THE FRACTURE SURFACE. MECHANICAL DAMAGE, A CUP-AND-CONE SHAPED FRACTURE AND MACROSCOPIC PLASTIC DEFORMATION OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDE ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED FROM MECHANICAL DEFORMATION, INDENTS AND SCRATCHES, DUE TO GRIPPING, BENDING AND OVERSTRESSING OF THE NEEDLE. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW THAT WOULD CAUSE PREMATURE FAILURE. DURING THE VISUAL INSPECTION OF TWO UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. NO NEEDLE BREAKAGE WAS OBSERVED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS WERE OBSERVED. ALSO, THE SAMPLES WERE TESTED BY RESISTANCE TEST TO NEEDLE AND MET THE REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLE RECEIVED, NO PERFORMANCE BREAKAGE NEEDLE WERE FOUND. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 4 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-90006, 2210968-2019-90007, AND 2210968-2019-90692. ANALYSIS RESULTS HAVE BEEN INCLUDED IN REPORTS FOR EACH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9443985&pc=GAM | 2210968-2019-90692 | 2018-07-01 | Malfunction | ETHICON INC. | 2019-12-10 | GAM | VICRYL CT VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TAHBSO IN (B)(6) 2018 AND SUTURE WAS USED FOR CLOSURE OF THE ABDOMINAL WALL FASCIA. DURING THE PROCEDURE, THE NEEDLE BROKE AT THE NEEDLE TIP WITH NO TRACES OF MATERIAL LEFT INSIDE THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT LE8CLRB0, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE NEEDLE BREAKAGE. A FAILED SUTURE NEEDLE WAS SUBMITTED FOR FRACTOGRAPHIC EVALUATION. THE COMPONENTS WERE IDENTIFIED AS PRODUCT CODE W9431. A FRACTURE WAS OBSERVED AT THE TIP OF SAMPLE A. ONE SIDE OF THE NEEDLE WAS RECEIVED, THE MATING FRACTURE SURFACE WAS NOT PROVIDED FOR THIS EVALUATION. A MICROSCOPE WAS USED TO EXAMINE THE FRACTURE SURFACE AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACE WAS EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. INTERGRANULAR FRACTURE WAS ALSO IDENTIFIED IN THIS REGION AND IN SMALL AREAS ACROSS THE FRACTURE SURFACE. MECHANICAL DAMAGE, A CUP-AND-CONE SHAPED FRACTURE AND MACROSCOPIC PLASTIC DEFORMATION OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDE ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE TIP OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. DURING THE VISUAL INSPECTION OF TWO UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. NO NEEDLE BREAKAGE WAS OBSERVED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS WERE OBSERVED. ALSO, THE SAMPLES WERE TESTED BY RESISTANCE TEST TO NEEDLE AND MET THE REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLE RECEIVED, NO PERFORMANCE BREAKAGE NEEDLE WERE FOUND. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 4 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-90006, 2210968-2019-90007, 2210968-2019-90692, AND 2210968-2019-90694. ANALYSIS RESULTS HAVE BEEN INCLUDED IN REPORTS FOR EACH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9444616&pc=GAM | 2210968-2019-90711 | 2019-10-03 | Malfunction | ETHICON INC. | 2019-12-10 | GAM | VICRYL RAP UND 75CM M1.5 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF SUTURE NEEDLE BROKE OFF WHEN SURGEON RELOADING IT. THE TIP OF NEEDLE WAS FOUND ON THE MAYO STAND AND WAS RECOVERED. THE NEEDLE WAS NOT IN THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT: AL9476 AND NO NON-CONFORMANCES WERE IDENTIFIED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CONFIRM EVENT/PROCEDURE DATE, RECEIVED DATE AND ALERT DATE FOR THIS COMPLAINT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9438193&pc=GAM | 2210968-2019-90655 | 2019-10-01 | Malfunction | ETHICON INC. | 2019-12-09 | GAM | VICRYL VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN OVARIECTOMY IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN LIGATING, THE DEVICE BREAKS IN ITS LENGTH WHEN TIGHTENING THE KNOT, WITHOUT EXERTING EXCESSIVE TENSION ON THE THREAD. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9438260&pc=GAM | 2210968-2019-90656 | 2019-10-01 | Malfunction | ETHICON INC. | 2019-12-09 | GAM | VICRYL VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN OVARIECTOMY IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, WHEN LIGATING, THE DEVICE BREAKS IN ITS LENGTH WHEN TIGHTENING THE KNOT, WITHOUT EXERTING EXCESSIVE TENSION ON THE THREAD. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: ADDITIONAL INVESTIGATION SUMMARY: THIRTY-TWO UNOPENED SAMPLES OF PRODUCT CODE JV461, LOT PBBCCMR0 WERE RECEIVED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RETURNED FOR EVALUATION. DURING THE VISUAL INSPECTION OF SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE REPRESENTATIVE SAMPLES, NO BREAKAGE SUTURE WAS FOUND AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PBBCCMR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9411144&pc=GAM | 2210968-2019-90532 |  | Malfunction | ETHICON INC. | 2019-12-04 | GAM | VICRYL PLUS UND 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND THE SUTURE WAS USED. IT WAS ALSO REPORTED THAT THE NEEDLE WAS MISSING FROM SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT (B)(4), AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY EVENT DETAILS: ¿SUTURE NEEDLES MISSING TWO SEPARATE TIMES FROM SUTURE¿ DID THE SUTURE SEPARATE FROM THE NEEDLE PRIOR OR DURING USE ON THE PATIENT? HOW MANY DEVICES WERE INVOLVED IN THIS SINGLE PROCEDURE? WHEN DID THE EVENT OCCURRED? NAME OF PROCEDURE? DEVICE RETURN STATUS/FOLLOW UP? TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9411247&pc=GAM | 2210968-2019-90534 |  | Malfunction | ETHICON INC. | 2019-12-04 | GAM | VICRYL PLUS UND 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND THE SUTURE WAS USED. IT WAS ALSO REPORTED THAT THE NEEDLE WAS MISSING FROM SUTURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PH7597, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY EVENT DETAILS: ¿SUTURE NEEDLES MISSING TWO SEPARATE TIMES FROM SUTURE¿ DID THE SUTURE SEPARATE FROM THE NEEDLE PRIOR OR DURING USE ON THE PATIENT? HOW MANY DEVICES WERE INVOLVED IN THIS SINGLE PROCEDURE? WHEN DID THE EVENT OCCURRED? NAME OF PROCEDURE? DEVICE RETURN STATUS/FOLLOW UP? TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412760&pc=GAM | 2210968-2019-90559 | 2019-11-12 | Malfunction | ETHICON INC. | 2019-12-04 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS BREAKING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE PROCEDURE DATE IS (B)(6) 2019. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT (B)(4), AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9403143&pc=GAM | 2210968-2019-90442 | 2019-11-13 | Malfunction | ETHICON INC. | 2019-12-03 | GAM | VICRYL CT UD 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A MINI PLASTIC ABDOMINAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, AT THE FIRST PASSING OF THE NEEDLE IN THE SKIN (WITHOUT FORCED), THE NEEDLE GOT BROKEN AT ABOUT 3 MM. ANOTHER LIKE DEVICE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PABDBRR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9395500&pc=GAM | 2210968-2019-90371 | 2019-11-11 | Malfunction | ETHICON INC. | 2019-12-02 | GAM | VICRYL VIO 2X70CM M3::USP2/0 NON NEEDLED | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. PART OF SUTURE WAS OUT OF PACKAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: SUMMARY: A SAMPLE OF PRODUCT WAS RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLE, COULD BE OBSERVED THE SUTURE MATERIAL EXTENDS COMPLETELY THROUGH THE ENTIRE SEAL MARGIN AND THE STERILE BARRIER IS COMPROMISED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING,PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED THE ASSIGNABLE CAUSE OF THE PACKAGING INTEGRITY IS CAUSED BY SUTURE IN SEAL AREA Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9375946&pc=GAM | 2210968-2019-90313 |  | Malfunction | ETHICON INC. | 2019-11-26 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TOTAL ABDOMINAL HYSTERECTOMY WITH SALPINGECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED TO CLOSE THE RECTAL SHEATH. DURING THE PROCEDURE, THE NEEDLE SNAPPED IN THE NEEDLE HOLDER WHILE PUSHING THROUGH TOUGH FASCIA. THE TWO PARTS TO THE NEEDLE WERE IMMEDIATELY IDENTIFIED AND WAS FOUND STRAIGHT AWAY. THE NEEDLE PIECES WERE REMOVED. THE SURGEON CONTINUED TO CLOSE THE ABDOMEN WITH ANOTHER SUTURE. THE SURGEON STATED A LARGE NEEDLE HOLDER WAS USED TO STITCH DUE TO THE PATIENT HAVING A TOUGHER FASCIA AND DEMONSTRATED THE JAWS WOULD SPAN MORE OF THE CURVED SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9369038&pc=GAM | 2210968-2019-90269 | 2019-11-05 | Malfunction | ETHICON INC. | 2019-11-25 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN OPHTHALMIC PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE CAME AWAY FROM THE NEEDLE AT THE SWAGE. A FEW DELAY TO THE PROCEDURE, AND REPLACEMENT SUTURE WAS OPENED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? NO ADVERSE TO ANY PATIENT. IT WAS REPORTED ¿SUTURE OFTEN COMES AWAY FROM NEEDLE AT THE SWAGE¿ IF THERE WERE PREVIOUS PATIENT EVENTS WITH THIS SUTURE, WERE THEY PREVIOUSLY REPORTED: NO PREVIOUS REPORT TO JJM, CANNOT PROVIDE EXACT NUMBER OF OCCURRENCE. HAVE SPOKEN TO REPORTER TO REPORT EACH CASE TO JJM MOVING FORWARD. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MK5059 BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? IT WAS REPORTED ¿SUTURE OFTEN COMES AWAY FROM NEEDLE AT THE SWAGE¿? IF THERE WERE PREVIOUS PATIENT EVENTS WITH THIS SUTURE, WERE THEY PREVIOUSLY REPORTED? IF YES, PLEASE PROVIDE PC NUMBERS. IF NOT PREVIOUSLY REPORTED, PLEASE CREATE A PRODUCT COMPLAINT FOR EACH PATIENT EVENT AND INCLUDE THE FOLLOWING INFORMATION: NAME AND DATE OF PROCEDURE, EVENT DATE, EVENT DESCRIPTION, WERE THERE ANY ADVERSE PATIENT CONSEQUENCES?, PRODUCT CODE AND LOT NUMBER, IS DEVICE INVOLVED IN THE EVENT AVAILABLE FOR RETURN?.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9365985&pc=GAM | 2210968-2019-90238 | 2019-10-25 | Malfunction | ETHICON INC. | 2019-11-22 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION SURGERY ON (B)(6) 2019 AND THE SUTURE WAS USED. THE TIP BROKE OFF DURING PROCEDURE. X-RAY WAS NOT DONE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9355491&pc=GAM | 2210968-2019-90082 | 2019-10-26 | Malfunction | ETHICON INC. | 2019-11-21 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SURGEON REPORTED THAT THE SUTURE BROKE WHILE KNOTTING DURING CLOSING MUCOSA. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). BATCH MANUFACTURING RECORDS OF NW2761 / T8006 WAS REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE FOUND TO MEET THE SPECIFICATION REQUIREMENT. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9355495&pc=GAM | 2210968-2019-90076 | 2019-10-26 | Malfunction | ETHICON INC. | 2019-11-21 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SURGEON REPORTED THAT THE SUTURE BROKE WHILE KNOTTING DURING CLOSING MUCOSA. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). BATCH MANUFACTURING RECORDS OF NW2761/T8006 WAS REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE FOUND TO MEET THE SPECIFICATION REQUIREMENT. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9355497&pc=GAM | 2210968-2019-90081 | 2019-10-26 | Malfunction | ETHICON INC. | 2019-11-21 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SURGEON REPORTED THAT THE SUTURE BROKE WHILE KNOTTING DURING CLOSING MUCOSA. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). BATCH MANUFACTURING RECORDS OF NW2761/T8006 WAS REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE FOUND TO MEET THE SPECIFICATION REQUIREMENT. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9358759&pc=GAM | 2210968-2019-90193 | 2019-11-02 | Malfunction | ETHICON INC. | 2019-11-21 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE WAS BROKEN WHILE KNOTTING WHEN BEING UTILIZED FOR CLOSING MUCOSA. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). BATCH MANUFACTURING RECORDS OF NW2761/T8010 WAS REVIEWED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE FOUND TO MEET THE SPECIFICATION REQUIREMENT. COMPLAINT SAMPLE: 1 EA OF OPEN COMPLAINT SAMPLE FOIL RECEIVED ALONG WITH PARTIALLY DISINTEGRATED SUTURE PIECES, NEEDLE, SINGLE BARRIER FOLDER FOR EVALUATION OF CODE NW2761/LOT T8010. COMPLAINT SAMPLE WAS VISUALLY INSPECTED AGAINST MENTIONED BREAKAGE SUTURE AND IT HAS BEEN OBSERVED THAT SUTURE RECEIVED WAS IN PARTIAL DISINTEGRATED CONDITION. AS THE COMPLAINT SAMPLE RECEIVED IN OPEN CONDITION FURTHER INVESTIGATION ON COMPLAINT SAMPLE CANNOT BE PERFORMED EXCEPT VISUAL INSPECTION. THE FOIL PACK WAS INSPECTED UNDER A 10 X MAGNIFICATION FOR ANY DAMAGE AND SHOWED A MULTITUDE OF WRINKLES OVER THE WHOLE FOIL ALONG WITH PINHOLE AT BOTTOM CAVITY SIDE OF THE PACK WERE OBSERVED. RETURNED NEEDLE WAS INSPECTED AGAINST 10X MAGNIFICATION AND FOUND TO SATISFACTORY. RETAINED SAMPLE EVALUATION: FIVE RETAIN SAMPLES OF INCIDENT CODES AND LOT NUMBER WERE RETRIEVED FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS LIKE PSEUDO SEAL, PIN HOLE, PRODUCT IN SEAL, PRESS MARKS ON PACK BUT NO SUCH DEFECTS WERE OBSERVED. THE PRIMARY PACKS WERE OPENED, AND SUTURES AND NEEDLES WERE FOUND INTACT. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE KINKS, WEAK SPOTS, BROKEN PIECE, FRAY, BRITTLENESS, DETACHED NEEDLES BUT NO SUCH DEFECTS WERE OBSERVED. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE BEND, CRACKED BARREL, FINS BUT NO SUCH DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WAS TESTED FOR KNOT PULL TENSILE STRENGTH TEST AND FOUND TO MEET THE SPECIFICATION. MANUFACTURING RECORDS REVIEW: AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED AND FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION. FROM THE ABOVE ANALYSIS, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. AS THE COMPLAINT IS RELATED TO BREAKAGE SUTURE, KNOT PULL TENSILE STRENGTH TEST IS APPLICABLE & MEASURABLE PARAMETER. KNOT PULL TENSILE STRENGTH TEST WAS PERFORMED OVER FIVE RETAINED SAMPLES TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. ALL THE INDIVIDUAL AS WELL AS AVERAGE KNOT PULL TENSILE STRENGTH VALUES OF RETAIN SAMPLE AND REFERENCE SAMPLE WERE FOUND TO MEET THE IN-HOUSE SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9347979&pc=GAM | 2210968-2019-90127 | 2019-09-09 | Malfunction | ETHICON INC. | 2019-11-20 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Break; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. PRIOR TO USE ON THE PATIENT, THE SUTURE BROKE WHEN IT WAS PULLED AT THE TIME OF PREPARING FOR THE SURGERY. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. UPON EVALUATION OF THE RETURNED DEVICE, THE HAD BEGUN WITH THE PROCESS OF DEGRADATION AND THE STRAND WAS BROKEN IN MULTIPLES PIECES. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AK5991, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE BREAKAGE SUTURE PRE-OP. AN EMPTY OPENED FOIL AND FOUR SUTURE PIECES OF PRODUCT CODE VR917, LOT # AK5991 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, THE SUTURE HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE STRAND WAS BROKEN IN MULTIPLES PIECES, THIS IS THE CAUSE THAT BREAKAGE SUTURE. THE PRODUCT CODE VR917 CONTAINS ABSORBABLE SUTURES AND AS THE SAMPLE WAS RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED AND NO FUNCTIONAL TEST CAN BE PERFORMED DUE TO SAMPLE CONDITION. THE FOIL WAS INSPECTED AND MULTIPLES HOLES IN CAVITY WERE NOTED. THIS CONDITION CONTRIBUTED TO DEGRADATION OF THE SUTURE. THE CONDITIONS OF THE HOLES COULD NOT BE DETERMINED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE PRE-OP, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9349191&pc=GAM | 2210968-2019-90143 | 2019-10-30 | Malfunction | ETHICON INC. | 2019-11-20 | GAM | VICRYL RAPIDE 0 1X90 V-34 | Break; Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND THE SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE DURING THE USE AND THE NEEDLE CAME OUT WHEN IT PULLED OFF FROM THE THREAD. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336123&pc=GAM | 2210968-2019-90038 | 2019-10-30 | Malfunction | ETHICON INC. | 2019-11-18 | GAM | VICRYL VIO 70CM M2 3/0 JB-1 VB | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A GASTRIC BYPASS PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE OF THE SUTURE BROKE OUT JUST BEFORE THE FIRST KNOT. ANTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IT WAS RECEIVED FOR ANALYSIS AN EMPTY OPENED FOLDER AND A DETACHED NEEDLE OF PRODUCT CODE W3110, LOT MPBBMDQ0. DURING THE VISUAL INSPECTION OF A DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED, MARKS WERE OBSERVED ON THE EDGE OF THE NEEDLE THAT APPEARS TO BE BY THE USE OF A SURGICAL INSTRUMENT, THIS CONDITION CAUSES THE NEEDLE PULL OFF. THE SUTURE WAS NOT RECEIVED FOR EVALUATION. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/ PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF THE PERFORMANCE PULL OFF - SUTURE NEEDLE IS AN IMPROPER HANDLING. TO AVOID THIS KIND OF DAMAGE: GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY WHAT IS MEANT BY ¿THE NEEDLE OF VICRYL SUTURE BROKE OUT JUST BEFORE THE FIRST KNOT¿. DID THE SUTURE DETACH FROM THE NEEDLE? OR DID THE SUTURE BREAK? OR DID THE NEEDLE BREAK? IT MEANS, THAT THE SUTURE DETACHED FROM THE NEEDLE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT MPBBMDQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY WHAT IS MEANT BY ¿THE NEEDLE OF VICRYL SUTURE BROKE OUT JUST BEFORE THE FIRST KNOT¿. DID THE SUTURE DETACH FROM THE NEEDLE? OR DID THE SUTURE BREAK? OR DID THE NEEDLE BREAK?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9325689&pc=GAM | 2210968-2019-89948 | 2019-10-23 | Malfunction | ETHICON INC. | 2019-11-15 | GAM | VICRYL VIO 150CM M4 1 LOOP CT | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A SURGERY ON (B)(6) 2019 AND THE SUTURE WAS USED TO CLOSE THE FASCIA. THE NEEDLE SEPARATED FROM THE SUTURE. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO DELAY IN SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: IT WAS RECEIVED FOR ANALYSIS TEN UNOPENED SAMPLES OF PRODUCT CODE V1408, LOT PD6457. DURING THE VISUAL INSPECTION OF TEN UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLE RECEIVED, NO PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLE MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PD6457 BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WAS THE NEEDLE SEPARATED FROM SUTURE ISSUE OCCURRED PRIOR OR DURING USE ON PATIENT, AND PROVIDE QUANTITY INVOLVED IF ANY? NEEDLE SEPARATED DURING USE. SEVERAL SUTURES WERE INVOLVED - PRECISE NUMBER IS UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9330340&pc=GAM | 2210968-2019-89971 | 2019-10-24 | Malfunction | ETHICON INC. | 2019-11-15 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN OB-GYN PROCEDURE ON (B)(6) 2019 AND THE SUTURE WAS USED. DURING USE, THE NEEDLE WAS DETACHED FROM THE SUTURE. IT WAS NOT A CONTROL RELEASE NEEDLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9330965&pc=GAM | 2210968-2019-89976 |  | Malfunction | ETHICON INC. | 2019-11-15 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT PRODUCT WAS NOT USED ON PATIENT. THE PRODUCT IS LABELED 'NOT FOR HUMAN USE'. NO ADVERSE PATIENT CONSEQUENCES. UPON DEVICE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED ON THE BOX. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NON-CONFORMANCE, WHICH WAS RAISED TO ADDRESS THE EVENT REPORTED. THE NECESSARY ACTIONS HAVE BEEN TAKEN TO ADDRESS THE ISSUE. DEVICE EVALUATION SUMMARY: UPON VISUAL INSPECTION OF ONE PICTURE #1, AN OPENED BOX OF PRODUCT CODE V372H WITH LOT# PHBBCJQ0 AND PRINTED AT THE SECONDARY PACKAGING THE LEGEND OF "NOT FOR HUMAN USE" AND COULD BE OBSERVED. IN THE PICTURE #2 AN UNOPENED FOIL AND NO DAMAGED OR DEFECT WAS NOTED. THE PRODUCTS MET ALL SPECIFICATIONS FOR USE ON PATIENTS. DEVICE EVALUATION SUMMARY: ONE OPENED BOX OF PRODUCT CODE V372H, LOT PHBBCJQ0 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE BOX, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. THE PRODUCTS MET ALL SPECIFICATIONS FOR USE ON PATIENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9331087&pc=GAM | 2210968-2019-89974 | 2019-08-27 | Malfunction | ETHICON INC. | 2019-11-15 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT PRODUCT WAS NOT USED ON PATIENT. THE PRODUCT IS LABELED 'NOT FOR HUMAN USE'. NO ADVERSE PATIENT CONSEQUENCES. UPON DEVICE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" ON THE ARTWORK OF THE PRODUCT COULD BE OBSERVED ON THE BOX. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE, WHICH WAS RAISED TO ADDRESS THE EVENT REPORTED. THE NECESSARY ACTIONS HAVE BEEN TAKEN TO ADDRESS THE ISSUE. DEVICE EVALUATION SUMMARY: UPON VISUAL INSPECTION OF ONE PICTURE, A SEALED BOX OF PRODUCT CODE V372H WITH LOT PHBBCJQ0 AND A PRINTED THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. DEVICE EVALUATION SUMMARY: TWENTY SEALED BOXES OF PRODUCT CODE V372, LOT PHBBCJQ0 WERE RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE BOXES, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. THE PRODUCTS MET ALL SPECIFICATIONS FOR USE ON PATIENTS. NO RISK TO PATIENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9322447&pc=GAM | 2210968-2019-89913 | 2019-10-22 | Malfunction | ETHICON INC. | 2019-11-14 | GAM | VICRYL RAP UND 75CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT LAPAROSCOPIC GASTRIC BYPASS ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE CAME LOOSE DURING SUTURING THE SKIN. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9323445&pc=GAM | 2210968-2019-89922 | 2019-10-21 | Malfunction | ETHICON INC. | 2019-11-14 | GAM | VICRYL UND 90CM M3 USP2/0::SGLE ARMED CT-1 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT TOTAL KNEE REPLACEMENT ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS RIPPING APART DURING CLOSING. A SIMILAR DEVICE WAS USED TO COMPLETE THE CASE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION: THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. THE POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH KP2010; EXPIRATION DATE: 11/30/2021. DATE OF MFG.: 12/01/2016. BATCH KM7023; EXPIRATION DATE: 10/31/2021. DATE OF MFG.: 11/14/2016. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE POSSIBLE LOTS FINISHED DEVICES KP2010 AND KM7023, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317402&pc=GAM | 2210968-2019-89896 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-11-13 | GAM | VICRYL VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY IN 2019 AND THE SUTURE WAS USED. THE RUPTURE OF THE NEEDLE OCCURRED DURING SURGICAL PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V347H; JD8CQJM0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317902&pc=GAM | 2210968-2019-89897 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-11-13 | GAM | VICRYL CT VIO 8X45CM M5 | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY IN 2019 AND THE SUTURE WAS USED. THE RUPTURE OF THE NEEDLE OCCURRED DURING SURGICAL PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V755G; PEBDBKR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9310160&pc=GAM | 2210968-2019-89844 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-11-12 | GAM | VICRYL RAP CT UD 75CM M3 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE 2 MM FROM THE NEEDLE. THE SURGEON CONFIRMED THAT NO EXCESSIVE TENSION WAS EXERTED. THE PROCEDURE WAS COMPLETED WITH THE SAME LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9310576&pc=GAM | 2210968-2019-89845 |  | Malfunction | ETHICON INC. | 2019-11-12 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Device Embedded In Tissue or Plaque | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE SUTURE BROKE OFF ,INTO THE PATIENT. THE NEEDLE TIP REMAINS IN THE PATIENT AND THERE ARE NO PLANS TO LOOK FOR IT FURTHER OR DO ANY OTHER INTERVENTION. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PRODUCT CODE AND LOT #? THE HOSPITAL DID NOT KEEP A NOTE OF THE LOT NUMBER OR CODE AND THERE IS NOTHING TO RETURN. PROCEDURE NAME? WAS THE TIP OF THE NEEDLE RETRIEVED FROM THE PATIENT? THE NEEDLE TIP REMAINS IN THE PATIENT IF YES, HOW? IF NOT, DOES THE PIECE OF NEEDLE STILL REMAIN IN PATIENT BODY? AND WHERE? THE NEEDLE TIP REMAINS IN THE PATIENT. ANY PLANS TO REMOVE NEEDLE FROM PATIENT BODY WITH MEDICAL/SURGICAL INTERVENTION? TO MY KNOWLEDGE THERE ARE NO PLANS TO LOOK FOR IT FURTHER OR DO ANY OTHER INTERVENTION. ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THE EVENT REPORT? I HAVE NOT BEEN ADVISED AS TO ANY PATIENT CONSEQUENCES AS A RESULT OF REPORTING THIS INCIDENT. PATIENT CURRENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9310633&pc=GAM | 2210968-2019-89848 | 2019-10-21 | Malfunction | ETHICON INC. | 2019-11-12 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE SURGERY, THE SUTURE WAS BREAKING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 11/26/2019. ADDITIONAL INVESTIGATION SUMMARY: COMPLAINT SAMPLE: 1 EA OF OPEN COMPLAINT SAMPLE FOIL RECEIVED ALONG WITH DISINTEGRATED SUTURE PIECES, ONE RUSTED NEEDLE, ZIPPER TRAY AND LID FOR EVALUATION OF CODE NW2382/LOT T5014. COMPLAINT SAMPLE WAS VISUALLY INSPECTED AGAINST MENTIONED BREAKAGE SUTURE AND IT HAS BEEN OBSERVED THAT SUTURE RECEIVED WAS IN COMPLETE DISINTEGRATED CONDITION. AS THE COMPLAINT SAMPLE RECEIVED IN OPEN CONDITION FURTHER INVESTIGATION ON COMPLAINT SAMPLE CANNOT BE PERFORMED EXCEPT VISUAL INSPECTION. THE FOIL PACK WAS INSPECTED UNDER A 10 X MAGNIFICATION FOR ANY DAMAGE AND SHOWED A MULTITUDE OF WRINKLES OVER THE WHOLE FOIL ALONG WITH SEVERAL PINHOLES AT THE TOP LABEL SIDE AS WELL AS BOTTOM CAVITY SIDE OF THE PACK WERE OBSERVED. RETURNED NEEDLE WAS INSPECTED AGAINST 10X MAGNIFICATION AND FOUND TO BE RUSTED. RETAINED SAMPLE EVALUATION: FIVE RETAIN SAMPLES OF INCIDENT CODES AND LOT NUMBER WERE RETRIEVED FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS LIKE PSEUDO SEAL, PIN HOLE, PRODUCT IN SEAL, PRESS MARKS ON PACK BUT NO SUCH DEFECTS WERE OBSERVED. THE PRIMARY PACKS WERE OPENED, AND SUTURES AND NEEDLES WERE FOUND INTACT. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE KINKS, WEAK SPOTS, BROKEN PIECE, FRAY, BRITTLENESS, DETACHED NEEDLES BUT NO SUCH DEFECTS WERE OBSERVED. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE BEND, CRACKED BARREL, FINS BUT NO SUCH DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WAS TESTED FOR KNOT PULL TENSILE STRENGTH TEST AND FOUND TO MEET THE SPECIFICATION. MANUFACTURING RECORDS REVIEW: AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED FOR ANY PROCESS DEVIATION, BUT NO DEVIATION WAS OBSERVED. FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION. FROM THE ABOVE ANALYSIS, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. AS THE COMPLAINT IS RELATED TO BREAKAGE SUTURE, KNOT PULL TENSILE STRENGTH TEST IS APPLICABLE & MEASURABLE PARAMETER. KNOT PULL TENSILE STRENGTH TEST WAS PERFORMED OVER FIVE RETAINED SAMPLES TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. ALL THE INDIVIDUAL AS WELL AS AVERAGE KNOT PULL TENSILE STRENGTH VALUES OF RETAIN SAMPLE AND REFERENCE SAMPLE WERE FOUND TO MEET THE USP SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS. BASED ON THE ANALYSIS, ¿THE DETECTED SUTURE BREAKAGE ISSUE MAY HAVE BEEN OBSERVED DUE TO INGRESS OF HUMIDITY THROUGH THE OBSERVED PINHOLES¿. THE OBSERVED PIN HOLES MIGHT HAVE GENERATED DURING IMPROPER STORAGE AND HANDLING OF THE PRODUCT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312429&pc=GAM | 2210968-2019-89878 | 2019-10-16 | Malfunction | ETHICON INC. | 2019-11-12 | GAM | VICRYL* UND 3-0 70CM (1)SC-20 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE BROKE FROM THE SUTURE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9305126&pc=GAM | 2210968-2019-89793 | 2019-10-16 | Malfunction | ETHICON INC. | 2019-11-11 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS DETACHED FROM THE NEEDLE DURING INTERRUPTED SUTURING. IT WAS NOT A CONTROL RELEASE NEEDLE. FURTHER DETAILS ARE NOT PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AND NO NON-CONFORMANCES WERE IDENTIFIED. AN EMPTY LABELED WINDING FORMER, A DETACHED NEEDLE AND MULTIPLES SUTURES PIECES OF PRODUCT WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, THE SWAGE AND ATTACHMENT AREA PRESENT MARKS AND THE BARREL HOLE WAS FLAT BY USE OF A SURGICAL INSTRUMENT. THE SUTURE HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE STRAND WAS BROKEN IN MULTIPLES PIECES, THIS IS THE CAUSE THAT BREAKAGE SUTURE. THE PRODUCT CODE CONTAINS ABSORBABLE SUTURE AND AS THE SAMPLE WAS RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED AND NO FUNCTIONAL TEST CAN BE PERFORMED DUE TO SAMPLE CONDITION. ACCORDING TO VISUAL INSPECTION THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306122&pc=GAM | 2210968-2019-89803 | 2019-10-16 | Malfunction | ETHICON INC. | 2019-11-11 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 3-0 UND | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2019, IT WAS FOUND IN THE INVENTORY THAT THERE WAS INDICATION OF ¿REVERSE CUTTING¿ ON THE BOX THOUGH THE SPEC OF THE NEEDLE SHAPE IS TAPER POINT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306123&pc=GAM | 2210968-2019-89802 | 2019-10-16 | Malfunction | ETHICON INC. | 2019-11-11 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 3-0 UND | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2019, IT WAS FOUND IN THE INVENTORY THAT THERE WAS INDICATION OF ¿REVERSE CUTTING¿ ON THE BOX THOUGH THE SPEC OF THE NEEDLE SHAPE IS TAPER POINT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306124&pc=GAM | 2210968-2019-89805 | 2019-10-16 | Malfunction | ETHICON INC. | 2019-11-11 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 3-0 UND | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2019, IT WAS FOUND IN THE INVENTORY THAT THERE WAS INDICATION OF ¿REVERSE CUTTING¿ ON THE BOX THOUGH THE SPEC OF THE NEEDLE SHAPE IS TAPER POINT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306136&pc=GAM | 2210968-2019-89804 | 2019-10-16 | Malfunction | ETHICON INC. | 2019-11-11 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 3-0 UND | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2019, IT WAS FOUND IN THE INVENTORY THAT THERE WAS INDICATION OF ¿REVERSE CUTTING¿ ON THE BOX THOUGH THE SPEC OF THE NEEDLE SHAPE IS TAPER POINT. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9307588&pc=GAM | 2210968-2019-89685 | 2019-10-15 | Malfunction | ETHICON INC. | 2019-11-11 | GAM | VICRYL CT UD 45CM M1.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND THE SUTURE WAS USED. THE NEEDLE GOT BROKEN IN TWO PARTS WHEN PERFORMING THE SUTURE OF THE LAST STITCH. THERE WERE NO PATIENT CONSEQUENCES REPORTED . Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V494H; PDBDCST0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9295448&pc=GAM | 2210968-2019-89675 | 2019-10-15 | Malfunction | ETHICON INC. | 2019-11-08 | GAM | VICRYL PLUS VIO 70CM::M2 USP3/0 SGLE ARMED SH-1 | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A THYROIDECTOMY SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE WHEN SEWING IN THE SURGERY OF THYROIDECTOMY. CHANGED ANOTHER DEVICE TO COMPLETE THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE DEVICE HAD PERFORMANCE BREAKAGE SUTURE. IT WAS RECEIVED FOR ANALYSIS AN A LABELED WINDING FORMER WITH A SUTURE PIECE PARTIALLY DISPENSED AND A NEEDLE / SUTURE PIECE OF PRODUCT CODE VCP311W. DURING THE VISUAL INSPECTION OF NEEDLE / SUTURE PIECE, MARKS ON THE BODY NEEDLE THAT APPEARS TO BE BY USE OF SURGICAL INSTRUMENTS WAS NOTED; ALSO, THE SUTURES PIECES PRESENT BODY FLUIDS AND WAS DAMAGED (BUSTED). THE SUTURE PIECE WAS DISPENSED AND EXAMINED ALONG OF STRAND AND THE END WAS NOTED DAMAGED (CUT) APPEARS TO BE BY USE OF A SURGICAL INSTRUMENT. THE MANUFACTURING RECORDS COULDN'T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE IS DAMAGED SUTURE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9295486&pc=GAM | 2210968-2019-89681 | 2019-10-15 | Malfunction | ETHICON INC. | 2019-11-08 | GAM | VICRYL CT UD 45CM M1.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE IN TWO PARTS WHEN PERFORMING THE SUTURE OF THE LAST STITCH. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). IT WAS REPORTED PERFORMANCE NEEDLE BREAKAGE. A FAILED SUTURE NEEDLE WAS SUBMITTED FOR FRACTOGRAPHIC EVALUATION. A FRACTURE WAS OBSERVED AT THE SUTURE ATTACHMENT OF SAMPLE A. SAMPLE A WAS RECEIVED IN TWO PIECES, ONLY ONE OF THE MATING FRACTURE SURFACES WAS EXAMINED FOR THIS EVALUATION. A MICROSCOPE WAS USED TO EXAMINE THE FRACTURE SURFACE AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACE WAS EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE ATTACHMENT AREA OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V494H; PDBDCST0 NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9295874&pc=GAM | 2210968-2019-89727 | 2019-10-01 | Malfunction | ETHICON INC. | 2019-11-08 | GAM | VICRYL CTD VIO 1 90CM CT PLUS | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND THE SUTURE WAS USED. THE NEEDLE TIP BROKE INSIDE THE PATIENT, SIGNIFICANT TIME WAS TAKEN TO RETRIEVE FRAGMENT. THE PROCEDURE COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH MMBDTJQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. DEVICE EVALUATION SUMMARY: A SUTURE NEEDLE WAS SUBMITTED FOR EVALUATION. A FRACTURE WAS OBSERVED AT THE TIP OF THE NEEDLE. THE NEEDLE WAS RECEIVED IN TWO PIECES, ONLY ONE OF THE MATING FRACTURE SURFACES WAS EXAMINED FOR THIS EVALUATION. A SCANNING ELECTRON MICROSCOPE (SEM) WAS USED TO EXAMINE THE FRACTURE SURFACES AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACES WERE EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. MECHANICAL DAMAGE OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDES ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE TIP OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. THE NEEDLE BREAKAGE WAS CONFIRMED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? UNKNOWN. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN UNKNOWN. WAS THE PROCEDURE COMPLETED SUCCESSFULLY? YES. ANY ADVERSE PATIENT CONSEQUENCES AND WHAT MEDICAL / SURGICAL INTERVENTION WAS PROVIDED? UNKNOWN. NAME OF THE PROCEDURE? UNKNOWN. WHAT IS THE PATIENT'S CURRENT STATUS? UNKNOWN. LOT NUMBER? MMBDTJQ0. DEVICE RETURN STATUS / FOLLOW UP? RECEIVED TO BE SHIPPED SHORTLY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9298242&pc=GAM | 2210968-2019-89751 | 2019-10-02 | Malfunction | ETHICON INC. | 2019-11-08 | GAM | VICRYL CT VIO 45CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. A SUTURE CAME LOOSE FROM THE NEEDLE AFTER TWO STITCHES. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A DETACHED NEEDLE WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF A DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE EDGE OF THE BARREL HOLE COULD BE OBSERVED WITH MARKS DUE TO USE OF A SURGICAL INSTRUMENT. IN ADDITION, A SUTURE PIECE WAS NOTED STILL ATTACHED TO NEEDLE AND THE END WAS CUT. TO AVOID THIS KIND OF DAMAGE: GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. IN HANDLING THIS OR ANY OTHER SUTURE MATERIAL, CARE SHOULD BE TAKEN TO AVOID DAMAGE FROM HANDLING. AVOID CRUSHING OR CRIMPING DAMAGE DUE TO APPLICATION OF SURGICAL INSTRUMENTS SUCH AS FORCEPS OR NEEDLE HOLDERS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF THE PERFORMANCE - PULL OFF SUTURE NEEDLE IS AN IMPROPER HANDLING. UNOPENED SAMPLES WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 2 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-89750. ANALYSIS RESULTS HAVE BEEN INCLUDED IN FOLLOW-UP REPORTS FOR EACH. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V392ZG; PCBCCTQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9298243&pc=GAM | 2210968-2019-89750 | 2019-10-02 | Malfunction | ETHICON INC. | 2019-11-08 | GAM | VICRYL CT VIO 45CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. UPON SUTURING, A SUTURE LOOSENED DIRECTLY FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). UNOPENED SAMPLES WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITIONS OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 2 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-89751. ANALYSIS RESULTS HAVE BEEN INCLUDED IN FOLLOW-UP REPORTS FOR EACH. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V392ZG; PCBCCTQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9299173&pc=GAM | 2210968-2019-89755 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-11-08 | GAM | VICRYL CT VIO 70CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND SUTURE WAS USED. DURING USAGE IN SURGERY SUTURE REATTACHED FROM NEEDLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. DEVICE EVALUATION SUMMARY: THE ACTUAL DEVICE WAS RECEIVED AS A DETACHED NEEDLE OF PRODUCT CODE V317H, LOT PHBCDMQ0 WAS RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE AS EXPECTED AND REMNANT OF THE SUTURE WAS NOTED INTO THE BARREL HOLE AND SLIGHTLY MARKS WERE OBSERVED ON THE EDGE OF THE NEEDLE THAT APPEARS TO BE BY THE USE OF A SURGICAL INSTRUMENT. IN ADDITION, THE SUTURE WAS NOT RECEIVED FOR ANALYSIS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE SAMPLE CONDITION THE ASSIGNABLE OF THE PERFORMANCE PULL OFF - SUTURE NEEDLE, SUGGEST AN IMPROPER HANDLING OF THE SAMPLE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS MEANT BY "DURING USAGE IN SURGERY SUTURE REATTACHED FROM NEEDLE"? DID THE NEEDLE PULL AWAY FROM THE SUTURE? WHAT IS THE PRODUCT CODE AND LOT# OF THE DEVICE? DID THIS EVENT OCCUR INTRA-OP? WHAT WAS THE NAME OF THE PROCEDURE? WHAT IS THE PROCEDURE DATE? WAS THERE ANY PATIENT CONSEQUENCE? WAS ANY TREATMENT RENDERED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9299237&pc=GAM | 2210968-2019-89766 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-11-08 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND SUTURE WAS USED. DURING USAGE IN SURGERY SUTURE REATTACHED FROM NEEDLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEVICE EVALUATION SUMMARY: ONE DETACHED BLACK NEEDLE OF PRODUCT CODE UNKNOWN WAS RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE OF THE NEEDLE WAS EXAMINED UNDER MAGNIFICATION, NO SUTURE REMNANT WAS OBSERVED. THE SUTURE WAS NOT RECEIVED FOR EVALUATION TO DETERMINE THE ASSIGNABLE CAUSE. THE MANUFACTURING RECORDS COULD NOT BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS MEANT BY "DURING USAGE IN SURGERY SUTURE REATTACHED FROM NEEDLE"? DID THE NEEDLE PULL AWAY FROM THE SUTURE? WHAT IS THE PRODUCT CODE AND LOT# OF THE DEVICE? DID THIS EVENT OCCUR INTRA-OP? WHAT WAS THE NAME OF THE PROCEDURE? WHAT IS THE PROCEDURE DATE? WAS THERE ANY PATIENT CONSEQUENCE? WAS ANY TREATMENT RENDERED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9288152&pc=GAM | 2210968-2019-89641 | 2019-10-11 | Malfunction | ETHICON INC. | 2019-11-06 | GAM | VICRYL CT VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CESAREAN SECTION PROCEDURE ON (B)(6) AND SUTURE WAS USED. THE NEEDLE WAS BROKEN DURING USE. ANOTHER NEEDLE WAS USED TO COMPLETE THE PROCEDURE. THE BROKEN PIECES WERE RETRIEVED FROM THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE ACTUAL SAMPLE WAS RETURNED FOR EVALUATION. A FRACTURE WAS OBSERVED IN THE BODY OF THE NEEDLE. THE FRACTURE SURFACES WERE EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE.THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. A CUP-AND-CONE SHAPED FRACTURE AND MACROSCOPIC PLASTIC DEFORMATION OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDE ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED FROM MECHANICAL DEFORMATION, INDENTS AND SCRATCHES, DUE TO GRIPPING, BENDING AND OVERSTRESSING OF THE NEEDLE. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW THAT WOULD CAUSE PREMATURE FAILURE. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9288478&pc=GAM | 2210968-2019-89649 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-11-06 | GAM | VICRYL VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORTHOPEDIC SURGICAL PROCEDURE IN 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD PULLED OFF OF THE NEEDLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV461; PDBBHLR0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9281114&pc=GAM | 2210968-2019-89455 | 2019-10-01 | Malfunction | ETHICON INC. | 2019-11-05 | GAM | VICRYL VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. AT THE BEGINNING OF THE SUTURE, THE NEEDLE PULLED OFF FROM THE THREAD AT THE JUNCTION POINT. THE NEEDLE HAS TO BE SEARCHED IN THE MUCOUS TISSUES AND WAS REMOVED WITH DIFFICULTY. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE (B)(4) NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9282133&pc=GAM | 2210968-2019-89560 | 2019-10-08 | Malfunction | ETHICON INC. | 2019-11-05 | GAM | VICRYL RAPIDE UND 2/0 70CM MH PLUS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN OB-GYN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS BROKEN DURING USE. ANOTHER NEEDLE-SUTURE WAS USED WITH NO PROBLEM. FURTHER DETAILS ARE NOT PROVIDED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PG1674H; PBCCZCE0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. AN EMPTY OPENED FOIL, A NEEDLE/SUTURE PIECE AND A SUTURE PIECE OF PRODUCT CODE PG1674, LOT # PBCCZCE0 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE, THE SWAGE AND ATTACHMENT AREA OF THE NEEDLE WAS NOTED TO BE AS EXPECTED. THE SUTURE WAS RECEIVED BROKEN (CUT) IN TWO SECTION APPEARS TO BE BY USE OF THE SURGICAL INSTRUMENT AND BODY FLUIDS WERE NOTED. THE PRODUCT CODE PG1674 CONTAINS AN ABSORBABLE SUTURE AND AS THE SAMPLE WAS RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURES TO THE ENVIRONMENT COULD NOT BE DETERMINED AND NO FUNCTIONAL TEST CAN BE PERFORMED DUE TO SAMPLE CONDITION. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED, THE ASSIGNABLE CAUSE OF THE BREAKAGE SUTURE WAS AN IMPROPER HANDLING. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9271750&pc=GAM | 2210968-2019-89434 |  | Malfunction | ETHICON INC. | 2019-11-04 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, AT THE TIME OF USE, THE EYE OF THE NEEDLE WAS BROKEN, ONE PART REMAINING ATTACHED TO THE THREAD, THE OTHER TO THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. . Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9273349&pc=GAM | 2210968-2019-89471 | 2019-10-06 | Malfunction | ETHICON INC. | 2019-11-04 | GAM | VICRYL CT VIO 45CM M1.5 | Break; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE BREAKAGE SUTURE AND ASSEMBLY MISSING COMPONENT / INCORRECT QUANTITY. AN OPENED SAMPLE OF PRODUCT CODE 18531, LOT # PJBDJSQ0 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, THE SWAGE AND ATTACHMENT AREA OF THE NEEDLE WERE NOTED TO BE AS EXPECTED AND THE SUTURE HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE STRAND WAS BROKEN IN MULTIPLES PIECES, THIS IS THE CAUSE THAT BREAKAGE SUTURE. THE PRODUCT CODE 18531 CONTAINS ABSORBABLE SUTURE AND AS THE SAMPLE WAS RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED AND NO FUNCTIONAL TEST CAN BE PERFORMED DUE TO SAMPLE CONDITION. THE FOIL WAS INSPECTED AND MULTIPLES HOLES IN CAVITY WERE NOTED. THIS CONDITION CONTRIBUTED TO DEGRADATION OF THE SUTURE. THE CONDITIONS OF THE HOLES COULD NOT BE DETERMINED. ACCORDING TO VISUAL INSPECTION NO MISSING COMPONENT / INCORRECT QUANTITY WERE NOTED AND THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9273354&pc=GAM | 2210968-2019-89472 | 2019-10-06 | Malfunction | ETHICON INC. | 2019-11-04 | GAM | VICRYL CT VIO 45CM M1.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON 10/06/2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY ==> IT WAS REPORTED PERFORMANCE BREAKAGE SUTURE AND ASSEMBLY MISSING COMPONENT / INCORRECT QUANTITY. A FOLDER WITH A DETACHED NEEDLE OF PRODUCT CODE 18531, LOT # MPBBWKQ0 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED AND MARKS COULD BE OBSERVED ON THE BODY NEEDLE THAT APPEAR TO BE BY USE OF A SURGICAL INSTRUMENT. IN ADDITION, THE DYED SUTURE WAS TRANSFERRED TO FOLDER WITH INDICATING THAT THE SUTURE WAS WINDING INTO THE FOLDER. DUE TO THE SUTURE WAS NOT RECEIVED, NO FUNCTIONAL TEST COULD BE PERFORMED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED NO CONCLUSION COULD BE REACHED AS TO WHAT MAY HAVE CAUSED THE REPORTED INCIDENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9273357&pc=GAM | 2210968-2019-89473 | 2019-10-06 | Malfunction | ETHICON INC. | 2019-11-04 | GAM | VICRYL CT VIO 45CM M1.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE BREAKAGE SUTURE AND ASSEMBLY MISSING COMPONENT / INCORRECT QUANTITY. AN OPENED SAMPLE OF PRODUCT CODE 18531, LOT # PJBDJSQ0 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, THE SWAGE AND ATTACHMENT AREA OF THE NEEDLE WERE NOTED TO BE AS EXPECTED AND THE SUTURE HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE STRAND WAS BROKEN IN MULTIPLES PIECES, THIS IS THE CAUSE THAT BREAKAGE SUTURE. THE PRODUCT CODE 18531 CONTAINS ABSORBABLE SUTURE AND AS THE SAMPLE WAS RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED AND NO FUNCTIONAL TEST CAN BE PERFORMED DUE TO SAMPLE CONDITION. ACCORDING TO VISUAL INSPECTION NO MISSING COMPONENT / INCORRECT QUANTITY WERE NOTED AND THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9273359&pc=GAM | 2210968-2019-89474 | 2019-10-06 | Malfunction | ETHICON INC. | 2019-11-04 | GAM | VICRYL CT VIO 45CM M1.5 | Break; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY ==> IT WAS REPORTED PERFORMANCE BREAKAGE SUTURE AND ASSEMBLY MISSING COMPONENT / INCORRECT QUANTITY. AN OPENED SAMPLE OF PRODUCT CODE 18531, LOT # MPBBWKQ0 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLE, THE SWAGE AND ATTACHMENT AREA OF THE NEEDLE WERE NOTED TO BE AS EXPECTED AND THE SUTURE HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THE STRAND WAS BROKEN IN MULTIPLES PIECES, THIS IS THE CAUSE THAT BREAKAGE SUTURE. THE PRODUCT CODE 18531 CONTAINS ABSORBABLE SUTURE AND AS THE SAMPLE WAS RECEIVED OPEN, THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED AND NO FUNCTIONAL TEST CAN BE PERFORMED DUE TO SAMPLE CONDITION. THE FOIL WAS INSPECTED AND MULTIPLES HOLES IN CAVITY WERE NOTED. THIS CONDITION CONTRIBUTED TO DEGRADATION OF THE SUTURE. THE CONDITIONS OF THE HOLES COULD NOT BE DETERMINED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO VISUAL INSPECTION NO MISSING COMPONENT / INCORRECT QUANTITY WERE NOTED AND THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE, WAS CAUSED BY SUTURE DEGRADATION EXPOSURE TO ENVIRONMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9273363&pc=GAM | 2210968-2019-89475 | 2019-10-06 | Malfunction | ETHICON INC. | 2019-11-04 | GAM | VICRYL CT VIO 45CM M1.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PERFORMANCE BREAKAGE SUTURE AND ASSEMBLY MISSING COMPONENT / INCORRECT QUANTITY. A FOLDER WITH A DETACHED NEEDLE OF PRODUCT CODE 18531, LOT # MPBBWKQ0 WAS RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED AND MARKS COULD BE OBSERVED ON THE BODY NEEDLE THAT APPEAR TO BE BY USE OF A SURGICAL INSTRUMENT. IN ADDITION, THE DYED SUTURE WAS TRANSFERRED TO FOLDER WITH INDICATING THAT THE SUTURE WAS WINDING INTO THE FOLDER. DUE TO THE SUTURE WAS NOT RECEIVED, NO FUNCTIONAL TEST COULD BE PERFORMED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLE RECEIVED NO CONCLUSION COULD BE REACHED AS TO WHAT MAY HAVE CAUSED THE REPORTED INCIDENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9266149&pc=GAM | 2210968-2019-88510 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-11-01 | GAM | VICRYL CT UD 90CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT OSTEOSYNTHESIS CLAVICLE PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SURGEON OPINED THAT THE QUALITY OF THE THREAD WAS DIFFERENT AND MORE BRITTLE WHEN MAKING THE KNOT. THERE WAS NO ADVERSE PATIENT CONSEQUENCE. Manufacturer Narrative: (B)(4). ¿ON OCTOBER 9 & 10 2019 CDRH EXPERIENCED INTERMITTENT ISSUE CAUSING MEDWATCH REPORTS TO REMAIN STUCK IN ACK 2. THIS REPORT WAS AFFECTED BY THIS ISSUE AND IS NOW BEING RESUBMITTED. THE ORIGINAL SUBMISSION WAS MADE WITHIN THE 30 DAY REPORTING TIMELINE ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PROCEDURE NAME: OSTEOSYNTHESIS CLAVICLE. THE AMOUNT OF TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION¿ ? THE TIME BETWEEN THE SUTURE PLACEMENT AND THE FINDING THE PROBLEM IS OF 60 DAYS. THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION"? UNKNOWN. EVENT OF POST OP ABSORPTION CAPTURED IN MW 2210968-2019-88509.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9267608&pc=GAM | 2210968-2019-89384 | 2019-10-07 | Malfunction | ETHICON INC. | 2019-11-01 | GAM | VICRYL* VLT PLUS 0 70CM (1)SH | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CESAREAN SECTION ON (B)(6) AND SUTURE WAS USED. IT WAS FOUND THAT THE SUTURE BROKE DURING SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9267649&pc=GAM | 2210968-2019-89385 | 2019-10-07 | Malfunction | ETHICON INC. | 2019-11-01 | GAM | VICRYL* VLT PLUS 0 70CM (1)SH | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CESAREAN SECTION ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS FOUND THAT THE SUTURE BROKE DURING SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AK6612, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE SAMPLE WAS SCRAPPED BY THE HOSPITAL. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9267670&pc=GAM | 2210968-2019-89387 | 2019-10-07 | Malfunction | ETHICON INC. | 2019-11-01 | GAM | VICRYL* VLT PLUS 0 70CM (1)SH | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CESAREAN SECTION ON (B)(6) 2019 AND SUTURE WAS USED. IT WAS FOUND THAT THE SUTURE BROKE DURING SURGERY. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AK6612, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE SAMPLE WAS SCRAPPED BY THE HOSPITAL. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9259841&pc=GAM | 2210968-2019-89261 | 2019-08-08 | Malfunction | ETHICON INC. | 2019-10-31 | GAM | VICRYL*VLT PLUS 3-0 70CM 1-SH | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A GASTRECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE POINT OF THE NEEDLE HAD DEFECT. THE SURGEON USED THE NEEDLE WHICH CAUSED DAMAGE TO THE TISSUE. THE PROCEDURE WAS DELAYED BY 15 MINUTES. ANOTHER SUTURE WAS USED TO REPAIR THE TISSUE DURING THE PROCEDURE. THERE WERE NO PIECES IN THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. UPON EVALUATION OF THE DEVICE, A FRACTURE WAS NOTED AT THE TIP OF THE NEEDLE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL6231 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. DEVICE EVALUATION SUMMARY: ONLY A PICTURE WAS RETURNED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PHOTO, TWO NEEDLES WITH BROKEN TIP COULD BE OBSERVED AND NO CONCLUSION COULD BE REACH. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DESCRIBE THE DEFECT OBSERVED AT THE TIP OF THE NEEDLE: ON ONE OF THE TIPS OF THE NEEDLE HCP CAN SEE HOW A DEFECT IN THE FINISH AS IF THE TIP HAD BEEN CUT. AND IN THE SECOND THE TIP HAD MELTED REMAINING AS A BALL. DESCRIBE IN DETAIL ¿THE DAMAGE CAUSED BY THE TIP IN THE TISSUE¿. THE TISSUE WAS SLIGHTLY TORN AT THE FIRST STEP OF THE NEEDLE. WHAT ACTION DID THE SURGEON TAKE TO CORRECT / CORRECT THE PROBLEM PRESENTED WITH THE DEVICE? HE IMMEDIATELY REMOVED THE NEEDLE AND PERFORMED TISSUE REPAIR WITH ANOTHER SUTURE. DEMOGRAPHIC DATA OF THE PATIENT (INITIALS BEGINNING WITH LAST NAME, AGE, GENDER): THE DATA IS NOT AVAILABLE. DOES THE DOCTOR CONSIDER THAT THE TOTAL SURGICAL PROCEDURE LASTED BEYOND WHAT WAS EXPECTED SUCH THAT THE ANESTHESIA DOSE HAD TO BE INCREASED / MODIFIED DUE TO THE PROBLEM WITH THE MEDICAL DEVICE AND THIS CAUSES DAMAGE / CONSEQUENCE IN THE PATIENT? (NO; YES, HOW LONG, CLINICAL EVOLUTION OF THE DAMAGE): AS IT IS PUT IN THE REPORT IT LASTED ABOUT 10 MORE MINUTES. WHAT IS THE CURRENT STATUS OF THE PATIENT? (FOR EXAMPLE: DISCHARGED WITHOUT CONSEQUENCES ARISING FROM THE PROBLEM WITH THE DEVICE, IN RECOVERY FROM SURGERY, CONTINUES TO BE HOSPITALIZED DUE TO¿ ETC.): THE DATA IS NOT AVAILABLE. SAMPLE A CAPTURED IN REPORT 2210968-2019-89260.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9260518&pc=GAM | 2210968-2019-89303 | 2019-09-12 | Malfunction | ETHICON INC. | 2019-10-31 | GAM | VICRYL* RAPIDE* VLT 45 6-0 P-1 | Break; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. AT THE MOMENT OF THE PROCEDURE, THE THREAD WAS BROKEN, DUE TO CONTACT WITH THE HAND AND THE THREAD WAS BROKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9262253&pc=GAM | 2210968-2019-89327 | 2019-10-04 | Malfunction | ETHICON INC. | 2019-10-31 | GAM | VICRYL VIO 75CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE AT THE SECOND PASSAGE DURING THE SUTURING. THERE WAS AN UNKNOWN DELAY OF THE PROCEDURE. NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ML8BBMR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THE NEEDLE PIECES REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? WAS THERE ANY CHANGE IN THE PATIENT POST OPERATIVE CARE DUE TO THE PROLONGED PROCEDURE? WHAT IS PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9256721&pc=GAM | 2210968-2019-89246 | 2019-10-04 | Malfunction | ETHICON INC. | 2019-10-30 | GAM | VICRYL RAP UND 90CM 2-0 V-37 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE SUTURE UNDERWENT AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE PULLED OFF FROM THE SUTURE STRAND DURING USE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE W9946, PEBCKDT0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9258071&pc=GAM | 2210968-2019-89260 | 2019-08-08 | Malfunction | ETHICON INC. | 2019-10-30 | GAM | VICRYL*VLT PLUS 3-0 70CM 1-SH | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A GASTRECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE POINT OF THE NEEDLE HAD DEFECT. THE SURGEON USED THE NEEDLE WHICH CAUSED DAMAGE TO THE TISSUE. THE PROCEDURE WAS DELAYED BY 15 MINUTES. ANOTHER SUTURE WAS USED TO REPAIR THE TISSUE DURING THE PROCEDURE. THERE WERE NO PIECES IN THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. UPON EVALUATION OF THE DEVICE, A FRACTURE WAS NOTED AT THE TIP OF THE NEEDLE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL5929 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. A SUTURE NEEDLE WAS SUBMITTED FOR EVALUATION. A FRACTURE WAS OBSERVED AT THE TIP OF SAMPLE A. ONE SIDE OF EACH NEEDLE WAS RECEIVED, THE MATING FRACTURE SURFACE WAS NOT PROVIDED FOR THIS EVALUATION. A SCANNING ELECTRON MICROSCOPE (SEM) WAS USED TO EXAMINE THE FRACTURE SURFACE AND SURROUNDING AREA OF THE NEEDLE. THE FRACTURE SURFACE WAS EXAMINED IN MULTIPLE LOCATIONS IN ORDER TO DETERMINE THE FRACTURE MODE. THE EVALUATION REVEALED THE FRACTURE WAS COMPOSED OF MICROVOID COALESCENCE, WHICH IS EVIDENCE OF A DUCTILE OVERLOAD FAILURE. MECHANICAL DAMAGE, A CUP-AND-CONE SHAPED FRACTURES AND MACROSCOPIC PLASTIC DEFORMATION OBSERVED COINCIDENTAL TO THE FRACTURE PROVIDE ADDITIONAL EVIDENCE THAT THE FAILURE WAS INDUCED BY MECHANICAL DEFORMATION LEADING TO DUCTILE OVERLOAD. THIS WAS A DUCTILE FRACTURE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE EVIDENCE OF THIS EXAMINATION INDICATES THAT THE BREAKAGE OCCURRED AT THE TIP OF THE NEEDLE DURING USE DUE TO TENSILE OVERLOAD. THERE IS NO EVIDENCE OF ANY MATERIAL FLAW OR DEFECT THAT WOULD CAUSE PREMATURE FAILURE. IN ORDER TO AVOID THIS KIND OF DAMAGE GRASP THE NEEDLE IN AN AREA ONE-THIRD (1/3) TO ONE-HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT. THE NEEDLE BREAKAGE WAS CONFIRMED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DESCRIBE THE DEFECT OBSERVED AT THE TIP OF THE NEEDLE: ON ONE OF THE TIPS OF THE NEEDLE HCP CAN SEE HOW A DEFECT IN THE FINISH AS IF THE TIP HAD BEEN CUT. AND IN THE SECOND THE TIP HAD MELTED REMAINING AS A BALL. DESCRIBE IN DETAIL ¿THE DAMAGE CAUSED BY THE TIP IN THE TISSUE¿. THE TISSUE WAS SLIGHTLY TORN AT THE FIRST STEP OF THE NEEDLE. WHAT ACTION DID THE SURGEON TAKE TO CORRECT / CORRECT THE PROBLEM PRESENTED WITH THE DEVICE? HE IMMEDIATELY REMOVED THE NEEDLE AND PERFORMED TISSUE REPAIR WITH ANOTHER SUTURE. DEMOGRAPHIC DATA OF THE PATIENT (INITIALS BEGINNING WITH LAST NAME, AGE, GENDER): THE DATA IS NOT AVAILABLE. DOES THE DOCTOR CONSIDER THAT THE TOTAL SURGICAL PROCEDURE LASTED BEYOND WHAT WAS EXPECTED SUCH THAT THE ANESTHESIA DOSE HAD TO BE INCREASED / MODIFIED DUE TO THE PROBLEM WITH THE MEDICAL DEVICE AND THIS CAUSES DAMAGE / CONSEQUENCE IN THE PATIENT? (NO; YES, HOW LONG, CLINICAL EVOLUTION OF THE DAMAGE): AS IT IS PUT IN THE REPORT IT LASTED ABOUT 10 MORE MINUTES. WHAT IS THE CURRENT STATUS OF THE PATIENT? (FOR EXAMPLE: DISCHARGED WITHOUT CONSEQUENCES ARISING FROM THE PROBLEM WITH THE DEVICE, IN RECOVERY FROM SURGERY, CONTINUES TO BE HOSPITALIZED DUE TO¿ ETC.): -THE DATA IS NOT AVAILABLE. NOTE: SAMPLE B CAPTURED IN REPORT 2210968-2019-89261.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9258115&pc=GAM | 2210968-2019-89263 | 2019-09-28 | Malfunction | ETHICON INC. | 2019-10-30 | GAM | VICRYL PLUS UND 70CM M3 USP2/0 SGLE ARMED PCLX PRIME | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE TIP OF THE NEEDLE WAS MISSING AFTER USE BY THE SURGEON. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PBK647 BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252052&pc=GAM | 2210968-2019-89196 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-29 | GAM | VICRYL CT UD 90CM M3 | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN ORTHOPEDIC PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE SUBCUTANEOUS SUTURE OF THE PATIENT HAD NOT RE-SORBED IN ITS ENTIRETY CAUSING A DELAY FOR THE SCARRING. THERE WAS NO TREATMENT REPORTED. THE PATIENT WOUND IS REPORTED AS HEALED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PROCEDURE NAME AND DATE: UNK. HOW MANY DAYS POST OP WAS THE ISSUE NOTED? UNK. HOW WAS THE DELAYED SUTURE ABSORPTION ISSUE MANAGED? IT WAS RETRIEVED. AFTER ABLATION OF THE SUTURE WHICH WAS NOT ABSORBED, THE WOUND USES TO SCAR ALONE MOST OF THE TIME. LOT NUMBER: UNK. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME AND DATE HOW MANY DAYS POST OP WAS THE ISSUE NOTED? HOW WAS THE DELAYED SUTURE ABSORPTION ISSUE MANAGED? EXPLAIN. LOT NUMBER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9244716&pc=GAM | 2210968-2019-89116 | 2019-09-01 | Malfunction | ETHICON INC. | 2019-10-28 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT C-SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE IN HALF. FRAGMENTS WERE RETRIEVED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9241668&pc=GAM | 2210968-2019-89064 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-25 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient; Device Embedded In Tissue or Plaque; Not Applicable | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. WHILE SUTURING THE UTERUS, THE TIP OF THE SUTURE NEEDLE BROKE OFF. THE PATIENT RECEIVED 3 X-RAYS TRYING TO LOCATE IT. THE NEEDLE WAS LOCATED, THEN SECOND X RAY IT MOVED, AND 3RD X RAY COULD NOT LOCATE. IT IS NOT KNOWN WHERE THE TIP IS BUT COULD NOT BE LOCATED ON 3RD X RAY. THE PATIENT ENDED UP WITH UTERUS OUTSIDE OF BODY FOR 5 HOURS. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: WEIGHT, AGE, BMI AT THE TIME OF INDEX PROCEDURE WHAT WAS THE DATE OF THE INITIAL PROCEDURE WITH DR. BARON? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? WHAT INSTRUMENTS WERE USED TO GRASP THE NEEDLE? WHERE WAS THE NEEDLE GRASPED DURING USE? WHERE WAS THE NEEDLE FRAGMENT LOCATED, IN WHAT TISSUE/STRUCTURE? WHAT PORTION OF THE NEEDLE TIP WAS RETAINED IN THE PATIENT (IN CM)? DOES THE NEEDLE FRAGMENT REMAIN RETAINED IN THE PATIENT? WERE ALL NEEDLE FRAGMENTS REMOVED DURING THE INITIAL PROCEDURE? WAS THERE ANY CHANGE TO THE PATIENTS POST OPERATIVE CARE DUE TO THE EXTENDED PROCEDURE? CAN YOU IDENTIFY THE PRODUCT CODE AND LOT NUMBER? WHAT IS THE SUTURE SIZE? WILL THE DEVICE INVOLVED IN THE EVENT BE RETURNING FOR EVALUATION? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? WHAT IS THE PATIENT¿S CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9218547&pc=GAM | 2210968-2019-88915 | 2019-09-11 | Malfunction | ETHICON INC. | 2019-10-22 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT CHILDBIRTH ON (B)(6) 2019 AND SUTURE WAS USED. DURING SUTURING ON THE PERINEAL INCISION SITE, THE SUTURE WAS BROKEN. ANOTHER PACKAGE WITH ANOTHER LOT NUMBER WAS USED WITH NO PROBLEM. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9219229&pc=GAM | 2210968-2019-88923 |  | Malfunction | ETHICON INC. | 2019-10-22 | GAM | VICRYL RAP UND 75CM M1.5 USP4/0 SGLE ARMED PC-3 PRIME | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT CHILDBIRTH ON AN UNKNOWN DATE AND SUTURE WAS USED ON THE PERINEUM. DURING THE PROCEDURE, WHEN SUTURING THE PERINEUM, THE NEEDLE CAME AWAY FROM THE SUTURE . THE SURGEON WAS UNABLE TO LOCATE THE NEEDLE IN SOFT TISSUE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL9625 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 10/22/2019. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE DETACHED NEEDLE RETRIEVED? WAS IT CONFIRMED THAT THE NEEDLE WAS NOT RETAINED IN THE PATIENT? WERE X-RAYS/RADIOLOGICAL TEST PERFORMED TO ENSURE THE NEEDLE IS NOT RETAINED IN THE PATIENT? IF NEEDLE IS RETAINED IN PATIENT- PLEASE PROVIDE LOCATION AND SIZE OF THE NEEDLE WHERE IT IS RETAINED. IS ANY SURGERY SCHEDULED TO REMOVE IT FROM THE PATIENT? WERE THERE ANY OTHER ADVERSE PATIENT CONSEQUENCES REPORTED AND IF YES, WHAT MEDICAL SURGICAL INTERVENTION WAS PROVIDED? PATIENT'S CURRENT CONDITION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9208642&pc=GAM | 2210968-2019-88517 |  | Malfunction | ETHICON INC. | 2019-10-18 | GAM | VICRYL CT VIO 90CM M5 | Inaccurate Information | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT PRODUCT WAS NOT USED ON PATIENT. THE PRODUCT IS LABELED 'NOT FOR HUMAN USE'. NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE, WHICH WAS RAISED TO ADDRESS THE EVENT REPORTED. THE NECESSARY ACTIONS HAVE BEEN TAKEN TO ADDRESS THE ISSUE. SUMMARY SUPPLEMENT: ONE SEALED BOX OF PRODUCT CODE V372H, LOT PHBBCJQ0 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE BOX, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. Manufacturer Narrative: (B)(4). EVALUATION: ONE SEALED BOX OF PRODUCT CODE (B)(4)., LOT PHBBCJQ0 WERE RETURNED FOR ANALYSIS. DURING VISUAL INSPECTION OF THE BOX, THE PRINTED LEGEND OF NOT FOR HUMAN USE WAS OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE WHICH WAS RAISED TO ADDRESS. THE NECESSARY ACTIONS WILL BE TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM ACCORDING TO THE SAMPLE CONDITION, THE REPORTED COMPLAINT WAS OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE NR- (B)(4). WHICH WAS RAISED TO ADDRESS THE EVENT REPORTED. IF ANY FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9204110&pc=GAM | 2210968-2019-88810 | 2019-09-01 | Malfunction | ETHICON INC. | 2019-10-17 | GAM | VICRYL*UND 4-0 45CM (1)P-3 PR | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE AND NEEDLE DISASSEMBLED AND IT WAS NECESSARY TO REQUEST FOR MORE SUTURES. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NUMBER 0839501001 IS INVALID FOR A LOT NUMBER OF THE REPORTED PRODUCT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9204123&pc=GAM | 2210968-2019-88809 | 2019-09-01 | Malfunction | ETHICON INC. | 2019-10-17 | GAM | VICRYL*UND 4-0 45CM (1)P-3 PR | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE AND NEEDLE DISASSEMBLED AND IT WAS NECESSARY TO REQUEST FOR MORE SUTURES. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NUMBER 0839501001 IS INVALID FOR A LOT NUMBER OF THE REPORTED PRODUCT. DID THE NEEDLES PULLED OFF THE SUTURE THREADS OR NEEDLES BROKE? DID BOTH NEEDLES HAD ISSUE IN THIS PROCEDURE? LOT NUMBER, DEVICE RETURN STATUS / FOLLOW UP?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9199167&pc=GAM | 2210968-2019-88747 | 2019-09-16 | Malfunction | ETHICON INC. | 2019-10-16 | GAM | VICRYL*UND 4-0 45CM (1)P-3 PR | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE IS DISASSEMBLED WHEN USED. A DIFFERENT DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE PROCEDURE COMPLETED SUCCESSFULLY? WAS THERE ANY ADVERSE CONSEQUENCE ASSOCIATED WITH THE PATIENT? NAME OF THE PROCEDURE? WHAT IS THE PATIENT'S CURRENT STATUS? HOW MANY DEVICES HAD ¿THE NEEDLE IS DISASSEMBLED¿ ISSUE IN THIS PROCEDURE? DID THE NEEDLE PULL OFF OR BROKE? PRODUCT RETURN STATUS / FOLLOW UP? THERE IS NO PHYSICAL OR PHOTOGRAPHIC EVIDENCE OF THE DEVICE. STEEL SUTURE CAPTURED IN MW 2210968-2019-88746.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9199170&pc=GAM | 2210968-2019-88745 | 2019-08-01 | Malfunction | ETHICON INC. | 2019-10-16 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED VIA USER FACILITY MEDWATCH FORM # 2200710000-2019-8050 THAT THE PATIENT UNDERWENT LAPAROSCOPIC TOTAL HYSTERECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. BILATERAL SALPINGO-OOPHORECTOMY, BILATERAL PELVIC SENTINEL LYMPH NODE MAPPING WITH PINPOINT, LYSIS OF ADHESION WAS ALSO PERFORMED DURING THE PROCEDURE. ABOUT 2/3 LENGTH OF 0-UR6 SUTURE NEEDLE TIP BROKE OFF INSIDE PATIENT'S FASCIA WHILE SUTURING. PORTABLE XRAY OF ABDOMEN CALLED TO AID FOR LOCATION OF SUTURE IN ABDOMINAL TISSUE. THE SURGEONS WERE ABLE TO REMOVE ALL MISSING SUTURE NEEDLE IN ONE PIECE. THE RADIOLOGIST WAS CALLED BACK INTO THE ROOM FOR NEGATIVE RESULT AFTER AND CONFIRMED NO NEEDLE FRAGMENTS WERE REMAINING WHEN THE X-RAY WAS COMPLETED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED AND IT WAS REPORTED THE EVENT HAD NO IMPACT ON THE TRAJECTORY OF THE PATIENT¿S CARE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PROCEDURE(S): LAPAROSCOPIC TOTAL HYSTERECTOMY, BILATERAL SALPINGO-OOPHORECTOMY, BILATERAL PELVIC SENTINEL LYMPH NODE MAPPING WITH PINPOINT, LYSIS OF ADHESION. PRODUCT CODE AND LOT NUMBER OF VICRYL SUTURE. LOT # IS UNKNOWN. WAS THE PROCEDURE COMPLETED SUCCESSFULLY? THE PROCEDURE WAS NOT IMPACTED BY THIS EVENT. WHAT IS MEANT BY ¿RADIOLOGIST CALLED BACK ROOM FOR NEGATIVE RESULT AFTER¿? THE RADIOLOGIST CONFIRMED THAT THERE WAS NO NEEDLE FRAGMENTS REMAINING WHEN THE X-RAY WAS COMPLETED. WAS ANY PIECE OF THE NEEDLE LEFT IN THE PATIENT? NO PIECE OF THE NEEDLE WAS LEFT IN THE PATIENT. WAS THERE ANY ADVERSE CONSEQUENCE ASSOCIATED WITH THE PATIENT? NOTHING ADVERSE IMPACTING THE PATIENT. WHAT IS THE PATIENT'S CURRENT STATUS? SUCCESSFUL OUTCOME TO THE SURGERY ¿ THIS BROKEN NEEDLE EVENT HAD NO IMPACT ON THE TRAJECTORY OF THE PATIENT¿S CARE. THIS MW IS RELATED TO USER FACILITY MEDWATCH FORM # 2200710000-2019-8050.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9192235&pc=GAM | 2210968-2019-88660 | 2019-06-05 | Malfunction | ETHICON INC. | 2019-10-15 | GAM | VICRYL* UND 3-0 70CM (1)SC-20 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ABDOMINOPLASTY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE OF THE SUTURE WAS DISASSEMBLED AT THE MOMENT THE SURGEON DRAGS THE SUTURE ON THE FIRST PASS THROUGH THE TISSUE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9179828&pc=GAM | 2210968-2019-88568 | 2019-09-16 | Malfunction | ETHICON INC. | 2019-10-11 | GAM | VICRYL RAP UND 75CM 5-0 RB-1 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BREAKDOWN FROM SWAGE POINT. IT WAS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9179866&pc=GAM | 2210968-2019-88569 | 2019-09-16 | Malfunction | ETHICON INC. | 2019-10-11 | GAM | VICRYL RAP UND 75CM 5-0 RB-1 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BREAKDOWN FROM SWAGE POINT. IT WAS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9181455&pc=GAM | 2210968-2019-88582 | 2019-09-16 | Malfunction | ETHICON INC. | 2019-10-11 | GAM | VICRYL CT VIO 70CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT LOWER RIGHT LOBECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING SUTURE OF THE DORSAL MUSCLE, AT THIRD STEP, THE RUPTURE OF THE NEEDLE INSIDE MUSCLE OCCURRED. THE SURGEON RECOVERED THE BROKEN NEEDLE. ANOTHER SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 10/11/2019. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V318H, PABDQHQ00 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9173758&pc=GAM | 2210968-2019-88528 | 2019-09-05 | Malfunction | ETHICON INC. | 2019-10-10 | GAM | VICRYL ABS SUR SUT USP SYN 4-0 70CM | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE DURING USE. NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING / PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. TO DATE NO SAMPLE HAS BEEN RETURNED. IF PRODUCT IS RETURNED OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9174801&pc=GAM | 2210968-2019-88530 | 2019-09-05 | Malfunction | ETHICON INC. | 2019-10-10 | GAM | VICRYL ABS SUR SUT USP SYN 4-0 70CM | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE DURING USE. NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9175955&pc=GAM | 2210968-2019-88534 | 2019-09-12 | Malfunction | ETHICON INC. | 2019-10-10 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING SURGERY THE SUTURE GOT DETACHED FROM THE NEEDLE AT THE SWAGED POINT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS ANY ADVERSE PATIENT CONSEQUENCES AND WHAT MEDICAL / SURGICAL INTERVENTION WAS PROVIDED TO MANGE THEM?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9177300&pc=GAM | 2210968-2019-88553 | 2019-01-20 | Malfunction | ETHICON INC. | 2019-10-10 | GAM | VICRYL SUTURE 27" (70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A COLECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS UNWOUND. THE NEEDLE WAS LOST INSIDE THE BODY, BUT THEN FOUND AND REMOVED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AK9777 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INVESTIGATION SUMMARY: ONE SEALED BOX WITH THIRTY-SIX SAMPLES AND UNOPENED SAMPLES OF PRODUCT CODE: J318, LOT: AK9977 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. ADDITIONAL INVESTIGATION SUMMARY: ONE EMPTY OPENED FOIL, A DETACHED NEEDLE AND DISPENSED SUTURE OF PRODUCT CODE: J318, LOT: AK9777 WERE RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE AS EXPECTED. THE BARREL HOLE WAS EXAMINED UNDER 20X MAGNIFICATION AND REMNANT SUTURE WAS NOTED. THE SUTURE END IS SEVERELY CUT, RESULTING IN A CLIP OFF DEFECT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE SAMPLE CONDITION THE ASSIGNABLE CAUSE OF THE NEEDLE PULL OFF, IS A CLIP OFF DEFECT, THIS DEFECT IS CAUSED BY EXCESSIVE PRESSURE ON THE SUTURE DURING SWAGING OF THE NEEDLE AND CONSEQUENTLY BREAKING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9177302&pc=GAM | 2210968-2019-88554 | 2019-01-20 | Malfunction | ETHICON INC. | 2019-10-10 | GAM | VICRYL SUTURE 27" (70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A COLECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS UNWOUND. THE NEEDLE WAS LOST INSIDE THE BODY, BUT THEN FOUND AND REMOVED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AK9777 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: ONE SEALED BOX WITH THIRTY-SIX SAMPLES AND UNOPENED SAMPLES OF PRODUCT CODE J318, LOT AK9977 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. INVESTIGATION SUMMARY: TWO EMPTY OPENED FOILS, A NEEDLE-SUTURE IN TWO SECTION OF PRODUCT CODE J318, LOT AK9777 WERE RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE (NEEDLE/SUTURE PIECE), THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED; BUT THE END OF THE SUTURE WAS OBSERVED CUT APPEARS TO BE BY USE OF A SURGICAL INSTRUMENT. THE OTHER SECTION OF THE SUTURE WAS EXAMINED, ONE END OF SUTURE WAS CUT AND IT MATCHED WITH THE EXTREME OF THE NEEDLE/SUTURE PIECE. PER THE SAMPLE CONDITION, NO FUNCTIONAL TEST COULD BE PERFORMED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO CONDITION OF THE SAMPLE, NO PULL OFF - SUTURE NEEDLE WAS FOUND ON THE SAMPLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9172411&pc=GAM | 2210968-2019-88509 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-09 | GAM | VICRYL CT UD 90CM M3 | Appropriate Term/Code Not Available | Scarring | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT OSTEOSYNTHESIS CLAVICLE PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. APPROXIMATELY 60 DAYS POST PROCEDURE, THE NURSE IN CHARGE OF THE PATIENT AT HOME INFORMED THAT THE SUTURE OF THE PATIENT HAD NOT RESORBED IN ITS ENTIRETY THUS CAUSING A DELAY FOR THE SCARRING. THERE WAS NO TREATMENT INDICATED FOR THE SCARRING. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PROCEDURE NAME: OSTEOSYNTHESIS CLAVICLE. THE AMOUNT OF TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION¿? THE TIME BETWEEN THE SUTURE PLACEMENT AND THE FINDING THE PROBLEM IS OF 60 DAYS. THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE "ABSORPTION"? UNKNOWN. DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING OR AFTER ITS USE IN THE PATIENT? YES, THE QUALITY OF THE THREAD WAS DIFFERENT AND MORE BRITTLE WHEN MAKING THE KNOT. REPORTED EVENT OF BRITTLE SUTURE CAPTURED IN MW 2210968-2019-88510.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9172942&pc=GAM | 2210968-2019-88516 |  | Malfunction | ETHICON INC. | 2019-10-09 | GAM | VICRYL CT VIO 90CM M5 | Inaccurate Information | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT PRODUCT WAS NOT USED ON PATIENT. THE PRODUCT IS LABELED 'NOT FOR HUMAN USE'. NO ADVERSE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). EVALUATION: ONE SEALED BOX OF PRODUCT CODE (B)(4), LOT PHBBCJQ0 WERE RETURNED FOR ANALYSIS. DURING VISUAL INSPECTION OF THE BOX, THE PRINTED LEGEND OF NOT FOR HUMAN USE WAS OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE WHICH WAS RAISED TO ADDRESS. THE NECESSARY ACTIONS WILL BE TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM ACCORDING TO THE SAMPLE CONDITION, THE REPORTED COMPLAINT WAS OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE NR-(B)(4) WHICH WAS RAISED TO ADDRESS THE EVENT REPORTED. IF ANY FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9173742&pc=GAM | 2210968-2019-88525 | 2019-09-06 | Malfunction | ETHICON INC. | 2019-10-09 | GAM | VICRYL ABS SUR SUT USP SYN 4-0 70CM | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE DURING USE. NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING / PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. TO DATE NO SAMPLE HAS BEEN RETURNED. IF PRODUCT IS RETURNED OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9173743&pc=GAM | 2210968-2019-88526 | 2019-09-06 | Malfunction | ETHICON INC. | 2019-10-09 | GAM | VICRYL ABS SUR SUT USP SYN 4-0 70CM | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PLASTIC SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE SUTURE DURING USE. NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4); 01/31/2023. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. TO DATE NO SAMPLE HAS BEEN RETURNED. IF PRODUCT IS RETURNED OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9167456&pc=GAM | 2210968-2019-88458 | 2019-05-28 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ELECTIVE LOWER SEGMENT CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. WHEN THE SURGEON GAVE THE NEEDLE BACK AFTER CLOSING THE UTERUS, THEY FOUND THE TIP OF THE NEEDLE WAS MISSING (APPROX.0.5 MM). INFORMED THE SURGEON, AND THEATRE FLOOR PERSON. ALSO X-RAYED THE PATIENT ON THE OPERATING TABLE WHICH SHOWED NO NEEDLE TIP ON THE PATIENT. THE SURGEON INFORMED THE PATIENT AS WELL. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 10/08/2019. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IS THE PRODUCT BEING RETURNED FOR EVALUATION? ACTUAL BROKEN NEEDLE OR REPRESENTATIVE SAMPLE? DEVICE RETURN STATUS / FOLLOW UP.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9167904&pc=GAM | 2210968-2019-88469 | 2019-07-24 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL RAP UND 70CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A UROLOGY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE CAME LOOSE FROM THE NEEDLE. A NEW SUTURE WAS TAKEN FROM A NEW BOX FOR USE. THERE WERE NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS TO RETRIEVE THE DEVICE WERE MADE. TO DATE NO SAMPLE HAS BEEN RETURNED. IF PRODUCT IS RETURNED OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168399&pc=GAM | 2210968-2019-88476 | 2019-01-04 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ELECTIVE CESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. ONE OF THE SURGEONS OPINED THAT THE SUTURE NEEDLE WAS BLUNT. AFTER INSPECTING THE NEEDLE IT WAS FOUND THAT THE NEEDLE LOST ITS TIP. X-RAY WAS DONE AND NO NEEDLE WAS VISIBLE IN THE PATIENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE VERIFY, DID THE DEVICE USED IN THIS PROCEDURE HAVE DULL NEEDLE OR THE NEEDLE TIP BROKE DURING USE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168548&pc=GAM | 2210968-2019-88477 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL CT VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND SUTURE WAS USED. THE NEEDLE BROKE INTRA-OP AND FELL INTO SITUS. THIS LED TO 60 MINUTES DELAY IN SURGERY TIME. THE NEEDLE PARTS WERE RETRIEVED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: HOW WERE THE NEEDLE PARTS RETRIEVED? WAS ANY ADDITIONAL SURGICAL INTERVENTION /RE-OPENING DONE ON PATIENT TO REMOVE THE NEEDLE PIECES? WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN. PATIENT CURRENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168921&pc=GAM | 2210968-2019-88479 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE SEPARATE FROM THE SUTURE STRAND WHILE THE PHYSICIAN WAS SUTURING THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168924&pc=GAM | 2210968-2019-88478 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE SEPARATE FROM THE SUTURE STRAND WHILE THE PHYSICIAN WAS SUTURING THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168925&pc=GAM | 2210968-2019-88480 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE SEPARATE FROM THE SUTURE STRAND WHILE THE PHYSICIAN WAS SUTURING THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168926&pc=GAM | 2210968-2019-88482 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE SEPARATE FROM THE SUTURE STRAND WHILE THE PHYSICIAN WAS SUTURING THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168928&pc=GAM | 2210968-2019-88481 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE SEPARATE FROM THE SUTURE STRAND WHILE THE PHYSICIAN WAS SUTURING THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168929&pc=GAM | 2210968-2019-88483 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE SEPARATE FROM THE SUTURE STRAND WHILE THE PHYSICIAN WAS SUTURING THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168933&pc=GAM | 2210968-2019-88485 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE SEPARATE FROM THE SUTURE STRAND WHILE THE PHYSICIAN WAS SUTURING THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9168939&pc=GAM | 2210968-2019-88484 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE SEPARATE FROM THE SUTURE STRAND WHILE THE PHYSICIAN WAS SUTURING THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9151608&pc=GAM | 2210968-2019-88328 | 2019-09-04 | Malfunction | ETHICON INC. | 2019-10-03 | GAM | VICRYL RAP UND 45CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE FELL OFF DURING SUTURING. THEY HAD TO SEARCH FOR THE NEEDLE FOR 10 MINUTES WITHIN PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE ACTUAL DEVICE BATCH NUMBER ASSOCIATED WITH THIS EVENT IS NOT KNOWN. TWO POSSIBLE BATCH NUMBERS ARE REPORTED AS FOLLOWS: BATCH AL7127 MFG DATE: UNK, EXP DATE: UNK, BATCH AK3530 MFG DATE: 03/09/2018, EXP DATE: 02/28/2023. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AK3530 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9151742&pc=GAM | 2210968-2019-88329 | 2019-08-01 | Malfunction | ETHICON INC. | 2019-10-03 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CCV PROCEDURE ON AN UNKNOWN DATE IN 2019 AND SUTURE WAS USED. IN (B)(6) THE PATIENT PRESENTED WITH DEHISCENCE. THE LESION IMPROVED AFTER THE SUTURE WAS REMOVED. IT WAS NOTED THAT THE SUTURE WAS COMPLETELY REMOVED FROM THE LESION, THEREFORE, IT'S POSSIBLE TO SEE THAT IT IS NOT BEING ABSORBED. THERE WAS NO INTERVENTION REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE NAME OF THE INITIAL PROCEDURE? WHAT WAS THE DATE OF THE INITIAL PROCEDURE? ON WHAT TISSUE WAS THE SUTURE USED? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? WHAT WAS THE ONSET DATE OF SYMPTOMS FROM THE INITIAL SURGICAL PROCEDURE? HOW MANY DAYS POST-OPERATIVELY WAS THE SUTURE REMOVED? WERE THE SUTURES REMOVED EARLIER THAN THE NORMAL ROUTINE REMOVAL TIME FRAME? WAS THE PATIENT'S POST- OPERATIVE CARE ALTERED IN ANY WAY AS A RESULT OF THIS EVENT? WHAT WAS DONE IMMEDIATELY UPON SUTURE REMOVAL? WAS THERE ANY TREATMENT? CAN YOU IDENTIFY THE PRODUCT CODE AND LOT # ? HOW WAS THE DEHISCENCE MANAGED / TREATED? WAS ANY RE-OPERATION/RE-SUTURING / REVISION SURGERY DONE ON PATIENT POST-OP INITIAL PROCEDURE? THE AMOUNT OF TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION? IF APPLICABLE, WILL PRODUCT BE RETURNED, RETURN DATE, TRACKING INFORMATION? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? WHAT IS THE PATIENT¿S CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9146982&pc=GAM | 2210968-2019-88296 |  | Malfunction | ETHICON INC. | 2019-10-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OP, THE PATIENT PRESENTED WITH A SUTURE DEHISCENCE. THE LESIONS IMPROVED AFTER THE SUTURE WAS REMOVED AND THE ENTIRE SUTURE EXITS, IE, IT IS NOT BEING ABSORBED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW WAS THE DEHISCENCE MANAGED? IF A SURGICAL INTERVENTION WAS PERFORMED, PROVIDE ADVISE OF DATE AND SPECIFICS. PLEASE PROVIDE THE LOT NUMBER REPORTED IN THIS EVENT? HOW IS THE PATIENT TODAY?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9146987&pc=GAM | 2210968-2019-88305 |  | Malfunction | ETHICON INC. | 2019-10-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OP, THE PATIENT PRESENTED WITH A SUTURE DEHISCENCE. THE LESIONS IMPROVED AFTER THE SUTURE WAS REMOVED FROM THE LESION, THEREFORE, IT WAS POSSIBLE TO SEE WHAT IS IT WAS NOT BEING ABSORBED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE NAME OF THE INITIAL PROCEDURE? WHAT WAS THE DATE OF THE INITIAL PROCEDURE? ON WHAT TISSUE WAS THE SUTURE USED? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? WHAT WAS THE ONSET DATE OF SYMPTOMS FROM THE INITIAL SURGICAL PROCEDURE? HOW MANY DAYS POST-OPERATIVELY WAS THE SUTURE REMOVED? WERE THE SUTURES REMOVED EARLIER THAN THE NORMAL ROUTINE REMOVAL TIME FRAME? WAS THE PATIENT'S POST- OPERATIVE CARE ALTERED IN ANY WAY AS A RESULT OF THIS EVENT? WHAT WAS DONE IMMEDIATELY UPON SUTURE REMOVAL? WAS THERE ANY TREATMENT? CAN YOU IDENTIFY THE PRODUCT CODE AND LOT #? HOW WAS THE DEHISCENCE MANAGED/TREATED? WAS ANY RE-OPERATION/RE-SUTURING/ REVISION SURGERY DONE ON PATIENT POST-OP INITIAL PROCEDURE? THE AMOUNT OF TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION? IF APPLICABLE, WILL PRODUCT BE RETURNED, RETURN DATE, TRACKING INFORMATION? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? WHAT IS THE PATIENT¿S CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9146988&pc=GAM | 2210968-2019-88306 |  | Malfunction | ETHICON INC. | 2019-10-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OP, THE PATIENT PRESENTED WITH A SUTURE DEHISCENCE ONE DAY AFTER THE PROCEDURE. THE LESIONS IMPROVED AFTER THE SUTURE WAS REMOVED AND THE ENTIRE SUTURE EXITS, IE, IT IS NOT BEING ABSORBED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. LOT #? WHAT DATE DID THE PATIENT PRESENT WITH DEHISCENCE (# POST OP DAYS)? WAS ANY SUTURE BREAKAGE NOTICED POST-OP? HOW WAS THE DEHISCENCE MANAGED/TREATED? WAS ANY MEDICAL / SURGICAL INTERVENTION REQUIRED TO TREAT THE PATIENT CONDITION? WAS ANY RE-OPERATION/RE-SUTURING/ REVISION SURGERY DONE ON PATIENT POST-OP INITIAL PROCEDURE? THE AMOUNT OF TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9146991&pc=GAM | 2210968-2019-88304 |  | Malfunction | ETHICON INC. | 2019-10-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POST-OP, THE PATIENT PRESENTED WITH A SUTURE DEHISCENCE. THE LESIONS IMPROVED AFTER THE SUTURE WAS REMOVED FROM THE LESION, THEREFORE, IT WAS POSSIBLE TO SEE WHAT IS IT WAS NOT BEING ABSORBED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE NAME OF THE INITIAL PROCEDURE? WHAT WAS THE DATE OF THE INITIAL PROCEDURE? ON WHAT TISSUE WAS THE SUTURE USED? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? WHAT WAS THE ONSET DATE OF SYMPTOMS FROM THE INITIAL SURGICAL PROCEDURE? HOW MANY DAYS POST-OPERATIVELY WAS THE SUTURE REMOVED? WERE THE SUTURES REMOVED EARLIER THAN THE NORMAL ROUTINE REMOVAL TIME FRAME? WAS THE PATIENT'S POST- OPERATIVE CARE ALTERED IN ANY WAY AS A RESULT OF THIS EVENT? WHAT WAS DONE IMMEDIATELY UPON SUTURE REMOVAL? WAS THERE ANY TREATMENT? CAN YOU IDENTIFY THE PRODUCT CODE AND LOT #? HOW WAS THE DEHISCENCE MANAGED/TREATED? WAS ANY RE-OPERATION/RE-SUTURING/ REVISION SURGERY DONE ON PATIENT POST-OP INITIAL PROCEDURE? THE AMOUNT OF TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION? IF APPLICABLE, WILL PRODUCT BE RETURNED, RETURN DATE, TRACKING INFORMATION? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? WHAT IS THE PATIENT¿S CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147019&pc=GAM | 2210968-2019-88310 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-10-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT PLASTIC SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE PATIENT PRESENTED WITH SUTURE DEHISCENCE AFTER USE OF THE DEVICE. THE LESIONS IMPROVED AFTER THE SUTURE WAS REMOVED AND ENTIRE SUTURE EXITS. THE SUTURE WAS NOT BEING ABSORBED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. LOT #? WHAT DATE DID THE PATIENT PRESENT WITH DEHISCENCE (# POST OP DAYS)? WAS ANY SUTURE BREAKAGE NOTICED POST-OP? HOW WAS THE DEHISCENCE MANAGED/TREATED? WAS ANY MEDICAL / SURGICAL INTERVENTION REQUIRED TO TREAT THE PATIENT CONDITION? WAS ANY RE-OPERATION/RE-SUTURING/ REVISION SURGERY DONE ON PATIENT POST-OP INITIAL PROCEDURE? THE AMOUNT OF TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION? THE SURGEON EXPECTED TIME BETWEEN THE SUTURE PLACEMENT AND THE NON ABSORPTION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9139458&pc=GAM | 2210968-2019-88258 | 2019-09-04 | Malfunction | ETHICON INC. | 2019-10-01 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN APPENDECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE WHILE SUTURING. A DIFFERENT SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9139459&pc=GAM | 2210968-2019-88257 | 2019-09-04 | Malfunction | ETHICON INC. | 2019-10-01 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN APPENDECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE WHILE SUTURING. A DIFFERENT SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9139537&pc=GAM | 2210968-2019-88259 | 2019-09-07 | Malfunction | ETHICON INC. | 2019-10-01 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN APPENDECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE DETACHED FROM THE NEEDLE WHILE SUTURING. ANOTHER SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136798&pc=GAM | 2210968-2019-88241 | 2019-09-04 | Malfunction | ETHICON INC. | 2019-09-30 | GAM | VICRYL CT VIO 75CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN EAR NOSE AND THROAT PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE, RESULTING IN A DETACHED NEEDLE DURING THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 11/13/2019. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV398, PEBDKHR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL H3 INVESTIGATION SUMMARY: ONLY PICTURES WERE RETURNED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, A BROKEN SUTURE AND SUTURE INSIDE OF THE BARREL HOLE COULD BE OBSERVED. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR EVALUATION. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. UPON THIS MEDWATCH BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131460&pc=GAM | 2210968-2019-88199 |  | Malfunction | ETHICON INC. | 2019-09-27 | GAM | VICRYL PLUS CT UD 18IN(45CM) USP5-0(M1) S/A PC-3 PRIME | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE STRAND DETACHED FROM SWAGE OF THE SURGICAL NEEDLE. THERE WAS NO ADVERSE PATIENT OUTCOME REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125697&pc=GAM | 2210968-2019-88156 | 2019-09-02 | Malfunction | ETHICON INC. | 2019-09-26 | GAM | VICRYL SUTURE 27" (70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE FIRST PASS THROUGH THE TISSUE IT WAS NOTED THAT THE NEEDLE HAD COME AWAY FROM THE SUTURE MATERIAL, AT THE CRIMP POINT. THE ENTIRE NEEDLE SECTION WAS RETRIEVED FOLLOWING X-RAY. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9117766&pc=GAM | 2210968-2019-88089 | 2019-08-22 | Malfunction | ETHICON INC. | 2019-09-25 | GAM | VICRYL RAP UND 90CM M3.5 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN ABDOMINAL SURGICAL PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED FOR SUBCUTANEOUS SUTURING. AT THE END OF THE PROCEDURE, IT WAS NOTICED THAT, AFTER THE THIRD POINT, THE NEEDLE TIP OF THE SUTURE WAS MISSING (2MM). THE WOUND AND SURROUNDING TISSUES WERE INSPECTED, WITHOUT IDENTIFICATION OF THE MISSING TIP. THE RADIOLOGY TECHNICIAN WAS ALERTED AND A CONTROL X-RAY WAS PERFORMED WITHOUT EVIDENCE OF FOREIGN BODIES. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. AFTER AWAKENING OF THE PATIENT, THERE WAS AN X-RAY CHECK IN TWO PROJECTIONS OF THE ABDOMINAL WALL: NO FOREIGN BODIES WERE FOUND. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? NO. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9118580&pc=GAM | 2210968-2019-88100 |  | Malfunction | ETHICON INC. | 2019-09-25 | GAM | VICRYL RAP UND 75CM M1.5 USP4/0 SGLE ARMED PC-3 PRIME | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND THE SUTURE WAS USED. DURING SURGERY, THE SUTURE WAS BREAKING, THE THREAD BREAKS FROM THE NEEDLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AJ5620 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. CLARIFY THE EVENT. DID THE NEEDLE PULL OFF THE SUTURE THREAD OR SUTURE THREAD BROKE? DID 2 DEVICES HAD ISSUE IN THIS SINGLE PROCEDURE? DEVICE RETURN STATUS / FOLLOW UP.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9118753&pc=GAM | 2210968-2019-88104 |  | Malfunction | ETHICON INC. | 2019-09-25 | GAM | VICRYL RAP UND 75CM M1.5 USP4/0 SGLE ARMED PC-3 PRIME | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND THE SUTURE WAS USED. DURING SURGERY, THE SUTURE WAS BREAKING, THE THREAD BREAKS FROM THE NEEDLE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AJ5620 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. CLARIFY THE EVENT. DID THE NEEDLE PULL OFF THE SUTURE THREAD OR SUTURE THREAD BROKE? DID 2 DEVICES HAD ISSUE IN THIS SINGLE PROCEDURE? DEVICE RETURN STATUS/ FOLLOW UP.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9119410&pc=GAM | 2210968-2019-88114 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-25 | GAM | VICRYL CT VIO 90CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED, THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9119440&pc=GAM | 2210968-2019-88115 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-25 | GAM | VICRYL UND 20CM M0.7 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED, THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9111735&pc=GAM | 2210968-2019-88045 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL UND 20CM M0.7 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING CAESAREAN SECTION, THE SUTURE BROKE SEVERAL TIMES. NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT PROLONGATION OF THE DURATION OF SURGERY. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9111763&pc=GAM | 2210968-2019-88046 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL CT VIO 90CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED. THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9111767&pc=GAM | 2210968-2019-88044 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL CT VIO 90CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED. THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9111783&pc=GAM | 2210968-2019-88047 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL CT VIO 90CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED. THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9111833&pc=GAM | 2210968-2019-88048 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL CT VIO 90CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED. THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED.. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9111857&pc=GAM | 2210968-2019-88049 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL CT VIO 90CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED. THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9111894&pc=GAM | 2210968-2019-88050 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL UND 20CM M0.7 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED. THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9111951&pc=GAM | 2210968-2019-88051 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL UND 20CM M0.7 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED. THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112019&pc=GAM | 2210968-2019-88052 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL UND 20CM M0.7 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED. THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112039&pc=GAM | 2210968-2019-88053 | 2019-08-12 | Malfunction | ETHICON INC. | 2019-09-24 | GAM | VICRYL UND 20CM M0.7 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE PROCEDURE. THE SURGEON NEED TO SUTURE THE UTERINE BRECCIA SEVERAL TIMES WITH A CONSEQUENT UNKNOWN PROLONGATION OF THE DURATION OF SURGERY. THE PROCEDURE WAS QUITE COMPLICATED. THE PATIENT BLED MORE FOR DIFFERENT PASSAGES OF THE NEEDLE TO THE UTERUS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9084002&pc=GAM | 2210968-2019-87829 | 2019-05-24 | Malfunction | ETHICON INC. | 2019-09-18 | GAM | VICRYL*RAPIDE*UND 3-0 75CM V-4 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE HAD A LOOSE THREAD. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHEN DID THE NEEDLE SEPARATE FROM THE SUTURE THREAD; IN THE PACKAGE, DURING DISPENSING (BEFORE USE ON PATIENT) OR DURING ACTUAL SUTURING (USE ON PATIENT)? DURING ACTUAL SUTURING, LOT NUMBER, HOW WAS CASE COMPLETED? ANOTHER THREAD HAS BEEN OPENED. WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? UNKNOWN, WAS PROCEDURE SUCCESSFULLY COMPLETED? UNKNOWN, WERE FRAGMENTS GENERATED? UNKNOWN, IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES :NO, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN,
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9084004&pc=GAM | 2210968-2019-87827 | 2019-05-24 | Malfunction | ETHICON INC. | 2019-09-18 | GAM | VICRYL*RAPIDE*UND 3-0 75CM V-4 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE HAD A LOOSE THREAD. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHEN DID THE NEEDLE SEPARATE FROM THE SUTURE THREAD; IN THE PACKAGE, DURING DISPENSING (BEFORE USE ON PATIENT) OR DURING ACTUAL SUTURING (USE ON PATIENT)? DURING ACTUAL SUTURING. LOT NUMBER. HOW WAS CASE COMPLETED? ANOTHER THREAD HAS BEEN OPENED. WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? UNKNOWN. WAS PROCEDURE SUCCESSFULLY COMPLETED? UNKNOWN. WERE FRAGMENTS GENERATED? UNKNOWN. IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? UNKNOWN. PATIENT STATUS/ OUTCOME / CONSEQUENCES: NO. WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9084082&pc=GAM | 2210968-2019-87835 | 2019-07-11 | Malfunction | ETHICON INC. | 2019-09-18 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING THE PROCEDURE. IMMEDIATELY CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT# (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9085080&pc=GAM | 2210968-2019-87837 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-09-18 | GAM | VICRYL RAP UND 70CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SCROTUM SEAM PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DOES NOT SEEM TO BE SHEATHED AS USUAL, THEY ARE TOO EASY TO PULL OFF THE THREAD. THE SURGERY WAS COMPLETED WITH THREADS FROM ANOTHER BATCH. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9085085&pc=GAM | 2210968-2019-87850 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-09-18 | GAM | VICRYL RAP UND 70CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SCROTUM SEAM PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DOES NOT SEEM TO BE SHEATHED AS USUAL, THEY ARE TOO EASY TO PULL OFF THE THREAD. THE SURGERY WAS COMPLETED WITH THREADS FROM ANOTHER BATCH. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9085089&pc=GAM | 2210968-2019-87851 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-09-18 | GAM | VICRYL RAP UND 70CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SCROTUM SEAM PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE DOES NOT SEEM TO BE SHEATHED AS USUAL, THEY ARE TOO EASY TO PULL OFF THE THREAD. THE SURGERY WAS COMPLETED WITH THREADS FROM ANOTHER BATCH. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9076221&pc=GAM | 2210968-2019-87663 | 2019-08-20 | Malfunction | ETHICON INC. | 2019-09-17 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS DETACHED FROM THE NEEDLE WITH WEAK TENSION. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9078475&pc=GAM | 2210968-2019-87697 | 2019-01-20 | Malfunction | ETHICON INC. | 2019-09-17 | GAM | VICRYL SUTURE 27" (70CM) 0 VIO | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A COLECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS UNWOUND. THE NEEDLE WAS LOST INSIDE THE BODY, BUT THEN FOUND AND REMOVED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9070998&pc=GAM | 2210968-2019-87541 | 2019-08-01 | Malfunction | ETHICON INC. | 2019-09-16 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD GOT BROKEN WHEN PASSING THE NEEDLE. THE THREAD BROKE FAR OF THE NEEDLE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9056641&pc=GAM | 2210968-2019-87368 | 2019-08-03 | Malfunction | ETHICON INC. | 2019-09-14 | GAM | VICRYL PLUS CT VIO 70CM M1.5::USP4-0 S/A TF PLUS | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A METAL REMOVAL OF THE PLATE OSTEOSYNTHESIS WITH ADHESIOLYSIS OF OS METACARPALE III PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SCAR STARTED TO BECOME INFLAMED ON (B)(6) 2019 AND THE PATIENT EXPERIENCED A POSSIBLE POST OP WOUND INFECTION. IT WAS REPORTED THAT THE SUTURE HAS NOT RESORBED OR HAS NOT COMPLETELY DISSOLVED AFTER MORE THAN ONE WEEK. A FURTHER OPERATION TO REMOVE THE SUTURE HAS NOT BEEN EXCLUDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED FOR EVALUATION, ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. UPDATE TO PATIENT CURRENT CONDITION. PLEASE ADVISE IF THE PATIENT HAS HAD ADDITIONAL OPERATION TO REMOVE THE SUTURE OR IS A DATE PLANNED? DID THE PATIENT EXPERIENCE AN INFECTION? WERE CULTURES PERFORMED? IF SO, WHAT ARE THE RESULTS? HAS THERE BEEN ANY MEDICATIONS PRESCRIBED TO TREAT THE INFECTION / INFLAMMATION? LOT NUMBER? WAS THERE ANY NON-CONFORMANCE NOTED WITH THE DEVICE BEFORE USE, DURING SUTURING, DURING POST-OP EXAMINATION OR POSSIBLE RE-SURGERY?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9027412&pc=GAM | 2210968-2019-87307 | 2019-07-01 | Malfunction | ETHICON INC. | 2019-09-13 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT THE DOG UNDERWENT AN ANIMAL SURGERY ON (B)(6) 2019 AND THE SUTURE WAS USED. DURING THE SURGERY, THE SUTURE GOT BROKEN WHEN PERFORMING KNOT. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: IT WAS RECEIVED FOR ANALYSIS TWO UNOPENED SAMPLES. DURING THE VISUAL INSPECTION OF TWO UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE REPRESENTATIVE SAMPLES, NO PERFORMANCE - BREAKAGE SUTURE WAS FOUND AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV453, ML8CJGR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9030808&pc=GAM | 2210968-2019-87315 | 2019-07-01 | Malfunction | ETHICON INC. | 2019-09-13 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT THE DOG UNDERWENT AN ANIMAL SURGERY ON (B)(6) 2019 AND THE SUTURE WAS USED. DURING THE SURGERY, THE SUTURE GOT BROKEN WHEN PERFORMING KNOT. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: IT WAS RECEIVED FOR ANALYSIS AN UNOPENED SAMPLE. DURING THE VISUAL INSPECTION OF THE UNOPENED SAMPLE, NO DEFECTS WERE FOUND ON THE PACKET. THE SAMPLE WAS OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURE WAS DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE REPRESENTATIVE SAMPLE, NO PERFORMANCE - BREAKAGE SUTURE WAS FOUND AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV453, ML8CJGR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9008026&pc=GAM | 2210968-2019-87148 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-09-12 | GAM | VICRYL VIO 70CM M0.7 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE KBZ620 BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9013101&pc=GAM | 2210968-2019-87199 | 2019-08-01 | Malfunction | ETHICON INC. | 2019-09-12 | GAM | VICRYL VIO 75CM M4 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON 08/1/2019 AND THE MESH WAS IMPLANTED. THE NEEDLE GOT BROKEN DURING THE USE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV988, MMBDZHQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME ANY ADVERSE PATIENT CONSEQUENCES AND HOW WERE THEY MANAGED? HOW WAS CASE COMPLETED? IS THE ACTUAL DEVICE BEING RETURNED? ARE ANY REPRESENTATIVE SAMPLES BEING RETURNED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8994828&pc=GAM | 2210968-2019-87036 |  | Malfunction | ETHICON INC. | 2019-09-11 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. IT WAS REPORTED THAT THE TIP OF NEEDLE BROKEN OFF WHEN REMOVED STOCK FROM SHELF. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME? SINGLE LOT OR MULTIPLE LOTS INVOLVED? IF MULTIPLE, WHAT ARE THE LOT #'S? WHEN DID THE EVENT OCCUR? PRE OP BEFORE USE ON PATIENT? OR INTRA OP DURING USE ON PATIENT WHILE SUTURING? ANY PATIENT CONSEQUENCE /AE OUTCOME AS A RESULT OF EVENT REPORT? WAS THE NEEDLE PIECE FOUND? AND WHERE? ANY SURGICAL INTERVENTION/RE-SUTURING DONE ON PATIENT POT-OP INITIAL PROCEDURE? PATIENT CURRENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8995179&pc=GAM | 2210968-2019-87046 |  | Malfunction | ETHICON INC. | 2019-09-11 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE WAS BROKEN EASILY DURING USE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE VR945, KP9BBSE0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8982494&pc=GAM | 2210968-2019-86910 |  | Malfunction | ETHICON INC. | 2019-09-10 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE WAS BENT DURING THE BEGINNING OF THE SUTURING PROCEDURE. THE SURGEON INTERN CONTINUED SUTURING THE UTERUS USING THE BENT NEEDLE, AND AFTER SEVERAL MORE STITCHES THE NEEDLE WAS BROKEN NEAR THE NEEDLE-HOLDER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965455&pc=GAM | 2210968-2019-86598 | 2019-08-05 | Malfunction | ETHICON INC. | 2019-09-05 | GAM | VICRYL CT VIO 70CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT SUTURING OF THE RIGHT TEMPORAL SIDE OF THE HEAD ON AN UNKNOWN DATE AND SUTURE WAS USED. THE EYE OF THE NEEDLE WAS MISSING AFTER SUTURING THE SUBCUTANEOUS AT THE RIGHT TEMPORAL SIDE OF THE HEAD. A CLEANING OF THE SURGICAL SITE WAS PERFORMED AND NO PIECE OF THE DEVICE WAS FOUND IN THE SURGICAL SITE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V279H, PBBCXPQ0 NUMBER, AND NO NON-CONFORMANCES RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY 'LACK OF NEEDLE WEDGE' THE "EYE" OF THE NEEDLE WAS MISSING AFTER THE PROCEDURE. WAS THE NEEDLE TIP FOUND MISSING/BROKEN DURING SUTURING? IF NO, EXPLAIN. NOT FOUND. HOW WAS CASE COMPLETED? A CLEANING OF THE SURGICAL SITE WAS PERFORMED AND NO PIECE OF THE DEVICE FOUND INTO THE SURGICAL SITE. ANY ADVERSE PATIENT CONSEQUENCES? NOT REPORTED WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED? CLEANING OF THE SURGICAL SITE ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8967416&pc=GAM | 2210968-2019-86629 | 2019-07-15 | Malfunction | ETHICON INC. | 2019-09-05 | GAM | VICRYL VIO M3 USP2/0 45CM NN | Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS FRAYED ON BOTH ENDS. IT WASN¿T POSSIBLE TO STEATHE THE NEEDLE. IT WASN´T POSSIBLE TO BRING THE SUTURE THROUGH THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. DEVICE EVALUATION SUMMARY: AN EMPTY OPENED FOIL, A WINDING FORMER WITH SIX UN-DISPENSED SUTURE OF PRODUCT CODE V1226, LOT # PABCTLR0 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE, THE SIX SUTURES PRESENT BROOMED END DEFECT, SINCE THE END HAS BECOME PARTIALLY UNRAVELED AND HAS A FLARED APPEARANCE. ALL SUTURES WERE DISPENSED WITHOUT PROBLEMS AND NO DEFECTS OR FRAYING WERE NOTED ALONG OF THE STRANDS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE SAMPLE CONDITION, THE ASSIGNABLE CAUSE OF THE PERFORMANCE FRAYING SUTURE INTRA-OP, WAS A DAMAGED SUTURE (BROOMED END). DEVICE EVALUATION SUMMARY: SIX UNOPENED SAMPLES OF PRODUCT CODE V1226, LOT # PABCTLR0 WERE RETURNED FOR ANALYSIS. THIS PRODUCT CODE IS MULTISTRAND COUNT AND EACH PACKET CONTAINS SIX STRANDS. DURING THE VISUAL INSPECTION OF SIX SAMPLES, NO DEFECTS WERE FOUND ON THE PACKETS, THE SAMPLES WERE OPENED, AND EACH PACKET CONTAINS SIX STRANDS. ALL STRANDS PRESENT BROOMED END DEFECT, SINCE THE END HAS BECOME PARTIALLY UNRAVELED AND HAS A FLARED APPEARANCE. ALL SUTURES WERE DISPENSED WITHOUT PROBLEMS AND NO DEFECTS OR FRAYING WERE NOTED ALONG OF THE STRANDS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE SAMPLES CONDITION, THE ASSIGNABLE CAUSE OF THE PERFORMANCE FRAYING SUTURE INTRA-OP, WAS DAMAGED SUTURE REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 6 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-85383, 2210968-2019-85384, 2210968-2019-85385, 2210968-2019-86626 AND 2210968-2019-86628. ANALYSIS RESULTS HAVE BEEN INCLUDED IN MEDWATCH REPORTS FOR EACH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8967439&pc=GAM | 2210968-2019-86628 | 2019-07-15 | Malfunction | ETHICON INC. | 2019-09-05 | GAM | VICRYL VIO M3 USP2/0 45CM NN | Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS FRAYED ON BOTH ENDS. IT WASN¿T POSSIBLE TO STEA THE NEEDLE. IT WASN´T POSSIBLE TO BRING THE SUTURE THROUGH THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. DEVICE EVALUATION SUMMARY: AN EMPTY OPENED FOIL, A WINDING FORMER WITH SIX UN-DISPENSED SUTURE OF PRODUCT CODE V1226, LOT # PABCTLR0 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE, THE SIX SUTURES PRESENT BROOMED END DEFECT, SINCE THE END HAS BECOME PARTIALLY UNRAVELED AND HAS A FLARED APPEARANCE. ALL SUTURES WERE DISPENSED WITHOUT PROBLEMS AND NO DEFECTS OR FRAYING WERE NOTED ALONG OF THE STRANDS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE SAMPLE CONDITION, THE ASSIGNABLE CAUSE OF THE PERFORMANCE FRAYING SUTURE INTRA-OP, WAS A DAMAGED SUTURE (BROOMED END). DEVICE EVALUATION SUMMARY: SIX UNOPENED SAMPLES OF PRODUCT CODE V1226, LOT # PABCTLR0 WERE RETURNED FOR ANALYSIS. THIS PRODUCT CODE IS MULTISTRAND COUNT AND EACH PACKET CONTAINS SIX STRANDS. DURING THE VISUAL INSPECTION OF SIX SAMPLES, NO DEFECTS WERE FOUND ON THE PACKETS, THE SAMPLES WERE OPENED, AND EACH PACKET CONTAINS SIX STRANDS. ALL STRANDS PRESENT BROOMED END DEFECT, SINCE THE END HAS BECOME PARTIALLY UNRAVELED AND HAS A FLARED APPEARANCE. ALL SUTURES WERE DISPENSED WITHOUT PROBLEMS AND NO DEFECTS OR FRAYING WERE NOTED ALONG OF THE STRANDS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE SAMPLES CONDITION, THE ASSIGNABLE CAUSE OF THE PERFORMANCE FRAYING SUTURE INTRA-OP, WAS DAMAGED SUTURE. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 6 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-85383, 2210968-2019-85384, 2210968-2019-85385, 2210968-2019-86626, 2210968-2019-86628 AND 2210968-2019-86629. ANALYSIS RESULTS HAVE BEEN INCLUDED IN MEDWATCH REPORTS FOR EACH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8967441&pc=GAM | 2210968-2019-86626 | 2019-07-15 | Malfunction | ETHICON INC. | 2019-09-05 | GAM | VICRYL VIO M3 USP2/0 45CM NN | Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS FRAYED ON BOTH ENDS. IT WASN¿T POSSIBLE TO STEATHE THE NEEDLE. IT WASN´T POSSIBLE TO BRING THE SUTURE THROUGH THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. DEVICE EVALUATION SUMMARY: AN EMPTY OPENED FOIL, A WINDING FORMER WITH SIX UN-DISPENSED SUTURE OF PRODUCT CODE V1226, LOT # PABCTLR0 WERE RETURNED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF THE SAMPLE, THE SIX SUTURES PRESENT BROOMED END DEFECT, SINCE THE END HAS BECOME PARTIALLY UNRAVELED AND HAS A FLARED APPEARANCE. ALL SUTURES WERE DISPENSED WITHOUT PROBLEMS AND NO DEFECTS OR FRAYING WERE NOTED ALONG OF THE STRANDS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE SAMPLE CONDITION, THE ASSIGNABLE CAUSE OF THE PERFORMANCE FRAYING SUTURE INTRA-OP, WAS A DAMAGED SUTURE (BROOMED END). DEVICE EVALUATION SUMMARY: SIX UNOPENED SAMPLES OF PRODUCT CODE V1226, LOT # PABCTLR0 WERE RETURNED FOR ANALYSIS. THIS PRODUCT CODE IS MULTISTRAND COUNT AND EACH PACKET CONTAINS SIX STRANDS. DURING THE VISUAL INSPECTION OF SIX SAMPLES, NO DEFECTS WERE FOUND ON THE PACKETS, THE SAMPLES WERE OPENED, AND EACH PACKET CONTAINS SIX STRANDS. ALL STRANDS PRESENT BROOMED END DEFECT, SINCE THE END HAS BECOME PARTIALLY UNRAVELED AND HAS A FLARED APPEARANCE. ALL SUTURES WERE DISPENSED WITHOUT PROBLEMS AND NO DEFECTS OR FRAYING WERE NOTED ALONG OF THE STRANDS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE SAMPLES CONDITION, THE ASSIGNABLE CAUSE OF THE PERFORMANCE FRAYING SUTURE INTRA-OP, WAS DAMAGED SUTURE. REPRESENTATIVE SAMPLES RETURNED TO SUPPORT INVESTIGATION OF 6 DEVICE ISSUES CAPTURED IN REPORTS 2210968-2019-85383, 2210968-2019-85384, 2210968-2019-85385, 2210968-2019-86628 AND 2210968-2019-86629. ANALYSIS RESULTS HAVE BEEN INCLUDED IN MEDWATCH REPORTS FOR EACH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8952789&pc=GAM | 2210968-2019-86461 | 2019-01-04 | Malfunction | ETHICON INC. | 2019-08-30 | GAM | VICRYL RAP UND90CM 2/0 CT-1 PLUS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EPISIOTOMY REPAIR PROCEDURE IN 2019 AND THE SUTURE WAS USED. DURING AN EPISIOTOMY REPAIR ON THE KNOT TIGHTENING STAGE, THE SUTURE BREAKAGE OCCURRED. TWO PIECES OF BROKEN SUTURE WERE THE PART THAT WAS TORN FROM THE NEEDLE TO THE POINT OF TEAR AND THE PART WHERE THE SUTURE WAS TORN TO THE PART WHERE THE SUTURE WAS CUT BY THE SURGEON. ANOTHER LIKE SUTURE WAS USED TO CLOSE THE TISSUE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8947107&pc=GAM | 2210968-2019-86442 | 2019-07-03 | Malfunction | ETHICON INC. | 2019-08-29 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Inflammation | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON UNKNOWN DATE AND THE SUTURE WAS USED. THERE WERE INFLAMMATION AT THE SUTURE SITE AND NO ABSORPTION OF THE SUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME? LOT #? ON WHAT TISSUE WAS THE SUTURE USED? WAS ANY MEDICAL / SURGICAL INTERVENTION PERFORMED? WAS THERE ANY OTHER TREATMENT ( MEDICAL OR SURGICAL INTERVENTION / ANTIBIOTICS OR PRESCRIPTION MEDICATION) / RESUTURING PROVIDED TO PATIENT FOR THE INFLAMMATION POST-OP ? PATIENT CURRENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8940542&pc=GAM | 2210968-2019-86377 | 2019-07-03 | Malfunction | ETHICON INC. | 2019-08-28 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE PATIENT EXPERIENCED INFLAMMATION AT THE SITE ON AN UNKNOWN DATE POST PROCEDURE AND NO ABSORPTION OF THE SUTURE. THE PATIENT CURRENT CONDITION IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME? LOT #? ON WHAT TISSUE WAS THE SUTURE USED? WAS ANY MEDICAL / SURGICAL INTERVENTION PERFORMED ON THE PATIENT? WAS THERE ANY OTHER TREATMENT ( MEDICAL OR SURGICAL INTERVENTION / ANTIBIOTICS OR PRESCRIPTION MEDICATION) / RE-SUTURING PROVIDED TO PATIENT FOR THE INFLAMMATION POST-OP ? PATIENT CURRENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8940625&pc=GAM | 2210968-2019-86379 | 2019-07-03 | Malfunction | ETHICON INC. | 2019-08-28 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE PATIENT EXPERIENCED INFLAMMATION AT THE SITE ON AN UNKNOWN DATE POST PROCEDURE AND NO ABSORPTION OF THE SUTURE. THE PATIENT CURRENT CONDITION IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME? LOT #? ON WHAT TISSUE WAS THE SUTURE USED? WAS ANY MEDICAL / SURGICAL INTERVENTION PERFORMED ON THE PATIENT? WAS THERE ANY OTHER TREATMENT ( MEDICAL OR SURGICAL INTERVENTION / ANTIBIOTICS OR PRESCRIPTION MEDICATION) / RE SUTURING PROVIDED TO PATIENT FOR THE INFLAMMATION POST-OP ? PATIENT CURRENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8940677&pc=GAM | 2210968-2019-86378 | 2019-07-03 | Malfunction | ETHICON INC. | 2019-08-28 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Inflammation | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON UNKNOWN DATE AND THE SUTURE WAS USED. THERE WERE INFLAMMATION AT THE SUTURE SITE AND NO ABSORPTION OF THE SUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME? LOT #? ON WHAT TISSUE WAS THE SUTURE USED? WAS ANY MEDICAL / SURGICAL INTERVENTION PERFORMED? WAS THERE ANY OTHER TREATMENT ( MEDICAL OR SURGICAL INTERVENTION / ANTIBIOTICS OR PRESCRIPTION MEDICATION) / RE-SUTURING PROVIDED TO PATIENT FOR THE INFLAMMATION POST-OP ? PATIENT CURRENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941447&pc=GAM | 2210968-2019-86390 |  | Malfunction | ETHICON INC. | 2019-08-28 | GAM | VICRYL RAP UND 45CM 4/0 FS-2S | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE NEEDLE LOOSENS/COMES OFF THE SUTURE DURING SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PABCDHT0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941448&pc=GAM | 2210968-2019-86392 |  | Malfunction | ETHICON INC. | 2019-08-28 | GAM | VICRYL RAP UND 45CM 4/0 FS-2S | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE NEEDLE LOOSENS/COMES OFF THE SUTURE DURING SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PABCDHT0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8922680&pc=GAM | MW5089240 | 2019-08-16 | Malfunction | ETHICON | 2019-08-21 | GAM | VICRYL PLUS ON A 3-0 SH NEEDLE | Break; Detachment of Device or Device Component | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: TIP OF 3-0 SH TAPER NEEDLE ON AN ETHICON, COATED VICRYL PLUS ANTIBACTERIAL SUTURE, BROKE OFF WHILE CLOSING FASCIA. NEEDLE AND TIP ACCOUNTED FOR IN THEIR ENTIRETY. FDA SAFETY REPORT ID# (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912692&pc=GAM | 2210968-2019-85854 |  | Malfunction | ETHICON INC. | 2019-08-20 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN OBGYN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE WAS BROKEN EASILY DURING USE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912708&pc=GAM | 2210968-2019-85855 |  | Malfunction | ETHICON INC. | 2019-08-20 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN OB-GYN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE WAS BROKEN EASILY DURING USE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8913349&pc=GAM | 2210968-2019-85875 | 2019-07-24 | Malfunction | ETHICON INC. | 2019-08-20 | GAM | VICRYL*VLT PLUS 1 90CM 1-CT | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ABDOMINAL COMPLEX SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE CUT. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A PHOTO UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8913352&pc=GAM | 2210968-2019-85874 | 2019-07-24 | Malfunction | ETHICON INC. | 2019-08-20 | GAM | VICRYL*VLT PLUS 1 90CM 1-CT | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ABDOMINAL COMPLEX SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE CUT. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A PHOTO UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8913734&pc=GAM | 2210968-2019-85883 |  | Malfunction | ETHICON INC. | 2019-08-20 | GAM | VICRYL RAP UND 75CM M1 USP5/0 SGLE ARMED V-18 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. AFTER SOME STITCHES, THE NEEDLE TOOK OFF THE STRAND. NO PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895385&pc=GAM | 2210968-2019-85576 |  | Malfunction | ETHICON INC. | 2019-08-14 | GAM | VICRYL VIO 75CM M3 | Break; Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE PULLED OFF AND BROKE EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PBBCCMR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895422&pc=GAM | 2210968-2019-85577 | 2019-05-31 | Malfunction | ETHICON INC. | 2019-08-14 | GAM | VICRYL VIO 75CM M4 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT COLOSTOMY PROCEDURE BY LAPAROSCOPY AND SMALL LAPAROTOMY ON (B)(6) 2019 AND SUTURE WAS USED. AT THE END OF THE PROCEDURE, THE NEEDLE PULLED OFF DURING THE SUTURE BY THE SURGEON. USE OF ANOTHER THREAD, THE PROCEDURE HAS BEEN FINISHED WITHOUT ANY ISSUE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE ML8BMZQ0 NUMBER, AND NO NON-CONFORMANCES RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8887297&pc=GAM | 2210968-2019-85459 | 2019-07-12 | Malfunction | ETHICON INC. | 2019-08-13 | GAM | VICRYL CT VIO 75CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE TIP OF THE NEEDLE BROKE DURING THE SUTURE. IT WAS FOUND IN THE FASCIA AND RECOVERED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV2017, PCBBKZR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889244&pc=GAM | 2210968-2019-85496 |  | Malfunction | ETHICON INC. | 2019-08-13 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE WHEN TYING. IT IS UNKNOWN HOW THE CASE WAS COMPLETED. NO DEVICE WILL BE RETURNED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED AND NO ADDITIONAL INFORMATION HAS BEEN RECEIVED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8890878&pc=GAM | 2210968-2019-85510 | 2019-07-17 | Malfunction | ETHICON INC. | 2019-08-13 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A VAGINAL HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8891335&pc=GAM | 2210968-2019-85519 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-08-13 | GAM | VICRYL CT UD 45CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND THE SUTURE WAS USED. DURING USE, THE NEEDLE PULL OFF OCCURRED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8881150&pc=GAM | 2210968-2019-85383 | 2019-07-15 | Malfunction | ETHICON INC. | 2019-08-09 | GAM | VICRYL VIO M3 USP2/0 45CM NN | Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS FRAYED ON BOTH ENDS. IT WASN'T POSSIBLE TO STEATHE THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. CAN YOU EXPLAIN " IT WASN´T POSSIBLE TO STEATHE THE NEEDLE"?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8881151&pc=GAM | 2210968-2019-85384 | 2019-07-15 | Malfunction | ETHICON INC. | 2019-08-09 | GAM | VICRYL VIO M3 USP2/0 45CM NN | Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS FRAYED ON BOTH ENDS. IT WASN'T POSSIBLE TO STEATHE THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. CAN YOU EXPLAIN " IT WASN´T POSSIBLE TO STEATHE THE NEEDLE"?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8881160&pc=GAM | 2210968-2019-85385 | 2019-07-15 | Malfunction | ETHICON INC. | 2019-08-09 | GAM | VICRYL VIO M3 USP2/0 45CM NN | Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS FRAYED ON BOTH ENDS. IT WASN'T POSSIBLE TO STEATHE THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. CAN YOU EXPLAIN " IT WASN´T POSSIBLE TO STEATHE THE NEEDLE"?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876327&pc=GAM | 2210968-2019-85327 |  | Malfunction | ETHICON INC. | 2019-08-08 | GAM | VICRYL RAP UND90CM 2/0 CT-1 PLUS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EPISIOTOMY REPAIR PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE SUTURE BROKE AFTER TYING THE KNOT AND SECURING IT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876355&pc=GAM | 2210968-2019-85328 | 2019-07-11 | Malfunction | ETHICON INC. | 2019-08-08 | GAM | VICRYL RAP UND90CM 2/0 CT-1 PLUS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EPISIOTOMY REPAIR PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BROKE WHEN TIGHTENING THE KNOT AND SECURING IT. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: A DISPENSED NEEDLE/SUTURE BROKEN IN TWO SECTION OF PRODUCT CODE V9450H WAS RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED AND THE SUTURE WAS RECEIVED BROKEN IN TWO SECTION. BODY FLUIDS WERE NOTED ON THE SUTURE PIECES AND THE ENDS WERE NOTED DETACHED DUE TO STRESS APPLIED ON THE STRAND. IN ADDITION, NO FUNCTIONAL TEST CAN BE PERFORMED DUE TO CONDITION OF THE SAMPLE RECEIVED. THE MANUFACTURING RECORDS COULDN'T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. ACCORDING TO VISUAL INSPECTION OF THE SAMPLE, THE ASSIGNABLE CAUSE OF THE BREAKAGE SUTURE IS AN IMPROPER HANDLING. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8860121&pc=GAM | 2210968-2019-85147 | 2019-06-01 | Malfunction | ETHICON INC. | 2019-08-05 | GAM | VICRYL RAP UND 75CM M1 USP5/0 SGLE ARMED C-3 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PAROTIDECTOMY AND FRENECTOMY PROCEDURE IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE OR SPLIT IN TWO PARTS. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. NOTE: EVENT REPORTED IN MEDWATCH REPORTS 2210968-2019-85146 AND 2210968-2019-85147.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8860124&pc=GAM | 2210968-2019-85146 | 2019-06-01 | Malfunction | ETHICON INC. | 2019-08-05 | GAM | VICRYL RAP UND 75CM M1 USP5/0 SGLE ARMED C-3 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PAROTIDECTOMY AND FRENECTOMY PROCEDURE IN (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE OR SPLIT IN TWO PARTS. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. NOTE: EVENT REPORTED IN MEDWATCH REPORTS 2210968-2019-85146 AND 2210968-2019-85147.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8858200&pc=GAM | 2210968-2019-85100 | 2019-07-04 | Malfunction | ETHICON INC. | 2019-08-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K183183 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND THE SUTURE WAS USED FOR THE SURGICAL SUTURING OF THE LAYER OF MUSCLE FIBER AT LUMBOSACRAL LEVEL. AFTER THE SUTURE, THE NEEDLE DID NOT PRESENT THE TIP. AN ACCURATE WASHING AND CAREFUL RESEARCH OF THE MISSING PART HAVE BEEN DONE IN SURGICAL SITE WITH NO PROOF. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER MJ8BHZQ0 AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY PATIENT CONSEQUENCES? WHAT MEDICAL/SURGICAL INTERVENTION WAS PROVIDED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8840645&pc=GAM | 2210968-2019-84853 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-30 | GAM | VICRYL 3-0 1X75CM FS-1 VIOLET | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A VETERINARIAN THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE SUTURE OF THE ABDOMINAL WALL OR OF THE SUBCUTANEOUS TISSUE, THE NEEDLE PULLED OFF FROM THE STRAND AFTER SOME TISSUE PASSAGES, WITHOUT ANY SPECIFIC FORCE APPLIED. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). FINISHED GOODS RECORDS OF JV452, MK8DPCR0 HAVE BEEN REVIEWED AND SHOW NO DEVIATIONS; ALL RESULTS MEET THE SPECIFIED QUALITY REQUIREMENTS. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836853&pc=GAM | 2210968-2019-84792 |  | Malfunction | ETHICON INC. | 2019-07-29 | GAM | VICRYL RAP UND 75 4/0 SGL ARM PS-2 PRIME | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE IN THE MIDDLE OF THE THREAD. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CONFIRM THE SPECIFIC NUMBER OF PATIENT EVENTS. HAVE ANY OF THESE EVENTS BEEN PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE COMPLAINT FILE NUMBER/ AFFILIATE NUMBER. FOR EACH PATIENT EVENT PLEASE PROVIDE ADDITIONAL COMPLAINT DETAILS: EVENT DATE, PRODUCT CODE AND LOT NUMBER INVOLVED, PROCEDURE, EVENT DESCRIPTION, QTY INVOLVED? WAS THERE ANY ADVERSE PATIENT OUTCOME? IF YES, WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED, WAS PROCEDURE COMPLETED SUCCESSFULLY AND HOW? ANY SAMPLE RETURN..ETC? RESPONSE RECEIVED: UNFORTUNATELY I DON¿T HAVE ANY FURTHER INFO. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ATTEMPTS WERE MADE TO OBTAIN THE INFORMATION. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8832976&pc=GAM | 2210968-2019-84709 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-26 | GAM | VICRYL*UND 4-0 45CM (1)P-3 PR | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY IN 2019 AND THE SUTURE WAS USED. DURING THE SURGERY, THE SUTURE WAS BROKEN WHEN HANDLED WITH NEEDLE HOLDERS. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE (B)(4) NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8822516&pc=GAM | 2210968-2019-84610 | 2019-06-20 04:00 | Malfunction | ETHICON INC. | 2019-07-24 04:00 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED V-34 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING USE THE SUTURE AND THE NEEDLE CAME OFF THE THREAD. THERE WAS NO PATIENT CONSEQUENCE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MPBBCMQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8821866&pc=GAM | 2210968-2019-84607 |  | Malfunction | ETHICON INC. | 2019-07-24 | GAM | VICRYL RAP UND 75CM 3-0 FS-2 | Material Frayed; Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN UROLOGY PROCEDURE ON AND UNKNOWN DATE AND SUTURE WAS USED. A SUTURE THAT WAS BEING USED AND CAME OFF THE NEEDLE. NO PATIENT CONSEQUENCES WERE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8817550&pc=GAM | 2210968-2019-84585 | 2019-06-17 | Malfunction | ETHICON INC. | 2019-07-23 | GAM | VICRYL* VLT PLUS 0 70CM (1)SH | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AND (B)(6) 2019 AND SUTURE WAS USED. DURING USE OF THE SUTURE, THE FILAMENTS BROKE. NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE FOLLOWING INFORMATION WAS REQUESTED, HOWEVER NO RESPONSE HAS BEEN RECEIVED TO DATE: IS THE PRODUCT AVAILABLE FOR EVALUATION? PLEASE PROVIDE DEVICE RETURN UPDATE OR FOLLOW UP AS REQUIRED. PLEASE ADVISE IF EVENTS HAVE BEEN REPORTED? SPECIFIC NUMBER OF SURGERIES IN WHICH THE ISSUE OCCURRED. HOW MANY SUTURES BROKE DURING EACH SURGERY? SURGERY DATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8817696&pc=GAM | 2210968-2019-84586 | 2019-06-25 | Malfunction | ETHICON INC. | 2019-07-23 | GAM | VICRYL*RAPID UND 2-0 90CM 1V34 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EPISIOTOMY REPAIR ON (B)(6) 2019 AND THE SUTURE WAS USED. WHEN TYING A KNOT, THE SUTURE SNAPPED WHEN SUTURING AT THE MUCOSA AREA. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT : (B)(4). DATE SENT TO THE FDA: 07/23/2019. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AK8607 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8818009&pc=GAM | 2210968-2019-84591 |  | Malfunction | ETHICON INC. | 2019-07-23 | GAM | VICRYL*RAPID UND 2-0 90CM 1V34 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND THE SUTURE WAS USED. WHEN TYING A KNOT, THE SUTURE SNAPPED WHEN SUTURING AT THE MUCOSA AREA. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AK8607 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8818086&pc=GAM | 2210968-2019-84592 |  | Malfunction | ETHICON INC. | 2019-07-23 | GAM | VICRYL*RAPID UND 2-0 90CM 1V34 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2019 AND THE SUTURE WAS USED. WHEN TYING A KNOT, THE SUTURE SNAPPED WHEN SUTURING AT THE MUCOSA AREA. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # : (B)(4). DATE SENT TO THE FDA: 07/23/2019. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AK8607 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8811606&pc=GAM | 2210968-2019-84531 | 2019-05-29 | Malfunction | ETHICON INC. | 2019-07-22 | GAM | VICRYL*RAPID UND 2-0 90CM 1CT1 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND THE SUTURE WAS USED. WHEN THE SUTURE WAS BEING USED, THE FILAMENTS RUPTURED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8814399&pc=GAM | 2210968-2019-84551 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-22 | GAM | VICRYL RAP UND 75CM M2 | Break | Wound Dehiscence; Not Applicable | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. APPROXIMATELY 2 TO 3 DAYS AFTER THE PROCEDURE, THE PATIENT EXPERIENCED FAST RESORPTION AND DISUNITY OF THE SUTURE. THE PATIENT UNDERWENT PERINEAL SUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AK3046 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS THE DEFECT WITH THE PRODUCT? FAST RESORPTION OF THE SUTURE. WAS IT NOTICED UPON OPENING THE PACKAGE OR DURING USE OR OTHER (SPECIFY)? THE ISSUE WAS OBSERVED 2 TO 3 DAYS AFTER THE PROCEDURE. ANY ADVERSE PATIENT CONSEQUENCES? DISUNITY OF THE SUTURE. WHAT MEDICAL/SURGICAL INTERVENTION PROVIDED TO MANAGE THEM? PERINEAL SUTURE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE. DATE AND NAME OF THE INDEX SURGICAL PROCEDURE. THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? ON WHAT TISSUE WAS THE SUTURE USED? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR TIN, CALCIFIED, FRAGILE, DISEASED? HOW WAS THE SUTURE PLACED, INTERRUPTED OR CONTINUOUS? WHAT WERE THE CURRENT SYMPTOMS FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? WHAT WAS THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? WHAT IS THE PATIENT¿S CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8809924&pc=GAM | 2210968-2019-84504 | 2019-06-25 | Malfunction | ETHICON INC. | 2019-07-20 | GAM | VICRYL*RAPID UND 2-0 90CM 1V34 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EPISIOTOMY REPAIR ON (B)(6) 2019 AND THE SUTURE WAS USED. WHEN TYING A KNOT, THE SUTURE SNAPPED WHEN SUTURING AT THE MUCOSA AREA. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8805844&pc=GAM | 2210968-2019-84475 |  | Malfunction | ETHICON INC. | 2019-07-19 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED VIA A JOURNAL ARTICLE: TITLE: PROSPECTIVE EVALUATION OF STAPLED HEMORRHOIDOPEXY VERSUS TRANSANAL HAEMORRHOIDAL DEARTERIALISATION FOR STAGE II AND III HAEMORRHOIDS: THREE-YEAR OUTCOMES AUTHORS: P. GIORDANO, P. NASTRO, A. DAVIES, G. GRAVANTE CITATION: TECH COLOPROCTOL. 2011; 15: 67¿73. DOI: 10.1007/S10151-010-0667-Z. THE AIM OF THE STUDY WAS TO COMPARE SHORT- AND MEDIUM-TERM OUTCOMES OF TRANSANAL HEMORRHOIDAL DEARTERIALIZATION (THD) VERSUS STAPLED HEMORRHOIDOPEXY (SH) FOR THE TREATMENT OF SECOND- AND THIRD-DEGREE HEMORRHOIDS. A TOTAL OF 28 PATIENTS (AGE RANGE: 23 TO 73 YEARS OLD; 20 MALE AND 8 FEMALE) UNDERWENT THD AND 24 PATIENTS (AGE RANGE: 35 TO 78 YEARS OLD; 16 MALE AND 8 FEMALE PATIENTS) UNDERWENT SH. DURING THE SURGICAL PROCEDURE IN THE THD GROUP, ONCE IDENTIFIED, THE HEMORRHOIDAL ARTERIES WERE TRANSFIXED AND LIGATED USING VICRYL 2-0 ABSORBABLE SUTURES (ETHICON) IN A FIGURE-OF-EIGHT STITCH. IN THE SH GROUP, THE DEDICATED PPH 03 CIRCULAR STAPLING DEVICE (ETHICON) WAS THEN USED FOR MUCOSECTOMY AND ANOPEXY. IN THE THD GROUP, REPORTED COMPLICATIONS INCLUDED POST-OPERATIVE PAIN VAS (N-2), DELAYED DISCHARGE (N-1), SUBMUCOSAL HEMATOMA (N-1), TECHNICAL PROBLEMS (N-2) IN WHICH THE TIP OF THE NEEDLE SNAPPED OFF DURING LIGATION AND WAS NOT RETRIEVED. THIS DID NOT CAUSE ANY SYMPTOMS AND BOTH PATIENTS HAD A SUCCESSFUL OUTCOME, HEMORRHOIDAL PROLAPSE (N-3), AND FREQUENT BLEEDING (N-1). IN THE SH GROUP, REPORTED COMPLICATIONS INCLUDED POST-OPERATIVE PAIN VAS (N-3.5), SEVERE POST-OPERATIVE PAIN (N-1) WHICH REQUIRED READMISSION, DELAYED DISCHARGE (N-1), TRANSIENT FECAL URGENCY (N-2), POST-OPERATIVE BLEEDING (N-2) WHICH REQUIRED READMISSION AND MANAGED CONSERVATIVELY, RECTAL STENOSIS (N-1), HEMORRHOIDAL PROLAPSE (N-2), FREQUENT BLEEDING (N-1), AND COMPLETE OCCLUSION OF THE RECTAL LUMEN FOLLOWING FIRING OF THE STAPLER (N-1) IN WHICH THE COMPLICATION WAS SUCCESSFULLY MANAGED ENDOSCOPICALLY, AND THE PATIENT WAS DISCHARGED HOME 2 DAYS LATER. SHORT-TERM RESULTS ALTHOUGH SIMILAR SEEM TO SUGGEST SH MAY RESULT IN INCREASED MORBIDITY WHILE RETURN TO WORK IS QUICKER AFTER THD. GIVEN ITS MEDIUM-TERM RESULTS COMPARABLE TO SH IN TERMS OF RECURRENCE RATES AND PATIENT SATISFACTION, THE AUTHORS SUGGESTED THAT THD COULD BE CONSIDERED A VALID ALTERNATIVE FIRST-LINE SURGICAL OPTION FOR THE TREATMENT OF SECOND- AND THIRD-DEGREE HEMORRHOIDS. Manufacturer Narrative: (B)(6). THIS REPORT IS RELATED TO A JOURNAL ARTICLE, THEREFORE NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE MANUFACTURING RECORD EVALUATION CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. CITATION: TECH COLOPROCTOL. 2011; 15: 67¿73. DOI: 10.1007/S10151-010-0667-Z. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8805825&pc=GAM | 2210968-2019-84473 |  | Malfunction | ETHICON INC. | 2019-07-18 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED VIA A JOURNAL ARTICLE: TITLE: PROSPECTIVE EVALUATION OF STAPLED HEMORRHOIDOPEXY VERSUS TRANSANAL HAEMORRHOIDAL DEARTERIALISATION FOR STAGE II AND III HAEMORRHOIDS: THREE-YEAR OUTCOMES AUTHORS: P. GIORDANO, P. NASTRO, A. DAVIES, G. GRAVANTE CITATION: TECH COLOPROCTOL. 2011; 15: 67¿73. DOI: 10.1007/S10151-010-0667-Z. THE AIM OF THE STUDY WAS TO COMPARE SHORT- AND MEDIUM-TERM OUTCOMES OF TRANSANAL HEMORRHOIDAL DEARTERIALIZATION (THD) VERSUS STAPLED HEMORRHOIDOPEXY (SH) FOR THE TREATMENT OF SECOND- AND THIRD-DEGREE HEMORRHOIDS. A TOTAL OF 28 PATIENTS (AGE RANGE: 23 TO 73 YEARS OLD; 20 MALE AND 8 FEMALE) UNDERWENT THD AND 24 PATIENTS (AGE RANGE: 35 TO 78 YEARS OLD; 16 MALE AND 8 FEMALE PATIENTS) UNDERWENT SH. DURING THE SURGICAL PROCEDURE IN THE THD GROUP, ONCE IDENTIFIED, THE HEMORRHOIDAL ARTERIES WERE TRANSFIXED AND LIGATED USING VICRYL 2-0 ABSORBABLE SUTURES (ETHICON) IN A FIGURE-OF-EIGHT STITCH. IN THE SH GROUP, THE DEDICATED PPH 03 CIRCULAR STAPLING DEVICE (ETHICON) WAS THEN USED FOR MUCOSECTOMY AND ANOPEXY. IN THE THD GROUP, REPORTED COMPLICATIONS INCLUDED POST-OPERATIVE PAIN VAS (N-2), DELAYED DISCHARGE (N-1), SUBMUCOSAL HEMATOMA (N-1), TECHNICAL PROBLEMS (N-2) IN WHICH THE TIP OF THE NEEDLE SNAPPED OFF DURING LIGATION AND WAS NOT RETRIEVED. THIS DID NOT CAUSE ANY SYMPTOMS AND BOTH PATIENTS HAD A SUCCESSFUL OUTCOME, HEMORRHOIDAL PROLAPSE (N-3), AND FREQUENT BLEEDING (N-1). IN THE SH GROUP, REPORTED COMPLICATIONS INCLUDED POST-OPERATIVE PAIN VAS (N-3.5), SEVERE POST-OPERATIVE PAIN (N-1) WHICH REQUIRED READMISSION, DELAYED DISCHARGE (N-1), TRANSIENT FECAL URGENCY (N-2), POST-OPERATIVE BLEEDING (N-2) WHICH REQUIRED READMISSION AND MANAGED CONSERVATIVELY, RECTAL STENOSIS (N-1), HEMORRHOIDAL PROLAPSE (N-2), FREQUENT BLEEDING (N-1), AND COMPLETE OCCLUSION OF THE RECTAL LUMEN FOLLOWING FIRING OF THE STAPLER (N-1) IN WHICH THE COMPLICATION WAS SUCCESSFULLY MANAGED ENDOSCOPICALLY, AND THE PATIENT WAS DISCHARGED HOME 2 DAYS LATER. SHORT-TERM RESULTS ALTHOUGH SIMILAR SEEM TO SUGGEST SH MAY RESULT IN INCREASED MORBIDITY WHILE RETURN TO WORK IS QUICKER AFTER THD. GIVEN ITS MEDIUM-TERM RESULTS COMPARABLE TO SH IN TERMS OF RECURRENCE RATES AND PATIENT SATISFACTION, THE AUTHORS SUGGESTED THAT THD COULD BE CONSIDERED A VALID ALTERNATIVE FIRST-LINE SURGICAL OPTION FOR THE TREATMENT OF SECOND- AND THIRD-DEGREE HEMORRHOIDS. Manufacturer Narrative: (B)(4). THIS REPORT IS RELATED TO A JOURNAL ARTICLE, THEREFORE NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE MANUFACTURING RECORD EVALUATION CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. CITATION: TECH COLOPROCTOL. 2011; 15: 67¿73. DOI: 10.1007/S10151-010-0667-Z. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8798946&pc=GAM | 2210968-2019-84402 | 2019-06-05 | Malfunction | ETHICON INC. | 2019-07-17 | GAM | VICRYL* UND 3-0 70CM (1)SC-20 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019, AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE FRACTURED, LEAVING THE FRAGMENT OUTSIDE THE PATIENT'S TISSUE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8794427&pc=GAM | 2210968-2019-84337 | 2019-06-06 | Malfunction | ETHICON INC. | 2019-07-16 | GAM | VICRYL PLUS CT VIO 90CM M4::USP1 S/A CT | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN SECTION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE DETACHED FROM THE NEEDLE. THE NEEDLE WAS REMOVED AND VERIFIED ITS INTEGRITY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8796742&pc=GAM | 2210968-2019-84366 | 2019-06-17 | Malfunction | ETHICON INC. | 2019-07-16 | GAM | VICRYL*UND 5-0 45CM (1)P-3 PR | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE BREAKS DURING THE PROCEDURE. IT IS UNKNOWN WHAT WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AL1709 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8787397&pc=GAM | 2210968-2019-84258 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-13 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL PELVIC PROCEDURE IN 2019 AND THE SUTURE WAS USED. WHILE THE TISSUE WAS PENETRATING DURING THE PROCEDURE, A NEEDLE SNAPPED IN THE MIDDLE. THE BROKEN PIECE WAS RETRIEVED AND ANOTHER LIKE SUTURE WAS USED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8787398&pc=GAM | 2210968-2019-84259 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-13 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL PELVIC PROCEDURE IN 2019 AND THE SUTURE WAS USED. WHILE THE TISSUE WAS PENETRATING DURING THE PROCEDURE, A NEEDLE SNAPPED IN THE MIDDLE. THE BROKEN PIECE WAS RETRIEVED AND ANOTHER SUTURE WAS USED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8787401&pc=GAM | 2210968-2019-84261 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-13 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL PELVIC PROCEDURE IN 2019 AND THE SUTURE WAS USED. WHILE THE TISSUE WAS PENETRATING DURING THE PROCEDURE, A NEEDLE SNAPPED IN THE MIDDLE. THE BROKEN PIECE WAS RETRIEVED AND ANOTHER LIKE SUTURE WAS USED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8787402&pc=GAM | 2210968-2019-84262 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-13 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL PELVIC PROCEDURE IN 2019 AND THE SUTURE WAS USED. WHILE THE TISSUE WAS PENETRATING DURING THE PROCEDURE, A NEEDLE SNAPPED IN THE MIDDLE. THE BROKEN PIECE WAS RETRIEVED AND ANOTHER LIKE SUTURE WAS USED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8784371&pc=GAM | 2210968-2019-84221 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-12 | GAM | VICRYL RAP UND 75CM 3-0 SH-1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE DETACHED EASILY FROM THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE W9974, ME8BCDR0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8786637&pc=GAM | 2210968-2019-84238 | 2019-06-18 | Malfunction | ETHICON INC. | 2019-07-12 | GAM | VICRYL RAPIDE UND 2/0 70CM MH PLUS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS BROKEN DURING PERINEAL SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PG1674H, MMCBMLE0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8786638&pc=GAM | 2210968-2019-84239 | 2019-06-18 | Malfunction | ETHICON INC. | 2019-07-12 | GAM | VICRYL RAPIDE UND 2/0 70CM MH PLUS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE WAS BROKEN DURING PERINEAL SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PG1674H, MMCBMLE0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8787151&pc=GAM | 2210968-2019-84251 | 2019-06-19 | Malfunction | ETHICON INC. | 2019-07-12 | GAM | VICRYL UND 75CM M2 USP3/0 SGL ARM FS-2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE DETACHED FROM THE SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8782285&pc=GAM | 2210968-2019-84117 | 2019-06-12 | Malfunction | ETHICON INC. | 2019-07-11 | GAM | VICRYL VIO 75CM M3.5::USP0 SGLE ARMED UCL | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE OFF. AN INTENSIFYING SCREEN WAS USED TO FIND THE NEEDLE IN ORDER TO RETRIEVE IT FROM THE PATIENT. THERE WAS AN UNKNOWN DELAY OF THE PROCEDURE, THE DURATION OF ANESTHESIA WAS INCREASED. THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV115, LOT LB7DKRN, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE? NAME OF INDEX SURGICAL PROCEDURE? WAS THE PATIENT¿S POST-OPERATIVE COURSE CHANGED DUE TO THE REPORTED EVENT? WILL THE DEVICE INVOLVED IN THE EVENT BE RETURNED FOR EVALUATION? WHAT IS THE PATIENT CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8782804&pc=GAM | 2210968-2019-84125 | 2019-06-07 | Malfunction | ETHICON INC. | 2019-07-11 | GAM | VICRYL* UND 4-0 45CM (1)E-19 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE RUPTURED AT THE TIME OF THE PATIENT¿S SKIN SYNTHESIS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). IT WAS REPORTED THAT THE DEVICE HAD A BREAKAGE SUTURE ISSUE. ONE OPENED BOX WITH UNOPENED SAMPLES WAS RETURNED FOR ANALYSIS. THE COMPLAINT ISSUE WAS NOT RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE REPRESENTATIVE SAMPLES, NO PERFORMANCE - BREAKAGE SUTURE WAS FOUND AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED (B)(4), AND NO NON-CONFORMANCES WERE IDENTIFIED ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8782805&pc=GAM | 2210968-2019-84123 | 2019-06-07 | Malfunction | ETHICON INC. | 2019-07-11 | GAM | VICRYL* UND 4-0 45CM (1)E-19 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE RUPTURED AT THE TIME OF THE PATIENT¿S SKIN SYNTHESIS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE (B)(4), AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8782808&pc=GAM | 2210968-2019-84124 | 2019-06-07 | Malfunction | ETHICON INC. | 2019-07-11 | GAM | VICRYL* UND 4-0 45CM (1)E-19 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE RUPTURED AT THE TIME OF THE PATIENT¿S SKIN SYNTHESIS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED (B)(4), AND NO NON-CONFORMANCES WERE IDENTIFIED ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8775617&pc=GAM | 2210968-2019-83865 | 2018-09-06 | Malfunction | ETHICON INC. | 2019-07-10 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. THE NEEDLE TOOK AWAY FROM THE THREAD AT THE END OF THE PROCEDURE AND THE SUTURE OPENED. THE SURGERY WAS DELAYED 10 MINUTES. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AJ8073 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW LONG WAS THE DELAY? LESS THAN 10 MINUTES. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? NO. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8775895&pc=GAM | 2210968-2019-83868 | 2019-06-14 | Malfunction | ETHICON INC. | 2019-07-10 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE FOR IMPLANT REMOVAL ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE CAME OUT FROM THE SWAGED POINT WHILE TYING THE KNOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8769034&pc=GAM | 2210968-2019-83748 | 2018-06-04 | Malfunction | ETHICON INC. | 2019-07-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF FROM THE THREAD. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE EVENT OF "THE NEEDLE PULLED OFF FROM THE THREAD 3 TIMES" OCCUR DURING USE ON PATIENT WHILE SUTURING OR BEFORE USE ON PATIENT WHILE DISPENSING/HANDLING THE DEVICE? NO FURTHER INFORMATION AVAILABLE. ANY SAMPLE AVAILABLE TO BE RETURNED FOR EVALUATION? THE DEVICE IS NOT AVAILABLE FOR ANALYSIS. NOTE: EVENT REPORTED IN 2210968-2019-83746 AND 2210968-2019-83747.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8769066&pc=GAM | 2210968-2019-83747 | 2018-06-04 | Malfunction | ETHICON INC. | 2019-07-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF FROM THE THREAD. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE EVENT OF "THE NEEDLE PULLED OFF FROM THE THREAD 3 TIMES" OCCUR DURING USE ON PATIENT WHILE SUTURING OR BEFORE USE ON PATIENT WHILE DISPENSING/HANDLING THE DEVICE? NO FURTHER INFORMATION AVAILABLE ANY SAMPLE AVAILABLE TO BE RETURNED FOR EVALUATION? THE DEVICE IS NOT AVAILABLE FOR ANALYSIS NOTE: EVENT REPORTED IN 2210968-2019-83746, 2210968-2019-83747 AND 2210968-2019-83748.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8769068&pc=GAM | 2210968-2019-83746 | 2018-06-04 | Malfunction | ETHICON INC. | 2019-07-08 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2018 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF FROM THE THREAD. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE EVENT OF "THE NEEDLE PULLED OFF FROM THE THREAD 3 TIMES" OCCUR DURING USE ON PATIENT WHILE SUTURING OR BEFORE USE ON PATIENT WHILE DISPENSING/HANDLING THE DEVICE? NO FURTHER INFORMATION AVAILABLE. ANY SAMPLE AVAILABLE TO BE RETURNED FOR EVALUATION? THE DEVICE IS NOT AVAILABLE FOR ANALYSIS. NOTE: EVENT REPORTED IN 2210968-2019-83747 AND 2210968-2019-83748.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8769499&pc=GAM | 2210968-2019-83755 | 2019-06-12 | Malfunction | ETHICON INC. | 2019-07-08 | GAM | VICRYL VIO 75CM M3 USP2/0 SGLE ARMED CPA | Break | Device Embedded In Tissue or Plaque | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING THE PROCEDURE. A RADIOLOGY WAS PERFORMED POST OP AND THE NEEDLE IS PLATED AGAINST THE PACEMAKER CASE. THE NEEDLE PIECE IS INSIDE THE PATIENT. THE NEEDLE WAS LEFT IN PLACE AND TWO RADIOLOGICAL CONTROL AT 15 DAYS AND 1 MONTH HAVE BEEN PLANNED. IT WAS REPORTED THAT THE PATIENT IS GOOD FOR THE MOMENT, A RADIOLOGICAL MONITORING IS PROGRAMMED BY THE SURGEON TO CHECK THAT THE NEEDLE PIECE DOESN'T MOVE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE MJ8BSSQ0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS RECEIVED: THE DEVICE IS NOT AVAILABLE FOR ANALYSIS : ONE PART OF THE DEVICE IS INSIDE THE PATIENT AND THE OTHER HAS BEEN DISCARDED THE PATIENT IS GOOD FOR THE MOMENT, A RADIOLOGICAL MONITORING IS PROGRAMMED BY THE SURGEON TO CHECK THAT THE NEEDLE PIECE DOESN'T MOVE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DID THE PATIENT SUFFER ANY ADVERSE CONSEQUENCES? IF, YES WHAT MEDICAL/ SURGICAL INTERVENTION WAS PROVIDED. ARE ANY PLANS IN PLACE TO REMOVE THE NEEDLE PIECE? PATIENT'S PRESENT CONDITION UPDATED US ON FINDINGS AT DAY 15 AND 1 MONTH. IS DEVICE AVAILABLE TO BE RETURNED FOR EVALUATION? IF YES, IS IT THE ACTUAL BROKEN NEEDLE OR REPRESENTATIVE/STERILE? PLEASE PROVIDE DEVICE RETURN STATUS OF REPRESENTATIVE SAMPLES. UPDATE QUESTIONS: WHAT WAS THE SIZE OF THE NEEDLE USED? WAS MORE THAN HALF THE NEEDLE RETAINED IN THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8755253&pc=GAM | 2210968-2019-83605 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-02 | GAM | VICRYL 3-0 1X75CM FS-1 VIOLET | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF FROM THE THREAD AFTER A FEW TISSUE PASSAGES, WITHOUT EXERTING EXCESSIVE TENSION ON THE THREAD. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8755307&pc=GAM | 2210968-2019-83607 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-02 | GAM | VICRYL 3-0 1X75CM FS-1 VIOLET | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF FROM THE THREAD AFTER A FEW TISSUE PASSAGES, WITHOUT EXERTING EXCESSIVE TENSION ON THE THREAD. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8755348&pc=GAM | 2210968-2019-83608 | 2019-01-01 | Malfunction | ETHICON INC. | 2019-07-02 | GAM | VICRYL 3-0 1X75CM FS-1 VIOLET | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF FROM THE THREAD AFTER A FEW TISSUE PASSAGES, WITHOUT EXERTING EXCESSIVE TENSION ON THE THREAD. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8748933&pc=GAM | 2210968-2019-83487 | 2019-05-01 | Malfunction | ETHICON INC. | 2019-07-01 | GAM | VICRYL VIO 90CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT A FEMALE CAT UNDERWENT AN OOPHORECTOMY ON (B)(6) 2019 AND THE SUTURE WAS USED ON ABDOMINAL WALL AND SUBCUTANEOUS TISSUES. THE NEEDLE PULLED OFF FROM THE THREAD AFTER SEVERAL TISSUE PASSAGE WITHOUT EXERTING EXCESSIVE TENSION. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AND BATCH NUMBER, AND NO NON-CONFORMANCE'S RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8748972&pc=GAM | 2210968-2019-83489 | 2019-05-01 | Malfunction | ETHICON INC. | 2019-07-01 | GAM | VICRYL VIO 90CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE CAT UNDERWENT AN OOPHORECTOMY ON (B)(6) 2019 AND THE SUTURE WAS USED ON ABDOMINAL WALL AND SUBCUTANEOUS TISSUES. THE NEEDLE PULLED OFF DURING THE EXTRACTION OF THE THREAD FROM THE SUPPORT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AND BATCH NUMBER, AND NO NON-CONFORMANCE'S RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8749007&pc=GAM | 2210968-2019-83491 | 2019-05-01 | Malfunction | ETHICON INC. | 2019-07-01 | GAM | VICRYL VIO 90CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE CAT UNDERWENT AN OOPHORECTOMY ON (B)(6) 2019 AND THE SUTURE WAS USED ON ABDOMINAL WALL AND SUBCUTANEOUS TISSUES. THE NEEDLE PULLED OFF DURING THE EXTRACTION OF THE THREAD FROM THE SUPPORT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AND BATCH NUMBER, AND NO NON-CONFORMANCE'S RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8749345&pc=GAM | 2210968-2019-83497 | 2019-05-01 | Malfunction | ETHICON INC. | 2019-07-01 | GAM | VICRYL VIO 90CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OOPHORECTOMY PROCEDURE IN (B)(6) 2019 AND THE SUTURE WAS USED FOR SUTURING ABDOMINAL WALL AND SUBCUTANEOUS TISSUES. DURING THE PROCEDURE, THE NEEDLE PULLED OFF FROM THE THREAD AFTER SEVERAL TISSUE PASSAGE WITHOUT EXERTING EXCESSIVE TENSION. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCE'S RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8749417&pc=GAM | 2210968-2019-83498 | 2019-05-01 | Malfunction | ETHICON INC. | 2019-07-01 | GAM | VICRYL VIO 90CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OOPHORECTOMY PROCEDURE IN (B)(6) 2019 AND THE SUTURE WAS USED FOR SUTURING ABDOMINAL WALL AND SUBCUTANEOUS TISSUES. DURING THE PROCEDURE, THE NEEDLE PULLED OFF FROM THE THREAD AFTER SEVERAL TISSUE PASSAGE WITHOUT EXERTING EXCESSIVE TENSION. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
The MAUDE database houses medical device reports submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers. Each year | the FDA receives several hundred thousand medical device reports (MDRs) of suspected device-associated deaths | serious injuries and malfunctions. The FDA uses MDRs to monitor device performance | detect potential device-related safety issues | and contribute to benefit-risk assessments of these products. The MAUDE database houses MDRs submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers.
Although MDRs are a valuable source of information | this passive surveillance system has limitations | including the potential submission of incomplete | inaccurate | untimely | unverified | or biased data. In addition | the incidence or prevalence of an event cannot be determined from this reporting system alone due to potential under-reporting of events and lack of information about frequency of device use. Because of this | MDRs comprise only one of the FDA's several important postmarket surveillance data sources.
MDR data alone cannot be used to establish rates of events | evaluate a change in event rates over time or compare event rates between devices. The number of reports cannot be interpreted or used in isolation to reach conclusions about the existence | severity | or frequency of problems associated with devices. Confirming whether a device actually caused a specific event can be difficult based solely on information provided in a given report. Establishing a cause-and-effect relationship is especially difficult if circumstances surrounding the event have not been verified or if the device in question has not been directly evaluated.
MAUDE data does not represent all known safety information for a reported medical device and should be interpreted in the context of other available information when making device-related or treatment decisions. Variations in trade | product | and company names affect search results. Searches only retrieve records that contain the search term(s) provided by the requester.Submission of a medical device report and the FDA's release of that information is not necessarily an admission that a product | user facility | importer | distributor | manufacturer | or medical personnel caused or contributed to the event. Certain types of report information are protected from public disclosure under the Freedom of Information Act (FOIA).
If a report contains trade secret or confidential business information | that text is replaced by "(b)(4)". If a report contains personnel or medical files information | that text is replaced by "(b)(6)". The designations "(b)(4)" and "(b)(6)" refer to the exemptions in the FOIA. For example | "(b)(4)" may be found in place of the product's composition and "(b)(6)" may be found in place of a patient's age. MAUDE is updated monthly and the search page reflects the date of the most recent update. The FDA seeks to include all reports received prior to the update but the inclusion of some reports may be delayed.
Manufacturers and importers must submit reports when they become aware of information that reasonably suggests that one of their marketed devices may have caused or contributed to a death or serious injury or has malfunctioned and the malfunction of the device or a similar device that they market would be likely to cause or contribute to a death or serious injury if the malfunction were to recur. Manufacturers must send reports of such deaths | serious injuries and malfunctions to the FDA. Importers must send reports of deaths and serious injuries to the FDA and the manufacturer | and reports of malfunctions to the manufacturer.
Device user facilities include hospitals | outpatient diagnostic or treatment facilities | nursing homes and ambulatory surgical facilities. Device user facilities must submit reports when they become aware of information that reasonably suggests that a device may have caused or contributed to a death or serious injury of a patient in their facility. Death reports must be sent to the FDA and the manufacturer | if known. Serious injury reports must be sent to the manufacturer or to the FDA | if the manufacturer is not known.