## Sheet1
| Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11072015&pc=GAM | 2210968-2020-10292 | 2020-10-22 | Malfunction | ETHICON INC. | 2020-12-24 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # ==> PC (B)(4). COMPONENT CODE: G07002 - DEVICE NOT RETURNED IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE NAME OF THE PROCEDURE? WHAT WAS BEING DONE WHEN THE NEEDLE PULLED-OFF (IN THE PACKAGE/ REMOVAL FROM PACKAGE / DURING PASSAGE THROUGH TISSUE ¿)? WAS THE NEEDLE PIECE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? WHAT TISSUE/LOCATION WAS SUTURING WHEN PULL OFF OCCURRED? WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THIS SUTURE NEEDLE PULL OFF EVENT? A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11072098&pc=GAM | 2210968-2020-10293 | 2020-12-02 | Malfunction | ETHICON INC. | 2020-12-24 | GAM | VICRYL CT VIO 75CM M3.5 | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A UTERO PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS USED TO COMPLETE THE PROCEDURE? ANOTHER NEEDLE. YES BY USING ANOTHER NEEDLE. DID THE SURGERY WAS COMPLETED SUCCESSFULLY? DID THE PATIENT SUFFER FROM ANY CONSEQUENCE DUE TO THE ISSUE (BREAKAGE NEEDLE)? IF YES PLEASE EXPLAIN. NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065156&pc=GAM | 2210968-2020-10243 | 2020-12-05 | Malfunction | ETHICON INC. | 2020-12-23 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TAH PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE THREAD WAS NOT STRONG AND GOT CUT IN BETWEEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? NO, ACTION TAKEN WHEN EVENT OCCURRED? THEY HAVE USED SAME SUTURES WITH DIFFERENT BATCH NUMBER., NO COMPLICATIONS WERE REPORTED. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065167&pc=GAM | 2210968-2020-10244 | 2020-12-05 | Malfunction | ETHICON INC. | 2020-12-23 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TAH PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE THREAD WAS NOT STRONG AND GOT CUT IN BETWEEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). DEVICE NOT RETURNED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? NO, ACTION TAKEN WHEN EVENT OCCURRED? THEY HAVE USED SAME SUTURES WITH DIFFERENT BATCH NUMBER., NO COMPLICATIONS WERE REPORTED. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065169&pc=GAM | 2210968-2020-10245 | 2020-12-05 | Malfunction | ETHICON INC. | 2020-12-23 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TAH PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE THREAD WAS NOT STRONG AND GOT CUT IN BETWEEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? NO, ACTION TAKEN WHEN EVENT OCCURRED? THEY HAVE USED SAME SUTURES WITH DIFFERENT BATCH NUMBER., NO COMPLICATIONS WERE REPORTED. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11066249&pc=GAM | 2210968-2020-10254 | 2020-12-09 | Malfunction | ETHICON INC. | 2020-12-23 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MWR-2210968-2020-10256. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW THE PROCEDURE WAS COMPLETE? ANY PATIENT CONSEQUENCES? PLEASE PROVIDE PROCEDURE NAME AND DATE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11066257&pc=GAM | 2210968-2020-10256 | 2020-12-09 | Malfunction | ETHICON INC. | 2020-12-23 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MWR-2210968-2020-10254. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW THE PROCEDURE WAS COMPLETE? ANY PATIENT CONSEQUENCES? PLEASE PROVIDE PROCEDURE NAME AND DATE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11059032&pc=GAM | 2210968-2020-10198 | 2020-10-09 | Malfunction | ETHICON INC. | 2020-12-22 | GAM | VICRYL* UND 3-0 70CM (1)SC-20 | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT REPLACEMENT AND CORRECTION OF ABDOMINAL AND UMBILICAL SCAR ON (B)(6) 2020 AND SUTURE WAS USED. WHEN CORRECTING THE NAVAL SCAR, A STITCH HAD TO BE PASSED TO THE FASCIA TO ENCIRCLE THE NAVEL. THIS STITCH WAS MADE WITH A SUTURE, WHICH DISASSEMBLED FROM THE BONE SUTURE THREAD, LEAVING THE NEEDLE IN THE CAVITY. THE NEEDLE WAS CAREFULLY SEARCHED FOR 30 MINUTES UNTIL IT WAS FOUND IN THE RECTUS ABDOMINIS CAVITY. THERE WAS NO REPORTED ADVERSE PATIENT CONSEQUENCES. IT WAS REPORTED THAT A PROBABLE CAUSE OF THE EVENT WAS HAVING TO ACCESS DIFFERENT STRUCTURES TO LOOK FOR THE NEEDLE, EVERY TIME IT MOVED. THE NEEDLE WAS FOUND AND ANOTHER SUTURE WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE PATIENT¿S POST-OPERATIVE CARE ALTERED DUE TO THE EVENT AND PROLONGED SURGERY? IF YES, PLEASE DESCRIBE. WHAT IS THE PATIENT'S CURRENT STATUS? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. WHAT IS THE PATIENT¿S CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11061311&pc=GAM | 2210968-2020-10228 | 2020-12-01 | Malfunction | ETHICON INC. | 2020-12-22 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN (B)(6) 2020 AND SUTURE WAS USED. DURING SURGERY, THE SUTURE BROKE OFF. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). BATCH MANUFACTURING RECORDS OF T8040 WAS REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE FOUND TO MEET THE SPECIFICATION REQUIREMENT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE INITIAL PROCEDURE? WHAT WAS USED TO COMPLETE THE PROCEDURE? DID THE SURGERY WAS COMPLETED SUCCESSFULLY? WHAT IS THE EVENT DAY AND PROCEDURE DATE? DID THE PATIENT SUFFER FROM ANY CONSEQUENCE DUE TO THE ISSUE (BREAKAGE SUTURE)? COULD YOU PLEASE CONFIRM HOW MANY DEVICES PRESENT THE ISSUE (BREAKAGE SUTURE)? NOTE: EVENTS REPORTED IN 2210968-2020-10229. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11061314&pc=GAM | 2210968-2020-10229 | 2020-12-01 | Malfunction | ETHICON INC. | 2020-12-22 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN (B)(6) 2020 AND SUTURE WAS USED. DURING SURGERY, THE SUTURE BROKE OFF. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). BATCH MANUFACTURING RECORDS OF T8040 WAS REVIEWED FOR ANY PROCESS DEVIATION BUT NO DEVIATION WAS OBSERVED. FINISHED GOODS TESTS REPORTS WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE FOUND TO MEET THE SPECIFICATION REQUIREMENT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE INITIAL PROCEDURE? WHAT WAS USED TO COMPLETE THE PROCEDURE? DID THE SURGERY WAS COMPLETED SUCCESSFULLY? WHAT IS THE EVENT DAY AND PROCEDURE DATE? DID THE PATIENT SUFFER FROM ANY CONSEQUENCE DUE TO THE ISSUE (BREAKAGE SUTURE)? COULD YOU PLEASE CONFIRM HOW MANY DEVICES PRESENT THE ISSUE (BREAKAGE SUTURE)? NOTE: EVENTS REPORTED IN 2210968-2020-10228. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11051537&pc=GAM | 2210968-2020-10149 | 2020-12-03 | Malfunction | ETHICON INC. | 2020-12-21 | GAM | VICRYL*RAPID UND 3-0 45CM 1PS2 | Material Separation | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN ORTHOPEDIC PROCEDURE ON (B)(6) 2020 AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY: IT WAS REPORTED BY THE ACCOUNT CONTACT THAT DURING A DENTAL SURGERY THAT TWO OF THE SUTURES IMMEDIATELY SEPARATED FROM THE NEEDLE. AN EMPTY OPENED BOX, TWO EMPTY LABELED WINDING FORMER AND FIVE UNOPENED SAMPLES OF PRODUCT CODE VR497, LOT # AM7267 WERE RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF FIVE UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. AS PART OF OUR QUALITY PROCESS, THE MANUFACTURING RECORDS OF THIS LOT-SERIAL NUMBER WERE REVIEWED, AND THE MANUFACTURING STANDARDS WERE MET PRIOR TO THE RELEASE OF THIS BATCH. THE DEVICE PERFORMED WITHOUT ANY DEFECT NOTED AND THE ACTUAL COMPLAINT SAMPLE WAS NOT RETURNED FOR ANALYSIS. COMPLAINT INFORMATION WILL BE INCLUDED IN COMPLAINT TRENDING THAT IS REVIEWED ON A REGULAR BASIS TO DETERMINE IF FURTHER ACTION IS NECESSARY. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MWR-2210968-2020-10148. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER AM7267, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11051550&pc=GAM | 2210968-2020-10148 | 2020-12-03 | Malfunction | ETHICON INC. | 2020-12-21 | GAM | VICRYL*RAPID UND 3-0 45CM 1PS2 | Material Separation | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN ORTHOPEDIC PROCEDURE ON (B)(6) 2020 AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MWR-2210968-2020-10148 AND 2210968-2020-10149. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER AM7267, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11052215&pc=GAM | 2210968-2020-10169 | 2020-12-01 | Malfunction | ETHICON INC. | 2020-12-21 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ANY ADVERSE EVENT CONSEQUENCES?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11036876&pc=GAM | 2210968-2020-10092 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-18 | GAM | VICRYL CT VIO 90CM M5 | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE THE NEEDLE WAS COMING OUT OF THE SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS NEEDLE COMING OUT FROM SUTURE DURING PROCEDURE OR WAS FOUND LIKE THIS IN PACKAGE? WHAT PROCEDURE WAS BEING PERFORMED? HOW MANY MINUTES WAS THE PROCEDURE DELAYED? WAS THERE ANY PATIENT CONSEQUENCES? CAN YOU CLARIFY THE LOT? STATUS OF DEVICE RETURN/FOLLOW-UP. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11030382&pc=GAM | 2210968-2020-10044 | 2020-11-24 | Malfunction | ETHICON INC. | 2020-12-17 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 4-0 UND | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HAND PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THE BROKEN NEEDLE FOUND AND REMOVE FROM THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). DEVICE NOT RETURNED. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? WHAT IS CURRENT CONDITION OF THE PATIENT? WHAT TISSUE STRUCTURE WAS IT LOCATED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11015647&pc=GAM | 2210968-2020-09953 | 2020-11-24 | Malfunction | ETHICON INC. | 2020-12-16 | GAM | VICRYL CT VIO 90CM M3.5 | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE FOR REMOVAL OF EXPANDER IMPLANT FOR DEFINITIVE IMPLANT FOLLOWING A CANCER ON THE RIGHT BREAST WITH MASTECTOMY ON (B)(6) 2020 AND SUTURE WAS USED. DURING SUTURING, THE NEEDLE BROKE 1/3 DOWNSTREAM FROM THE CRIMPING ZONE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE OTHER NEEDLE PIECE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? A PIECE OF NEEDLE IS FALLEN INSIDE THE PATIENT BUT HAS BEEN RECOVERED WITH A IMAGE INTENSIFIER. THE PIECE OF NEEDLE HAS BEEN REMOVED DURING THE SAME PROCEDURE. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THIS NEEDLE PIECE SEARCH? NO. WAS THERE ANY PATIENT CONSEQUENCE OR INJURY? ADDITIONAL X-RAY EXPOSURE DUE TO NEEDLE FINDING WITH IMAGE INTENSIFIER. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11015742&pc=GAM | 2210968-2020-09952 | 2020-11-24 | Malfunction | ETHICON INC. | 2020-12-16 | GAM | VICRYL CT VIO 90CM M3.5 | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE FOR REMOVAL OF EXPANDEUR IMPLANT FOR DEFINITIVE IMPLANT FOLLOWING A CANCER ON THE RIGHT BREAST WITH MASTECTOMY ON (B)(6) 2020 AND SUTURE WAS USED. DURING SUTURING, THE NEEDLE BROKE 1/3 DOWNSTREAM FROM THE CRIMPING ZONE. THE NEEDLE WAS LOST IN THE OPERATIVE SITE BUT FOUND USING THE IMAGE INTENSIFIER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). COMPONENT CODE: (B)(4) - DEVICE NOT RETURNED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THE OTHER NEEDLE PIECE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE?: A PIECE OF NEEDLE IS FALLEN INSIDE THE PATIENT BUT HAS BEEN RECOVERED WITH A IMAGE INTENSIFIER. THE PIECE OF NEEDLE HAS BEEN REMOVED DURING THE SAME PROCEDURE. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THIS NEEDLE PIECE SEARCH? NO. WAS THERE ANY PATIENT CONSEQUENCE OR INJURY? : ADDITIONAL X-RAY EXPOSURE DUE TO NEEDLE FINDING WITH IMAGE INTENSIFIER A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11016354&pc=GAM | 2210968-2020-09970 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-16 | GAM | VICRYL CT VIO 90CM M5 | Material Frayed; Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: NEEDLE SEPARATED FROM SUTURE DURING TISSUE PASSAGE. THE SUTURE APPEARS TO BE UNRAVELED OR FRAYED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: AN OPENED BOX WITH EIGHT UNOPENED SAMPLES OF PRODUCT CODE V372H, LOT #QJBBDWR0 WERE RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE DEVICE PERFORMED WITHOUT ANY DEFECT NOTED AND THE ACTUAL COMPLAINT SAMPLE WAS NOT RETURNED FOR ANALYSIS. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY, HOW MANY DEVICES ARE INVOLVED OF EACH LOT NUMBER? CUSTOMER IS UNABLE TO SPECIFY EXACT NUMBER. HOW THE PROCEDURE WAS COMPLETE? UNKNOWN. ANY PATIENT CONSEQUENCES NO HARM TO PATIENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA 2210968-2020-09964, 2210968-2020-09966 AND 2210968-2020-09968.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11016428&pc=GAM | 2210968-2020-09966 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-16 | GAM | VICRYL CT VIO 90CM M5 | Material Frayed; Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL EXTRAPERITONEAL (TEP) REPAIR PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE SEPARATED FROM SUTURE DURING TISSUE PASSAGE. IT WAS ALSO REPORTED THAT THE SUTURE APPEARS TO BE UNRAVELED OR FRAYED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). COMPONENT CODE: G07002 ¿ CONFORMING DEVICE A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH QDBDSGR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: A SEALED BOX AND AN OPENED BOX WITH TEN UNOPENED SAMPLES OF PRODUCT CODE V372H, LOT # QDBDSGR0 WERE RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE DEVICE PERFORMED WITHOUT ANY DEFECT NOTED AND THE ACTUAL COMPLAINT SAMPLE WAS NOT RETURNED FOR ANALYSIS. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY, HOW MANY DEVICES ARE INVOLVED OF EACH LOT NUMBER? CUSTOMER IS UNABLE TO SPECIFY EXACT NUMBER. HOW THE PROCEDURE WAS COMPLETE? UNKNOWN. ANY PATIENT CONSEQUENCES NO HARM TO PATIENT. EVENTS WERE SUBMITTED VIA 2210968-2020-09964, 2210968-2020-09968 AND 2210968-2020-09970.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11016572&pc=GAM | 2210968-2020-09968 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-16 | GAM | VICRYL CT VIO 90CM M5 | Material Frayed; Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL EXTRAPERITONEAL (TEP) REPAIR PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE SEPARATED FROM SUTURE DURING TISSUE PASSAGE. IT WAS ALSO REPORTED THAT THE SUTURE APPEARS TO BE UNRAVELED OR FRAYED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4)3 A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. AN OPENED BOX WITH EIGHT UNOPENED SAMPLES OF PRODUCT CODE V372H, LOT #QJBBDWR0 WERE RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE DEVICE PERFORMED WITHOUT ANY DEFECT NOTED AND THE ACTUAL COMPLAINT SAMPLE WAS NOT RETURNED FOR ANALYSIS. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: *PLEASE CLARIFY, HOW MANY DEVICES ARE INVOLVED OF EACH LOT NUMBER? CUSTOMER IS UNABLE TO SPECIFY EXACT NUMBER. *HOW THE PROCEDURE WAS COMPLETE? UNKNOWN *ANY PATIENT CONSEQUENCES NO HARM TO PATIENT. EVENTS WERE SUBMITTED VIA 2210968-2020-09964, 2210968-2020-09966, 2210968-2020-09968 AND 2210968-2020-09970.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11016581&pc=GAM | 2210968-2020-09964 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-16 | GAM | VICRYL CT VIO 90CM M5 | Material Frayed; Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL EXTRAPERITONEAL (TEP) REPAIR PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE NEEDLE SEPARATED FROM SUTURE DURING TISSUE PASSAGE. IT WAS ALSO REPORTED THAT THE SUTURE APPEARS TO BE UNRAVELED OR FRAYED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH QDBDSGR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: A SEALED BOX AND AN OPENED BOX WITH THIRTEEN UNOPENED SAMPLES OF PRODUCT CODE V372H, LOT # QDBDSGR0 WERE RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE DEVICE PERFORMED WITHOUT ANY DEFECT NOTED AND THE ACTUAL COMPLAINT SAMPLE WAS NOT RETURNED FOR ANALYSIS. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY, HOW MANY DEVICES ARE INVOLVED OF EACH LOT NUMBER? CUSTOMER IS UNABLE TO SPECIFY EXACT NUMBER. HOW THE PROCEDURE WAS COMPLETE? UNKNOWN ANY PATIENT CONSEQUENCES NO HARM TO PATIENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA 2210968-2020-09966, 2210968-2020-09968 AND 2210968-2020-09970.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11023333&pc=GAM | 2210968-2020-10011 | 2020-11-25 | Malfunction | ETHICON INC. | 2020-12-16 | GAM | VICRYL VIO 75CM M1.5 | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT A MASTECTOMY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. AFTER A FEW TISSUE PASSAGES IN THE SUBCUTANEOUS TISSUE, THE NEEDLE PULLED OFF, WITHOUT EXCESSIVE TENSION. THE PROCEDURE WAS COMPLETED USING A NEW LIKE DEVICE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11023510&pc=GAM | 2210968-2020-10012 | 2020-11-25 | Malfunction | ETHICON INC. | 2020-12-16 | GAM | VICRYL VIO 90CM M4 | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESAREAN, HYSTEROGRAPHY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. THE NEEDLE PULLED OFF FROM THE THREAD AND FELL INSIDE THE LOWER INCISED SEGMENT OF THE UTERUS. THE SURGEON WAS ABLE TO RECOVER IT IMMEDIATELY. THE PROCEDURE WAS FINISHED WITH SUCCESS AN THERE WERE NO PATIENT CONSEQUENCES FOLLOWING THE ISSUE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). COMPONENT CODE: (B)(4) - DEVICE NOT RETURNED ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE SURGERY WAS COMPLETED SUCCESSFULLY?: YES, THE PROCEDURE HAS BEEN FINISHED WITH SUCCESS. NO PATIENT CONSEQUENCES FOLLOWING THE ISSUE. THE NEEDLE HAS BEEN RECOVERED IMMEDIATELY. DID THE PATIENT SUFFER FROM ANY CONSEQUENCE DUE TO THE RECOVERY OF THE NEEDLE?: NO. BUT STRESS FOR THE SURGEON. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11003740&pc=GAM | 2210968-2020-09893 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-14 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN AESTHETIC PROCEDURE ON AN UNKNOWN DATE IN 2020 AND A SUTURE WAS USED. DURING EVALUATION, ADDITIONAL DETACHED NEEDLE WAS OBSERVED; THE SUTURE WAS NOT SENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). DATE SENT TO THE FDA: 12/14/2020. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER AM7405, AND NO NON- CONFORMANCES / MANUFACTURING IRREGULARITIES WERE IDENTIFIED. A DETACHED NEEDLE OF PRODUCT CODE VR2296, LOT # AM7405 WAS RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF A DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED, MARKS BY USE OF SURGICAL INSTRUMENT WERE NOTED ON BODY NEEDLE. THE SUTURE WAS NOT RECEIVED TO DETERMINE THE ASSIGNABLE CAUSE AND NO FUNCTIONAL TESTS COULD BE PERFORMED. NO CONCLUSION COULD BE REACHED AS THE NEEDLES OF THE BOX PULLED OFF. IT SHOULD BE NOTED THAT AS PART OF OUR QUALITY PROCESS, EACH DEVICE IS VISUALLY INSPECTED AND FUNCTIONALLY TESTED DURING MANUFACTURING TO ENSURE THE DEVICE MEETS THE REQUIRED SPECIFICATIONS PRIOR TO SHIPMENT. EVALUATION OF RECEIVED 1ST NEEDLE WAS SUBMITTED VIA 2210968-2020-08129. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE PROVIDE PROCEDURE NAME AND DATE - AESTHETIC SURGERY : SEVERAL PROCEDURE IMPACTED BUT EXACT NUMBER UNKNOWN. HOW LONG WAS THE DELAY? YES, 15 MIN. SEVERAL NEEDLES BY PROCEDURE (EXACT NUMBER UNKNOWN). WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? NO. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN - NO. PATIENT¿S CURRENT STATUS? GOOD. A NEW COMPLAINT HAS BEEN OPENED BECAUSE SEVERAL PROCEDURE IMPACTED. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL EVENT REPORTED VIA MW # 2210968-2020-08830. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10992581&pc=GAM | 2210968-2020-09859 | 2020-11-16 | Malfunction | ETHICON INC. | 2020-12-11 | GAM | VICRYL CT VIO 90CM M5 | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT KNEE SURGERY ON (B)(6) 2020 AND SUTURE WAS USED. THE NEEDLE LOOSENED FROM THE THREAD INTRA-OP. THE PROCEDURE WAS COMPLETED WITH A LIKE DEVICE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4) A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V1059H; QHBDJPQ0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10985364&pc=GAM | 2210968-2020-09815 | 2020-11-12 | Malfunction | ETHICON INC. | 2020-12-10 | GAM | VICRYL UND 45CM M2 | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: NEEDLE PULL OFF. PROCEDURE COMPLETE WITH OPENING OF A NEW SUTURE (NEEDLE PULL OFF HAPPENED TWICE). PLASTIC SURGERY BUT PROCEDURE NAME AND DATE ARE UNKNOWN. NO ADVERSE CONSEQUENCE ASSOCIATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10985426&pc=GAM | 2210968-2020-09816 | 2020-11-12 | Malfunction | ETHICON INC. | 2020-12-10 | GAM | VICRYL UND 45CM M2 | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: NEEDLE PULL OFF PROCEDURE COMPLETE WITH OPENING OF A NEW SUTURE (NEEDLE PULL OFF HAPPENED TWICE) PLASTIC SURGERY BUT PROCEDURE NAME AND DATE ARE UNKNOWN NO ADVERSE CONSEQUENCE ASSOCIATED
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10987420&pc=GAM | 2210968-2020-09823 | 2020-10-24 | Malfunction | ETHICON INC. | 2020-12-10 | GAM | VICRYL PLUS VIO 70CM::M2 USP3/0 SGLE ARMED SH-1 | Material Separation | No Clinical Signs, Symptoms or Conditions | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10987786&pc=GAM | 2210968-2020-09830 | 2020-10-30 | Malfunction | ETHICON INC. | 2020-12-10 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE UNKNOWN PROD | Break | No Clinical Signs, Symptoms or Conditions | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE TINY TIP OF FS-NEEDLE RETURNED BROKEN FROM SURGEON TO SCRUB NURSE, TIP WHEREABOUTS UNKNOWN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4) TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10959753&pc=GAM | 2210968-2020-09639 | 2020-10-10 | Malfunction | ETHICON INC. | 2020-12-07 | GAM | VICRYL RAP UND 90CM M3 USP2/0 SGLE ARMED V-34 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT PROCEDURE FOR EXCISION OF SKIN NEOPLASMS ON (B)(6) 2020 AND SUTURE WAS USED. THE SUTURE BROKE WHEN SEWING DURING THE EXCISION OF SKIN NEOPLASMS ON THE BACK. CHANGED TO ANOTHER DEVICE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). DEVICE NOT RETURNED A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE EMW9962; QDCCCXB0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10963688&pc=GAM | 2210968-2020-09692 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-07 | GAM | VICRYL CT VIO 90CM M5 | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A COW UNDERWENT A CESAREAN SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE SUTURING, WHEN THE VETERINARY PULL THE THREAD WITHOUT EXCESSIVE TENSION AFTER THE PASSAGE ON THE SKIN, THE DEVICE BROKE IN THE LENGTH OF THE THREAD. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: HOW MANY MEDICAL DEVICES WERE IMPACTED FOR THIS SPECIFIC CASE? HOW MANY PROCEDURES WERE AFFECTED BY THIS ISSUE? WHERE THE OTHER EVENTS REPORTED? (B)(4) (2210968-2020-09694). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10963689&pc=GAM | 2210968-2020-09694 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-07 | GAM | VICRYL CT VIO 90CM M5 | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A COW UNDERWENT A CESAREAN SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE SUTURING, WHEN THE VETERINARY PULL THE THREAD WITHOUT EXCESSIVE TENSION AFTER THE PASSAGE ON THE SKIN, THE DEVICE BROKE IN THE LENGTH OF THE THREAD. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: HOW MANY MEDICAL DEVICES WERE IMPACTED FOR THIS SPECIFIC CASE? HOW MANY PROCEDURES WERE AFFECTED BY THIS ISSUE? WHERE THE OTHER EVENTS REPORTED? (B)(4) ( 2210968-2020-09692). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10945705&pc=GAM | 2210968-2020-09571 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-04 | GAM | VICRYL CT VIO 90CM M5 | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT KNEE SURGERY ON AN UNKNOWN DATE AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE LOOSENED FROM THE THREAD. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10945744&pc=GAM | 2210968-2020-09572 | 2020-11-11 | Malfunction | ETHICON INC. | 2020-12-04 | GAM | VICRYL CT VIO 90CM M5 | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2020 AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE LOOSENED FROM THE THREAD. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PC (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER V1059H; QHBDJPQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10939573&pc=GAM | 2210968-2020-09545 | 2020-11-05 | Malfunction | ETHICON INC. | 2020-12-03 | GAM | VICRYL RAP UND 90CM M3 2/0 V-34 | Material Separation | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EPISIOTOMY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. THE NEEDLE DETACHED FROM THE THREAD WITH VERY MINIMAL TENSION DURING SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 12/02/2020. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? NO, WAS PROCEDURE SUCCESSFULLY COMPLETED? YES, WERE FRAGMENTS GENERATED? NO, PATIENT STATUS/ OUTCOME / CONSEQUENCES? NO. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10934913&pc=GAM | 2210968-2020-09509 | 2020-11-11 | Malfunction | ETHICON INC. | 2020-12-02 | GAM | VICRYL*RAPID UND 4-0 45CM 1PS2 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THERE WAS A FRAILTY RUPTURE IN THE PROXIMAL SEGMENT AND THE MIDDLE SEGMENT OF THE WIRE.. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MEDWATCH REPORTS 2210968-2020-09505, 2210968-2020-09505, 2210968-2020-09506, 2210968-2020-09507, 2210968-2020-09508, 2210968-2020-09510. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10934916&pc=GAM | 2210968-2020-09507 | 2020-11-11 | Malfunction | ETHICON INC. | 2020-12-02 | GAM | VICRYL*RAPID UND 4-0 45CM 1PS2 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THERE WAS A FRAILTY RUPTURE IN THE PROXIMAL SEGMENT AND THE MIDDLE SEGMENT OF THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MEDWATCH REPORTS 2210968-2020-09506, 2210968-2020-09508,2210968-2020-09509, 2210968-2020-09510, 2210968-2020-09511, 2210968-2020-09512 TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10934933&pc=GAM | 2210968-2020-09510 | 2020-11-11 | Malfunction | ETHICON INC. | 2020-12-02 | GAM | VICRYL*RAPID UND 4-0 45CM 1PS2 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THERE WAS A FRAILTY RUPTURE IN THE PROXIMAL SEGMENT AND THE MIDDLE SEGMENT OF THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MEDWATCH REPORTS 2210968-2020-09506, 2210968-2020-09507, 2210968-2020-09508,2210968-2020-09509, 2210968-2020-09511, 2210968-2020-09512. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10934972&pc=GAM | 2210968-2020-09508 | 2020-11-11 | Malfunction | ETHICON INC. | 2020-12-02 | GAM | VICRYL*RAPID UND 4-0 45CM 1PS2 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THERE WAS A FRAILTY RUPTURE IN THE PROXIMAL SEGMENT AND THE MIDDLE SEGMENT OF THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MEDWATCH REPORTS 2210968-2020-09506, 2210968-2020-09507, 2210968-2020-09509, 2210968-2020-09510, 2210968-2020-09511, 2210968-2020-09512. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10934975&pc=GAM | 2210968-2020-09512 | 2020-11-11 | Malfunction | ETHICON INC. | 2020-12-02 | GAM | VICRYL*RAPID UND 4-0 45CM 1PS2 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THERE WAS A FRAILTY RUPTURE IN THE PROXIMAL SEGMENT AND THE MIDDLE SEGMENT OF THE WIRE.. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO FDA: 12/02/2020. EVENTS REPORTED IN MEDWATCH REPORTS 2210968-2020-09506, 2210968-2020-09507, 2210968-2020-09508,2210968-2020-09509, 2210968-2020-09510, 2210968-2020-09511. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10934977&pc=GAM | 2210968-2020-09506 | 2020-11-11 | Malfunction | ETHICON INC. | 2020-12-02 | GAM | VICRYL*RAPID UND 4-0 45CM 1PS2 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THERE WAS A FRAILTY RUPTURE IN THE PROXIMAL SEGMENT AND THE MIDDLE SEGMENT OF THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MEDWATCH REPORTS 2210968-2020-09507, 2210968-2020-09508,2210968-2020-09509, 2210968-2020-09510, 2210968-2020-09511, 2210968-2020-09512 TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10935008&pc=GAM | 2210968-2020-09511 | 2020-11-11 | Malfunction | ETHICON INC. | 2020-12-02 | GAM | VICRYL*RAPID UND 4-0 45CM 1PS2 | Break | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THERE WAS A FRAILTY RUPTURE IN THE PROXIMAL SEGMENT AND THE MIDDLE SEGMENT OF THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). EVENTS REPORTED IN MEDWATCH REPORTS 2210968-2020-09506, 2210968-2020-09507, 2210968-2020-09508,2210968-2020-09509, 2210968-2020-09510, 2210968-2020-09512. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10936467&pc=FTL | 2210968-2020-09518 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-12-02 | FTL | VICRYL MESH 26.5X34CM | Delivered as Unsterile Product | No Clinical Signs, Symptoms or Conditions | K810428 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2020 AND THE MESH WAS USED. IT WAS REPORTED THAT BEFORE USING THE MESH, THE STERILE PACKAGE WAS DAMAGE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 12/02/2020. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10921158&pc=GAM | 2210968-2020-09440 | 2020-11-02 | Malfunction | ETHICON INC. | 2020-11-30 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A VAGINAL HYSTERECTOMY ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WERE THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? YES THEY RETRIEVED NEEDLE SAME TIME DURING PROCEDURE WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? NO PROCEDURE DATE? (B)(6) 2020. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10921267&pc=GAM | 2210968-2020-09439 | 2020-11-02 | Malfunction | ETHICON INC. | 2020-11-30 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A VAGINAL HYSTERECTOMY ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 11/30/2020. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WERE THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? YES, THEY RETRIEVED NEEDLE SAME TIME DURING PROCEDURE. WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? NO. PROCEDURE DATE? (B)(6) 2020. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10908102&pc=GAM | 2210968-2020-09340 | 2020-11-01 | Malfunction | ETHICON INC. | 2020-11-26 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. DURING THE SURGERY, DEVICE BREAKAGE FROM SWAGE POINT OCCURRED. THE PROCEDURE WAS DELAYED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4) A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DO YOU MEAN THE SUTURE DETACHED FROM THE NEEDLE OR SUTURE BREAKAGE OR NEEDLE BREAKAGE DURING THE PROCEDURE? WHAT WAS USED TO COMPLETE THE PROCEDURE? WHAT IS THE PATIENT'S CONDITION? PROCEDURE DATE? WERE THERE ANY PATIENT CONSEQUENCES DUE TO THE "BREAKAGE FROM SWAGE POINT" INTRA-OP ISSUE? DEVICE RETURN STATUS/FOLLOW UP? EVENTS WERE SUBMITTED VIA 2210968-2020-09339, 2210968-2020-09341 AND 2210968-2020-09344
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10908149&pc=GAM | 2210968-2020-09344 | 2020-11-01 | Malfunction | ETHICON INC. | 2020-11-26 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. DURING THE SURGERY, DEVICE BREAKAGE FROM SWAGE POINT OCCURRED. THE PROCEDURE WAS DELAYED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DO YOU MEAN THE SUTURE DETACHED FROM THE NEEDLE OR SUTURE BREAKAGE OR NEEDLE BREAKAGE DURING THE PROCEDURE? WHAT WAS USED TO COMPLETE THE PROCEDURE? WHAT IS THE PATIENT'S CONDITION? PROCEDURE DATE? WERE THERE ANY PATIENT CONSEQUENCES DUE TO THE "BREAKAGE FROM SWAGE POINT" INTRA-OP ISSUE? DEVICE RETURN STATUS/FOLLOW UP? EVENTS WERE SUBMITTED VIA 2210968-2020-09339, 2210968-2020-09340, 2210968-2020-09341.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10908308&pc=GAM | 2210968-2020-09339 | 2020-11-01 | Malfunction | ETHICON INC. | 2020-11-26 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. DURING THE SURGERY, DEVICE BREAKAGE FROM SWAGE POINT OCCURRED. THE PROCEDURE WAS DELAYED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DO YOU MEAN THE SUTURE DETACHED FROM THE NEEDLE OR SUTURE BREAKAGE OR NEEDLE BREAKAGE DURING THE PROCEDURE? WHAT WAS USED TO COMPLETE THE PROCEDURE? WHAT IS THE PATIENT'S CONDITION? PROCEDURE DATE? WERE THERE ANY PATIENT CONSEQUENCES DUE TO THE "BREAKAGE FROM SWAGE POINT" INTRA-OP ISSUE? DEVICE RETURN STATUS/FOLLOW UP? EVENTS WERE SUBMITTED VIA 2210968-2020-09340, 2210968-2020-09341 AND 2210968-2020-09344.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10908416&pc=GAM | 2210968-2020-09341 | 2020-11-01 | Malfunction | ETHICON INC. | 2020-11-26 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED CTX | Material Separation | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. DURING THE SURGERY, DEVICE BREAKAGE FROM SWAGE POINT OCCURRED. THE PROCEDURE WAS DELAYED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DO YOU MEAN THE SUTURE DETACHED FROM THE NEEDLE OR SUTURE BREAKAGE OR NEEDLE BREAKAGE DURING THE PROCEDURE? WHAT WAS USED TO COMPLETE THE PROCEDURE? WHAT IS THE PATIENT'S CONDITION? PROCEDURE DATE? WERE THERE ANY PATIENT CONSEQUENCES DUE TO THE "BREAKAGE FROM SWAGE POINT" INTRA-OP ISSUE? DEVICE RETURN STATUS/FOLLOW UP? EVENTS WERE SUBMITTED VIA 2210968-2020-09339, 2210968-2020-09340, AND 2210968-2020-09344
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10902255&pc=GAM | 2210968-2020-09289 | 2020-11-03 | Malfunction | ETHICON INC. | 2020-11-25 | GAM | VICRYL RAP UND 90CM M3 2/0 V-34 | Material Separation | No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EPISIOTOMY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. THE NEEDLE DETACHED AT SWAGE WHILE SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? NO. WAS PROCEDURE SUCCESSFULLY COMPLETED? YES. WERE FRAGMENTS GENERATED? NO. PATIENT STATUS/ OUTCOME / CONSEQUENCES. NO.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10898004&pc=GAM | 2210968-2020-09246 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-11-24 | GAM | VICRYL VIO 90CM M4 | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. ANOTHER LIKE DEVICE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER QCBCJQR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10898156&pc=GAM | 2210968-2020-09245 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-11-24 | GAM | VICRYL VIO 90CM M4 | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. ANOTHER LIKE DEVICE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER QCBCJQR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE, OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10898161&pc=GAM | 2210968-2020-09244 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-11-24 | GAM | VICRYL VIO 90CM M4 | Break | No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. ANOTHER LIKE DEVICE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER QCBCJQR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10875298&pc=GAM | 2210968-2020-09159 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-11-20 | GAM | VICRYL RAP CT UD 70CM M2 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE INTRA-OP DURING THE PROCEDURE. THE SURGERY WAS COMPLETED WITH A NEW SUTURE. THERE WAS NO HARM TO PATIENT AND NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V2150H; MK8BTSS0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY DEVICES BROKE? CUSTOMER IS UNABLE TO PROVIDE EXACT NUMBER OF INVOLVED SUTURES. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED VIA MW 2210968-2020-09158, 2210968-2020-09159. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10875308&pc=GAM | 2210968-2020-09158 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-11-20 | GAM | VICRYL RAP CT UD 70CM M2 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE INTRA-OP DURING THE PROCEDURE. THE SURGERY WAS COMPLETED WITH A NEW SUTURE. THERE WAS NO HARM TO PATIENT AND NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V2150H; MK8BTSS0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY DEVICES BROKE? CUSTOMER IS UNABLE TO PROVIDE EXACT NUMBER OF INVOLVED SUTURES. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED VIA MW 2210968-2020-09158, 2210968-2020-09159. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10848073&pc=GAM | 2210968-2020-09064 | 2020-10-22 | Malfunction | ETHICON INC. | 2020-11-16 | GAM | VICRYL*VLT PLUS 5-0 70CM 1-RB1 | Break; Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TUBAL RECANALIZATION PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE THREAD BROKE. ADDITIONALLY, THE NEEDLE FELL INTO THE CAVITY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? DID THE SUTURE BREAK AND DID THE NEEDLE PULLED OFF FROM THE THREAD DURING THIS PATIENT DURING THE PROCEDURE ON ONE OF THREE OCCASIONS CORRECT? WAS THE NEEDLE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? WHAT TISSUE/LOCATION WAS SUTURING WHEN PULL OFF OCCURRED? WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THIS SUTURE NEEDLE PULL OFF EVENT? WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME, 20-25 MINUTES.? IF YES, PLEASE EXPLAIN. TO DATE IT HAS BEEN REPORTED THAT THE DEVICE WILL NOT BE RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10848113&pc=GAM | 2210968-2020-09068 | 2020-10-22 | Malfunction | ETHICON INC. | 2020-11-16 | GAM | VICRYL*VLT PLUS 5-0 70CM 1-RB1 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TUBAL RECANALIZATION PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE THREAD BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? DID THE SUTURE BREAK AND DID THE NEEDLE PULLED OFF FROM THE THREAD DURING THIS PATIENT DURING THE PROCEDURE ON ONE OF THREE OCCASIONS CORRECT? WAS THE NEEDLE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? WHAT TISSUE/LOCATION WAS SUTURING WHEN PULL OFF OCCURRED? WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THIS SUTURE NEEDLE PULL OFF EVENT? WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME, 20-25 MINUTES? IF YES, PLEASE EXPLAIN. TO DATE IT HAS BEEN REPORTED THAT THE DEVICE WILL NOT BE RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10848118&pc=GAM | 2210968-2020-09067 | 2020-10-22 | Malfunction | ETHICON INC. | 2020-11-16 | GAM | VICRYL*VLT PLUS 5-0 70CM 1-RB1 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A TUBAL RECANALIZATION PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE THREAD BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? DID THE SUTURE BREAK AND DID THE NEEDLE PULLED OFF FROM THE THREAD DURING THIS PATIENT DURING THE PROCEDURE ON ONE OF THREE OCCASIONS CORRECT? WAS THE NEEDLE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? WHAT TISSUE/LOCATION WAS SUTURING WHEN PULL OFF OCCURRED? WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THIS SUTURE NEEDLE PULL OFF EVENT? WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME, 20-25 MINUTES.? IF YES, PLEASE EXPLAIN. TO DATE IT HAS BEEN REPORTED THAT THE DEVICE WILL NOT BE RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10839830&pc=GAM | 2210968-2020-09040 | 2020-10-22 | Malfunction | ETHICON INC. | 2020-11-13 | GAM | VICRYL CT VIO 90CM M3.5 | Break | No Consequences Or Impact To Patient; Foreign Body In Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A C-SECTION PROCEDURE ON 10/22/2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE TIP OF THE NEEDLE IS BROKEN OFF DURING SURGERY. WHILE PERFORMING THE CORNER SUTURES OF THE UTERUS IT WAS NOTICED THAT THE TIP OF THE NEEDLE WAS MISSING, APPROX. 0,5MM. THE MISSING PIECE COULD NOT BE RECOVERED AND IS STILL IN THE PATIENT¿S BODY. SURGERY WAS COMPLETED SUCCESSFULLY WITHOUT DELAY. NO X-RAYS WERE PERFORMED. ACCORDING TO SURGEON PIECE IS TOO SMALL TO BE SHOWN ON X-RAY. NO ADVERSE CONSEQUENCES FOR THE PATIENT WERE REPORTED. NO FURTHER PROCEDURE IS PLANNED. Manufacturer Narrative: (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10832953&pc=GAM | 2210968-2020-09007 | 2018-03-03 | Malfunction | ETHICON INC. | 2020-11-12 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Pain | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A KNEE TUMOR REMOVAL SURGERY ON (B)(6) 2018 AND THE SUTURE WAS USED. THE PATIENT EXPERIENCED PAIN SYNDROME TWO AND HALF YEARS AFTER SURGERY. IT WAS ALSO REPORTED THAT THE RESTS OF SUTURE REMAIN AFTER TWO AND HALF YEARS IN THE KNEE ON SUBCUTIS AND FASCIA. IT WAS REPORTED THAT THE SUTURE REMAINS ARE VISIBLE ON RECENT MRI SCAN AND POSSIBLY RELATED TO CRPS PAIN SYNDROME. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE? ON WHAT TISSUE WAS THE VICRYL SUTURE USED? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? PRODUCT CODE AND LOT NUMBER? PLEASE PROVIDE DETAILS OF MRI FINDINGS? HOW WAS IT CONFIRMED THAT SUTURE REMAINS BELONG TO THE VICRYL SUTURE? ONSET DATE/TIME OF THE PAIN FROM INDEX SURGERY? LOCATION AND CHARACTER OF THE PAIN? WHAT MEDICAL INTERVENTION WAS GIVEN FOR THE PAIN MANAGEMENT? RESULTS? OTHER RELEVANT PATIENT FACTORS/COMORBIDITIES/CONCURRENT PROCEDURES/CONCOMITANT MEDICATIONS? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? WHAT IS THE PATIENT¿S CURRENT STATUS? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10822369&pc=GAM | 2210968-2020-08949 | 2020-09-11 | Malfunction | ETHICON INC. | 2020-11-11 | GAM | VICRYL VIO 90CM M5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE DEVICE WAS DEFECTIVE. THE THREAD PULLED OFF FROM THE NEEDLE AT THE LEVEL THE EYE OF A NEEDLE, DURING THE SUTURE GRIP WITH THE NEEDLE HOLDER. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER QGBDXGR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT THE SUTURE FAILED (DEFECTIVE DEVICE. PLEASE EXPLAIN WHAT FAILED ON THE SUTURE. DID THE SUTURE BROKE/SUTURE SEPARATED FROM THE NEEDLE/NEEDLE BREAKAGE? PLEASE EXPLAIN IN MORE DETAIL. WHAT WAS BEING DONE WHEN THE SUTURE FAILED (REMOVAL FROM PACKAGE / DURING PASSAGE THROUGH TISSUE/ DURING TYING / POST-OP)? THE THREAD PULLED OFF FROM THE NEEDLE AT THE LEVEL THE EYE OF A NEEDLE, DURING THE SUTURE GRIP WITH THE NEEDLE HOLDER. PROCEDURE NAME AND DATE? (B)(6) 2020. EVENT DATE? (B)(6) 2020. WHAT WAS USED TO COMPLETE THE PROCEDURE? USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE. PATIENT'S CONDITION? NO PATIENT CONSEQUENCE DEVICE RETURN STATUS/FOLLOW UP? DEVICE NOT AVAILABLE FOR ANALYSIS. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10809005&pc=GAM | 2210968-2020-08830 |  | Malfunction | ETHICON INC. | 2020-11-09 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN AESTHETIC PROCEDURE ON AN UNKNOWN DATE; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF THE SUTURE. THE PROCEDURE WAS COMPLETED WITH A 15 MINUTE DELAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10802048&pc=GAM | 2210968-2020-08788 | 2020-01-09 | Malfunction | ETHICON INC. | 2020-11-06 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10791512&pc=GAM | 2210968-2020-08708 | 2020-10-26 | Malfunction | ETHICON INC. | 2020-11-05 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LESION EXCISION PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE DURING SKIN CLOSURE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY WITH A MINIMAL BUT UNSPECIFIED DELAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10791515&pc=GAM | 2210968-2020-08709 | 2020-10-26 | Malfunction | ETHICON INC. | 2020-11-05 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LESION EXCISION PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE DURING SKIN CLOSURE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY WITH A MINIMAL BUT UNSPECIFIED DELAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10782248&pc=GAM | 2210968-2020-08632 |  | Malfunction | ETHICON INC. | 2020-11-03 | GAM | VICRYL* UND 3-0 70CM (1)SC-20 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED. PLEASE CLARIFY WHAT IS THE SPECIFIC QUALITY ISSUE WITH THE SUTURE? THE SUTURE SEPARATES FROM THE NEEDLE OR SUTURE BROKE WERE FRAGMENTS OF THE DEVICE LEFT INSIDE THE PATIENT? WAS THERE ANY PATIENT CONSEQUENCE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10771291&pc=GAM | 2210968-2020-08547 | 2020-10-14 | Malfunction | ETHICON INC. | 2020-11-02 | GAM | VICRYL*UND 5-0 45CM (1)P-3 PR | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE DISASSEMBLED FROM THE SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW WAS THE PROCEDURE COMPLETED? PLEASE PROVIDE LOT NUMBER. PLEASE PROVIDE PROCEDURE NAME AND DATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10771292&pc=GAM | 2210968-2020-08548 | 2020-10-14 | Malfunction | ETHICON INC. | 2020-11-02 | GAM | VICRYL*UND 5-0 45CM (1)P-3 PR | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE DISASSEMBLED FROM THE SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW WAS THE PROCEDURE COMPLETED? PLEASE PROVIDE LOT NUMBER PLEASE PROVIDE PROCEDURE NAME AND DATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10771310&pc=GAM | 2210968-2020-08549 | 2020-10-14 | Malfunction | ETHICON INC. | 2020-11-02 | GAM | VICRYL*UND 5-0 45CM (1)P-3 PR | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE DISASSEMBLED FROM THE SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW WAS THE PROCEDURE COMPLETED? PLEASE PROVIDE LOT NUMBER PLEASE PROVIDE PROCEDURE NAME AND DATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10762244&pc=GAM | 2210968-2020-08487 | 2020-10-09 | Malfunction | ETHICON INC. | 2020-10-30 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN EXCISION OF A CUTANEOUS MASS PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF THE SUTURE AFTER A FEW TISSUE PASSAGES WITHOUT EXERTING EXCESSIVE TENSION WHILE SUTURING THE SUBCUTANEOUS TISSUES. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10741834&pc=GAM | 2210968-2020-08363 |  | Malfunction | ETHICON INC. | 2020-10-27 | GAM | VICRYL CT VIO 90CM M4 | Break; Material Separation | Not Applicable; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CESAREAN SECTION PROCEDURE ON AN UNKNOWN DATE; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE AND PULLED OFF THE SUTURE. THE BROKEN NEEDLE WAS RETRIEVED THROUGH X-RAY. THERE WAS A 30 MINUTE DELAY IN THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? DOES A PIECE OF THE NEEDLE REMAIN IN THE PATIENT¿S TISSUE? WHAT TISSUE STRUCTURE IS IT LOCATED? SIZE OF THE NEEDLE PIECE RETAINED? ARE ANY PLANS IN PLACE TO REMOVE THE NEEDLE PIECE IN FUTURE? WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? WAS MORE THAN HALF THE NEEDLE RETAINED IN THE PATIENT? A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PDBDSCR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10736577&pc=GAM | 2210968-2020-08336 |  | Malfunction | ETHICON INC. | 2020-10-26 | GAM | VICRYL PLUS POLYGLACTIN 910 SUTURE UNKNOWN PROD | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. FOLLOWING THE PROCEDURE, THE CUSTOMER OPINED OF WHY THE SUTURE IS DISSOLVING MUCH SLOWER THAN THE STANDARD VERSION. IT'S TAKING 8-10 WEEKS INSTEAD OF SIX WEEKS. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: WAS THERE ANY PATIENT CONSEQUENCES? PROCEDURE, EVENT DAY, LOT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10726684&pc=GAM | 2210968-2020-08276 |  | Malfunction | ETHICON INC. | 2020-10-23 | GAM | VICRYL* UND 3-0 45CM (1)E-24 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2020, AND THE SUTURE WAS USED. DURING SURGERY, THE THREAD IS BREAKING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). IF THE DEVICE, OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE, NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10726685&pc=GAM | 2210968-2020-08275 |  | Malfunction | ETHICON INC. | 2020-10-23 | GAM | VICRYL* UND 3-0 45CM (1)E-24 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2020 AND THE SUTURE WAS USED. DURING SURGERY, THE THREAD IS BREAKING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? PATIENT'S CONDITION? EVENT AND PROCEDURE DATE? PROCEDURE NAME? HOW MANY SUTURES BROKE IN THIS SINGLE PROCEDURE DURING USE? PLEASE PROVIDE JUSTIFICATION OF ALERT DATE BEING 9/30/2020. DEVICE RETURN STATUS/FOLLOW UP?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10726688&pc=GAM | 2210968-2020-08277 |  | Malfunction | ETHICON INC. | 2020-10-23 | GAM | VICRYL* UND 3-0 45CM (1)E-24 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2020 AND THE SUTURE WAS USED. DURING SURGERY, THE THREAD IS BREAKING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? PATIENT'S CONDITION? EVENT AND PROCEDURE DATE? PROCEDURE NAME? HOW MANY SUTURES BROKE IN THIS SINGLE PROCEDURE DURING USE? PLEASE PROVIDE JUSTIFICATION OF ALERT DATE BEING 9/30/2020. DEVICE RETURN STATUS/FOLLOW UP?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10726689&pc=GAM | 2210968-2020-08278 |  | Malfunction | ETHICON INC. | 2020-10-23 | GAM | VICRYL* UND 3-0 45CM (1)E-24 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2020 AND THE SUTURE WAS USED. DURING SURGERY, THE THREAD IS BREAKING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? PATIENT'S CONDITION? EVENT AND PROCEDURE DATE? PROCEDURE NAME? HOW MANY SUTURES BROKE IN THIS SINGLE PROCEDURE DURING USE? PLEASE PROVIDE JUSTIFICATION OF ALERT DATE BEING 9/30/2020. DEVICE RETURN STATUS/FOLLOW UP?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10726693&pc=GAM | 2210968-2020-08279 |  | Malfunction | ETHICON INC. | 2020-10-23 | GAM | VICRYL* UND 3-0 45CM (1)E-24 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN 2020 AND THE SUTURE WAS USED. DURING SURGERY, THE THREAD IS BREAKING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? PATIENT'S CONDITION? EVENT AND PROCEDURE DATE? PROCEDURE NAME? HOW MANY SUTURES BROKE IN THIS SINGLE PROCEDURE DURING USE? PLEASE PROVIDE JUSTIFICATION OF ALERT DATE BEING 9/30/2020. DEVICE RETURN STATUS/FOLLOW UP?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10729881&pc=GAM | 2210968-2020-08296 | 2020-09-23 | Malfunction | ETHICON INC. | 2020-10-23 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10730408&pc=GAM | 2210968-2020-08299 | 2020-09-08 | Malfunction | ETHICON INC. | 2020-10-23 | GAM | VICRYL CTD VIO 1 90CM CT PLUS | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED: "THE NEEDLE OF A FIRST WIRE VICRYL PURPLE PULLED OFF AND BROKE ON THE POINT OF CRIMPING" PLEASE CLARIFY: THE NEEDLE BROKE AND ALSO THE NEEDLE AND SUTURE PULLED OFF DURING THE PROCEDURE? ACCORDING TO THE DESCRIPTION THE THREAD GOT BROKEN AT THE LEVEL OF ATTACHMENT END (THE NEEDLE DID NOT BROKE BUT PULLED FROM THE THREAD). IT WAS REPORTED ALSO "AND THEN A NEEDLE OF A SECOND WIRE (SAME LOT) BROKE INTO 2 PARTS" PLEASE CLARIFY: ANOTHER ONE NEEDLE BROKE DURING THE INTRA OP EVENT OF SAME LOT MD8DSXP0 AND PRODUCT CODE V359H? PLEASE CLARIFY IF PRODUCT CODE IS ALSO THE SAME. A SECOND DEVICE FROM THE SAME LOT HAS BEEN USED AND THE NEEDLE GOT BROKEN IN TWO PART. THE BREAK OCCURRED AT THE LEVEL OF GRASP AREA (GRASP THE NEEDLE IN AN AREA ONE THIRD (1/3) TO ONE HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT.) WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? NO PATIENT CONSEQUENCES. THE TWO PART OF THE NEEDLE HAS BEEN RECOVERED AND NO METALLIC PART HAVE BEEN OBSERVED BY RADIOGRAPHY. WHAT IS CURRENT CONDITION OF THE PATIENT? NO PATIENT CONSEQUENCES. PROCEDURE DATE? (B)(6) 2020. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10730475&pc=GAM | 2210968-2020-08298 | 2020-09-08 | Malfunction | ETHICON INC. | 2020-10-23 | GAM | VICRYL CTD VIO 1 90CM CT PLUS | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. AN X-RAYS CONTROL WAS DONE ON THE OPERATING TABLE. NO METAL RESIDUES WERE IDENTIFIED ON THE X RAYS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED: "THE NEEDLE OF A FIRST WIRE VICRYL PURPLE PULLED OFF AND BROKE ON THE POINT OF CRIMPING" PLEASE CLARIFY: THE NEEDLE BROKE AND ALSO THE NEEDLE AND SUTURE PULLED OFF DURING THE PROCEDURE? ACCORDING TO THE DESCRIPTION THE THREAD GOT BROKEN AT THE LEVEL OF ATTACHMENT END (THE NEEDLE DID NOT BROKE BUT PULLED FROM THE THREAD). IT WAS REPORTED ALSO "AND THEN A NEEDLE OF A SECOND WIRE (SAME LOT) BROKE INTO 2 PARTS" PLEASE CLARIFY: ANOTHER ONE NEEDLE BROKE DURING THE INTRA OP EVENT OF SAME LOT MD8DSXP0 AND PRODUCT CODE V359H? PLEASE CLARIFY IF PRODUCT CODE IS ALSO THE SAME. A SECOND DEVICE FROM THE SAME LOT HAS BEEN USED AND THE NEEDLE GOT BROKEN IN TWO PART. THE BREAK OCCURRED AT THE LEVEL OF GRASP AREA (GRASP THE NEEDLE IN AN AREA ONE THIRD (1/3) TO ONE HALF (1/2) OF THE DISTANCE FROM THE ATTACHMENT END TO THE POINT.) WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? NO PATIENT CONSEQUENCES. THE TWO PART OF THE NEEDLE HAS BEEN RECOVERED AND NO METALLIC PART HAVE BEEN OBSERVED BY RADIOGRAPHY. WHAT IS CURRENT CONDITION OF THE PATIENT? NO PATIENT CONSEQUENCES. PROCEDURE DATE? (B)(6) 2020. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10719282&pc=GAM | 2210968-2020-08205 | 2020-10-05 | Malfunction | ETHICON INC. | 2020-10-22 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10704780&pc=GAM | 2210968-2020-08129 |  | Malfunction | ETHICON INC. | 2020-10-20 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient; Not Applicable; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE. THERE WAS A UNSPECIFIED DELAY IN THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE PROVIDE PROCEDURE NAME AND DATE. HOW LONG WAS THE DELAY? WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN. PATIENT¿S CURRENT STATUS? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10705965&pc=GAM | 2210968-2020-08144 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-10-20 | GAM | VICRYL CT VIO 75CM M2 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND SUTURE WAS TO BE USED. THE CUSTOMER NOTICED THAT THE LABEL WAS MISSING ON THE INDIVIDUAL PACKAGING. NO PROCEDURE IMPACTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. COULD YOU PLEASE SPECIFY HOW MANY LABELS WERE MISSING OF EACH PRODUCT CODE WERE OBSERVED: PRODUCT CODE C395 QUANTITY OF PRODUCT INVOLVED IS 36. PRODUCT CODE C394 QUANTITY OF PRODUCT INVOLVED IS 180. PRODUCT CODE C389 QUANTITY OF PRODUCT INVOLVED IS 72. PRODUCT CODE JV398 QUANTITY OF PRODUCT INVOLVED IS 72. DID THESE ISSUES PRESENT DURING A PROCEDURE? IF YES. WHAT WAS THE INITIAL PROCEDURE NAME AND DATE? WHEN DID THE ISSUES OCCURRED (PRE-OP, INTRA-OP OR POST-OP)? WHAT WAS THE EVENT DAY DATE? PLEASE PROVIDE LOT NUMBERS FOR EACH PRODUCT CODES: PRODUCT CODE C395 LOT NUMBER? PRODUCT CODE C394 LOT NUMBER? PRODUCT CODE C389 LOT NUMBER? PRODUCT CODE JV398 LOT NUMBER?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10689622&pc=GAM | 2210968-2020-08070 | 2020-09-23 | Malfunction | ETHICON INC. | 2020-10-16 | GAM | VICRYL RAP UND 75CM M2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020, AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10677273&pc=GAM | 2210968-2020-07979 | 2020-09-23 | Malfunction | ETHICON INC. | 2020-10-14 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DELIVERY SURGERY ON (B)(6) 2020, AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. ANOTHER DEVICE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4) TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10681219&pc=GAM | 2210968-2020-08028 | 2020-09-23 | Malfunction | ETHICON INC. | 2020-10-14 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT SUTURE THREADS WERE TORN APART. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS REPORTED "THE SUTURE THREADS ARE TORN APART, 3 SACHETS IAIN 1 BOX" PLEASE CLARIFY: PLEASE CONFIRM THE QUANTITY OF BROKEN SUTURES DURING THIS SINGLE PROCEDURE ON ONE PATIENT? WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? PATIENT'S CONDITION? PROCEDURE NAME AND DATE? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA 2210968-2020-08027 AND 2210968-2020-08029.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10681221&pc=GAM | 2210968-2020-08029 | 2020-09-23 | Malfunction | ETHICON INC. | 2020-10-14 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT SUTURE THREADS WERE TORN APART. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS REPORTED "THE SUTURE THREADS ARE TORN APART, 3 SACHETS IAIN 1 BOX" PLEASE CLARIFY: PLEASE CONFIRM THE QUANTITY OF BROKEN SUTURES DURING THIS SINGLE PROCEDURE ON ONE PATIENT? WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? PATIENT'S CONDITION? PROCEDURE NAME AND DATE? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA 2210968-2020-08027 AND 2210968-2020-08028.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10681230&pc=GAM | 2210968-2020-08027 | 2020-09-23 | Malfunction | ETHICON INC. | 2020-10-14 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT SUTURE THREADS WERE TORN APART. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS REPORTED "THE SUTURE THREADS ARE TORN APART, 3 SACHETS IAIN 1 BOX" PLEASE CLARIFY: PLEASE CONFIRM THE QUANTITY OF BROKEN SUTURES DURING THIS SINGLE PROCEDURE ON ONE PATIENT? WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? WHAT WAS USED TO COMPLETE THE PROCEDURE? PATIENT'S CONDITION? PROCEDURE NAME AND DATE? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA 2210968-2020-08028 AND 2210968-2020-08029.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10671905&pc=GAM | 2210968-2020-07931 | 2020-09-18 | Malfunction | ETHICON INC. | 2020-10-13 | GAM | VICRYL CT VIO 45CM M2 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DENTAL PROCEDURE ON (B)(6) 2020, AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD LOOSED FROM THE NEEDLE WHILE SUTURING, AND THE NEEDLE FEEL DOWN ON THE PATIENT'S CLOTHES. THE NEEDLE WAS REMOVED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER PABGLDQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PHOTO UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PHOTO EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WAS THE NEEDLE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? THE NEEDLE WAS REMOVED, IT FELL DOWN ON THE PATIENTS CLOTHES. PROCEDURE DATE? (B)(6) 2020. HOW WAS CASE COMPLETED? GOT ASSISTANCE FROM THE QUALITY TEAM. ANY ADVERSE PATIENT CONSEQUENCES, AND WHAT MEDICAL/SURGICAL INTERVENTION WAS PERFORMED TO MANAGE THEM? NO CONSEQUENCES. ARE SAMPLES BEING RETURNED FOR EVALUATION? THE BROKEN SUTURE HAS BEEN THROWN AWAY, THEY WILL SENT SUTURE FROM THE SAME BOX.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10662846&pc=GAM | 2210968-2020-07871 | 2020-09-20 | Malfunction | ETHICON INC. | 2020-10-12 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CONFIRM LOT NUMBER. LOT QDBBWKRO APPEARED INVALID. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY DEVICES PRESENT THE ISSUE (BREAKAGE SUTURE) IN THIS PROCEDURE? 1 PIECE OF SUTURE BROKE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10662852&pc=GAM | 2210968-2020-07870 | 2020-09-20 | Malfunction | ETHICON INC. | 2020-10-12 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CONFIRM LOT NUMBER. LOT QDBBWKRO APPEARED INVALID. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW MANY DEVICES PRESENT THE ISSUE (BREAKAGE SUTURE) IN THIS PROCEDURE? 1 PIECE OF SUTURE BROKE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10653807&pc=GAM | 2210968-2020-07436 | 2020-09-17 | Malfunction | ETHICON INC. | 2020-10-08 | GAM | VICRYL RAP UND 45CM M1.5 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020, AND SUTURE WAS USED. THE THREAD BROKE WITH THE SLIGHTEST PRESSURE NEAR THE NEEDLE OR SOMEWHERE IN THE MIDDLE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE, NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. COULD YOU PLEASE CONFIRM IF IT OCCURRED DURING A PROCEDURE? IF YES WHAT WAS THE INITIAL PROCEDURE? DID THIS ISSUE OCCURRED ONCE OR IT HAPPEN IN OTHER PROCEDURES? IN CASE THIS OCCURRED IN SEVERAL PROCEDURES, COULD YOU PLEASE CONFIRM IF THESE WAS REPORTED IN OTHER PC? PLEASE PROVIDE THE OTHER PC'S NUMBERS. DID THIS ISSUE OCCURRED PRE-OP, INTRA-OP OR POST-OP? PLEASE CONFIRM IF IN THIS PARTICULAR CASE DID THE SUTURE BROKE OR IF IT PRESENT PULL OFF SUTURE NEEDLE? DID THE PATIENT SUFFER FROM ANY CONSEQUENCE DUE TO THIS ISSUE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10645247&pc=GAM | 2210968-2020-07756 | 2020-08-28 | Malfunction | ETHICON INC. | 2020-10-07 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED V-37 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHERE THE NEEDLES DETACHED IN THE PACKAGE OR AT THE MOMENT OF SUTURING? THE SUTURE WAS DETACHED AT THE MOMENT OF SUTURING HOW MANY DEVICES WERE INVOLVED IN THIS SINGLE PROCEDURE? THE SUTURE WAS DETACHED AT THE MOMENT OF SUTURING SO THEY HAVE TO USE 3 EA IN THAT CASE. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10645339&pc=GAM | 2210968-2020-07757 | 2020-08-28 | Malfunction | ETHICON INC. | 2020-10-07 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED V-37 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHERE THE NEEDLES DETACHED IN THE PACKAGE OR AT THE MOMENT OF SUTURING? THE SUTURE WAS DETACHED AT THE MOMENT OF SUTURING HOW MANY DEVICES WERE INVOLVED IN THIS SINGLE PROCEDURE? THE SUTURE WAS DETACHED AT THE MOMENT OF SUTURING SO THEY HAVE TO USE 3 EA IN THAT CASE. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10645359&pc=GAM | 2210968-2020-07758 | 2020-08-28 | Malfunction | ETHICON INC. | 2020-10-07 | GAM | VICRYL VIO 90CM M4 USP1 SGLE ARMED V-37 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHERE THE NEEDLES DETACHED IN THE PACKAGE OR AT THE MOMENT OF SUTURING? THE SUTURE WAS DETACHED AT THE MOMENT OF SUTURING HOW MANY DEVICES WERE INVOLVED IN THIS SINGLE PROCEDURE? THE SUTURE WAS DETACHED AT THE MOMENT OF SUTURING SO THEY HAVE TO USE 3 EA IN THAT CASE. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10636199&pc=GAM | 2210968-2020-07678 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-10-06 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE IN 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS REPORTED THAT THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT DO YOU MEAN BY "SUTURE BROKE AT ATTACHMENT AREA"? SUTURE SEPARATED FROM NEEDLE. PLEASE CONFIRM HOW MANY DEVICES PRESENT THE ISSUE. NUMBER OF AFFECTED SUTURES IS UNKNOWN. WHAT IS THE EVENT DAY AND PROCEDURE DATE? UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10636200&pc=GAM | 2210968-2020-07677 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-10-06 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE IN 2020, AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS REPORTED THAT THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT DO YOU MEAN BY "SUTURE BROKE AT ATTACHMENT AREA"? SUTURE SEPARATED FROM NEEDLE. PLEASE CONFIRM HOW MANY DEVICES PRESENT THE ISSUE. NUMBER OF AFFECTED SUTURES IS UNKNOWN. WHAT IS THE EVENT DAY AND PROCEDURE DATE? UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10636207&pc=GAM | 2210968-2020-07684 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-10-06 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE IN 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, IT WAS REPORTED THAT THE SUTURE GOT DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT DO YOU MEAN BY "SUTURE BROKE AT ATTACHMENT AREA"? SUTURE SEPARATED FROM NEEDLE. PLEASE CONFIRM HOW MANY DEVICES PRESENT THE ISSUE. NUMBER OF AFFECTED SUTURES IS UNKNOWN. WHAT IS THE EVENT DAY AND PROCEDURE DATE? UNKNOWN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10597078&pc=GAM | 2210968-2020-07413 | 2020-09-14 | Malfunction | ETHICON INC. | 2020-09-28 | GAM | VICRYL VIO 90CM M5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS DETACHED FROM THE SUTURE DURING THE CLOSURE OF THE UTERUS. THE DEVICE WAS CHANGE, RISK OF HEMORRHAGE NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. COULD YOU PLEASE CONFIRM IF THE PATIENT SUFFER FROM ANY HEMORRHAGE? IF YES, PLEASE ADVICE HOW MANY BLOOD DID THE PATIENT LOST?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10580759&pc=GAM | 2210968-2020-07280 | 2020-08-03 | Malfunction | ETHICON INC. | 2020-09-24 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Fracture | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT WHEN PREPARING FOR PERINEAL LATERAL INCISION REPAIR SURGERY ON (B)(6) 2020, THE SUTURE BROKE EASILY. THE PROCEDURE WAS COMPLETED USING A LIKE DEVICE, AND THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B(4). ANALYSIS SUMMARY: IT WAS RECEIVED FOR ANALYSIS AN OPENED SAMPLE OF PRODUCT CODE W9923, LOT AF1459. DURING THE VISUAL INSPECTION OF SAMPLE, THE SWAGE AND ATTACHMENT AREA OF THE NEEDLE WERE NOTED TO BE AS EXPECTED. THE SUTURE COULD NOT BE DISPENSED SINCE HAS BEGUN WITH THE DEGRADATION PROCESS AND THIS IS THE CAUSE THAT BREAKAGE SUTURE. THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED AND THE FUNCTIONAL TEST CANNOT BE PERFORMED. ALSO, THE FOIL WAS EXAMINED AND SHOWED MULTIPLES WRINKLES AND HOLES ON FOIL PACKAGE DUE TO EXCESSIVE MANIPULATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10580914&pc=GAM | 2210968-2020-07282 | 2020-08-03 | Malfunction | ETHICON INC. | 2020-09-24 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Fracture | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT WHEN PREPARING FOR PERINEAL LATERAL INCISION REPAIR SURGERY ON (B)(6) 2020, THE SUTURE BROKE EASILY. THE PROCEDURE WAS COMPLETED USING A LIKE DEVICE AND THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ANALYSIS SUMMARY: IT WAS RECEIVED FOR ANALYSIS AN OPENED SAMPLE OF PRODUCT CODE W9923, LOT AF1284. DURING THE VISUAL INSPECTION OF SAMPLE, THE SWAGE AND ATTACHMENT AREA OF THE NEEDLE WERE NOTED TO BE AS EXPECTED. THE SUTURE COULD NOT BE DISPENSED SINCE HAS BEGUN WITH THE DEGRADATION PROCESS AND THIS IS THE CAUSE THAT BREAKAGE SUTURE. THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED AND THE FUNCTIONAL TEST CANNOT BE PERFORMED. ALSO, THE FOIL WAS EXAMINED AND SHOWED MULTIPLES WRINKLES AND HOLES ON FOIL PACKAGE DUE TO EXCESSIVE MANIPULATION. IN ADDITION, A REVIEW OF THE BATCH MANUFACTURING RECORDS WAS CONDUCTED AND THE BATCH MET ALL FINISHED GOODS RELEASE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10570232&pc=GAM | 2210968-2020-07218 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-09-22 | GAM | VICRYL VIO 75CM M4 1 CT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE INITIAL PROCEDURE? DID THE SURGERY COMPLETED SUCCESSFULLY? WHAT WAS USED TO COMPLETE THE PROCEDURE? DID THE PATIENT SUFFER FROM ANY CONSEQUENCE DUE TO THE BREAKAGE NEEDLE? WHAT IS THE EVENT DAY AND PROCEDURE DATE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10549369&pc=GAM | 2210968-2020-07057 | 2020-07-15 | Malfunction | ETHICON INC. | 2020-09-17 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN OOPHORECTOMY ON (B)(6) 2020 AND A SUTURE WAS USED. DURING THE PROCEDURE, AFTER FEW TISSUE STITCHES IN THE ABDOMINAL WALL, THE NEEDLE PULLED OFF THE THREAD WITHOUT ANY EXCESSIVE FORCE APPLICATION. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV452; PMBDKGR0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE NEEDLE PULLED OFF THE THREAD. A USED NEEDLE OF PRODUCT CODE JV587, LOT PMBDKGR0 WERE RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED; HOWEVER MARKS APPEAR TO BE BY USE OF SURGICAL INSTRUMENT COULD BE OBSERVE. THE BARREL HOLE WAS EXAMINED UNDER 20X MAGNIFICATION FOR SUTURE REMNANT AND NONE WAS NOTED. IN ADDITION, THE SUTURE WAS NOT RETURNED. AS PART OF OUR QUALITY PROCESS, THE MANUFACTURING RECORDS OF THIS LOT-SERIAL NUMBER WERE REVIEWED, AND THE MANUFACTURING STANDARDS WERE MET PRIOR TO THE RELEASE OF THIS BATCH. IT WAS RECEIVED FOR ANALYSIS FOURTEEN UNOPENED SAMPLES OF PRODUCT CODE JV587, LOT PMBDKGR0. DURING THE VISUAL INSPECTION OF THIRTEEN SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10549420&pc=GAM | 2210968-2020-07058 | 2020-07-15 | Malfunction | ETHICON INC. | 2020-09-17 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN OOPHORECTOMY ON (B)(6) 2020 AND A SUTURE WAS USED. DURING THE PROCEDURE, AFTER FEW TISSUE STITCHES IN THE ABDOMINAL WALL, THE NEEDLE PULLED OFF THE THREAD WITHOUT ANY EXCESSIVE FORCE APPLICATION. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV452; PMBDKGR0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE NEEDLE PULLED OFF THE THREAD. A USED NEEDLE OF PRODUCT CODE JV587, LOT PMBDKGR0 WERE RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION OF NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED; HOWEVER MARKS APPEAR TO BE BY USE OF SURGICAL INSTRUMENT COULD BE OBSERVE. THE BARREL HOLE WAS EXAMINED UNDER 20X MAGNIFICATION FOR SUTURE REMNANT AND NONE WAS NOTED. IN ADDITION, THE SUTURE WAS NOT RETURNED. AS PART OF OUR QUALITY PROCESS, THE MANUFACTURING RECORDS OF THIS LOT-SERIAL NUMBER WERE REVIEWED, AND THE MANUFACTURING STANDARDS WERE MET PRIOR TO THE RELEASE OF THIS BATCH. IT WAS RECEIVED FOR ANALYSIS FOURTEEN UNOPENED SAMPLES OF PRODUCT CODE JV587, LOT PMBDKGR0. DURING THE VISUAL INSPECTION OF THIRTEEN SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10539697&pc=GAM | 2210968-2020-06949 | 2020-07-15 | Malfunction | ETHICON INC. | 2020-09-16 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN OOPHORECTOMY ON (B)(6) 2020 AND A SUTURE WAS USED. DURING THE PROCEDURE, AFTER FEW TISSUE STITCHES IN THE ABDOMINAL WALL, THE NEEDLE PULLED OFF THE THREAD WITHOUT ANY EXCESSIVE FORCE APPLICATION. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER PMBDKGR0, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE NEEDLE PULLED OFF THE THREAD. ONE DETACHED NEEDLE AND A SUTURE PIECES OF PRODUCT CODE JV587, LOT PMBDKGR0. DURING THE VISUAL INSPECTION OF NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE WAS EXAMINED UNDER MAGNIFICATION FOR SUTURE REMNANT AND NONE WAS NOTED. ADDITIONALLY, THERE ARE SLIGHTLY MARKS ON BODY THAT APPEARS TO BE BY USE OF SURGICAL INSTRUMENT. THE SUTURE END WAS EXAMINED AND THERE ISN¿T SUFFICIENT IMPRESSION AT THE SWAGE END, RESULTING IN THE NEEDLE PULL OFF DUE TO LIGHT SWAGE. THIS DEFECT HAS BEEN CORRELATED TO THE MANUFACTURING PROCESS. AS PART OF OUR QUALITY PROCESS, THE MANUFACTURING RECORDS OF THIS LOT-SERIAL NUMBER WERE REVIEWED, AND THE MANUFACTURING STANDARDS WERE MET PRIOR TO THE RELEASE OF THIS BATCH. IT WAS RECEIVED FOR ANALYSIS FOURTEEN UNOPENED SAMPLES OF PRODUCT CODE JV587, LOT PMBDKGR0. DURING THE VISUAL INSPECTION OF SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10539706&pc=GAM | 2210968-2020-06950 | 2020-07-15 | Malfunction | ETHICON INC. | 2020-09-16 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN OOPHORECTOMY ON (B)(6) 2020 AND A SUTURE WAS USED. DURING THE PROCEDURE, AFTER FEW TISSUE STITCHES IN THE ABDOMINAL WALL, THE NEEDLE PULLED OFF THE THREAD WITHOUT ANY EXCESSIVE FORCE APPLICATION. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER PMBDKGR0, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED THAT THE NEEDLE PULLED OFF THE THREAD. ONE DETACHED NEEDLE AND A SUTURE PIECES OF PRODUCT CODE JV587, LOT PMBDKGR0. DURING THE VISUAL INSPECTION OF NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE WAS EXAMINED UNDER MAGNIFICATION FOR SUTURE REMNANT AND NONE WAS NOTED. ADDITIONALLY, THERE ARE SLIGHTLY MARKS ON BODY THAT APPEARS TO BE BY USE OF SURGICAL INSTRUMENT. THE SUTURE END WAS EXAMINED AND THERE ISN'T SUFFICIENT IMPRESSION AT THE SWAGE END, RESULTING IN THE NEEDLE PULL OFF DUE TO LIGHT SWAGE. THIS DEFECT HAS BEEN CORRELATED TO THE MANUFACTURING PROCESS. AS PART OF OUR QUALITY PROCESS, THE MANUFACTURING RECORDS OF THIS LOT-SERIAL NUMBER WERE REVIEWED, AND THE MANUFACTURING STANDARDS WERE MET PRIOR TO THE RELEASE OF THIS BATCH. IT WAS RECEIVED FOR ANALYSIS FOURTEEN UNOPENED SAMPLES OF PRODUCT CODE JV587, LOT PMBDKGR0. DURING THE VISUAL INSPECTION OF SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRANDS AND NO DEFECTS WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES, NO PERFORMANCE PULL OFF SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10527282&pc=GAM | 2210968-2020-06830 | 2020-08-14 | Malfunction | ETHICON INC. | 2020-09-14 | GAM | VICRYL RAP UND 90CM M3.5 | Material Separation | Pain; No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE TO REPAIR A TEAR ON THE PERINEUM AFTER CHILDBIRTH ON (B)(6) 2020, AND SUTURE WAS USED. DURING THE PROCEDURE, THE THREAD PULLED OFF FROM THE NEEDLE ON THE PATIENT¿S PERINEUM. THE NEEDLE WAS REMOVED WITH DIFFICULTY FROM THE PATIENT'S PERINEUM. THE PATIENT HAD PAIN AND ADAPTATION OF PAINKILLERS. THE SURGEON IMMEDIATELY CHANGED THE DEVICE, RE-PERFORMED THE SUTURE, AND COMPLETED THE PROCEDURE SUCCESSFULLY WITH NO ADVERSE PATIENT CONSEQUENCES. THE PATIENT FOLLOW-UP HAD BEEN PHYSIOLOGIC WITHOUT ANY ADDITIONAL PAIN KILLERS. THERE WAS NO COMPLICATION ASSOCIATED TO THE REPORTED EVENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS THE TISSUE CONDITION? WAS ANY SURGICAL INTERVENTION OR RE-SUTURING PERFORMED? ANY MEDICAL TREATMENT PROVIDED? IF YES, PLEASE PROVIDE MEDICINE NAME, STRENGTH AND DOSE. COULD YOU PLEASE CONFIRM DEVICE'S AVAILABILITY? DID TISSUE DAMAGE OCCUR? WAS THE NEEDLE RETRIEVED DURING THE INITIAL SURGERY OR WAS AN ADDITIONAL SURGICAL INTERVENTION NEEDED? THE THREAD WAS USED FOR THE SUTURE OF THE PERINEUM SO THE TEAR WAS CLEAN, MORE OR LESS REGULAR. THE NEEDLE WAS WITHDRAWN WITH DIFFICULTY WITH THE NEEDLE HOLDER BECAUSE IT WAS TAKE DEEP IN THE GLUTEAL MUSCLE. THERE WAS NO TREATMENT TO PUT IN PLACE. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHEN DID PAIN OCCUR INTRA-OP OR POST-OP? WHAT DO YOU MEAN BY ADAPTATION OF PAIN KILLERS? WERE ADDITIONAL PAIN KILLERS GIVEN TO THE PATIENT? IF SO, WHEN? WHAT IS MEANT BY ¿THE SUTURE HAD TO BE RE-PERFORMED¿? WAS THE PATIENT RE-SUTURED DURING THE INITIAL PROCEDURE? PLEASE EXPLAIN AND PROVIDE CLARIFICATION. THERE WAS NO COMPLICATION ASSOCIATED TO THE REPORTED EVENT. THE USER IMMEDIATELY CHANGED THE DEVICE, AND COMPLETED THE PROCEDURE SUCCESSFULLY. THE PATIENT FOLLOW-UP HAD BEEN PHYSIOLOGIC WITHOUT ANY ADDITIONAL PAIN KILLERS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10527829&pc=GAM | 2210968-2020-06835 | 2020-08-18 | Malfunction | ETHICON INC. | 2020-09-14 | GAM | VICRYL VIO 90CM M3 2/0 MH PLUS | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ABDOMINAL CLOSURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, WHILE THE SURGEON WAS CLOSING THE ABDOMEN, THE SUTURE BECAME DETACHED FROM THE NEEDLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE W9440; PMBBQMS0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE SUTURE DETACHED FROM THE NEEDLE DURING THE PROCEDURE? YES THAT IS CORRECT. PLEASE CONFIRM TWO SUTURES WERE DETACHED FROM THE NEEDLE? YES 2. HOW MANY SUTURES IN TOTAL WILL BE COMING BACK SINCE QUANTITY TO BE RETURNED IS 2 AND PER EVENT DESCRIPTION "3 BOXES"? - THE TWO SUTURES WILL BE RETURNED WITH NO NEEDLE. THE THREE BOXES ARE ALL OF THE SAME LOT NUMBER SO WILL BE COLLECTED AND SENT FOR TESTING. TOTAL OF 38 SUTURES. WAS THE NEEDLE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? YES. WAS THERE ANY PATIENT CONSEQUENCE OR INJURY? NO CONSEQUENCE OR INJURY, JUST NEED TO USE ANOTHER SUTURE. WHAT IS THE CONDITION OF THE PATIENT? THE PATIENT LEFT THEATRE / RECOVERY (PACU) IN A SATISFACTORY CONDITION. PROCEDURE DATE? (B)(6) 2020. LOT NUMBER? PMBBQMS0. NOTE: EVENTS REPORTED IN 2210968-2020-06836. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10527835&pc=GAM | 2210968-2020-06836 | 2020-08-18 | Malfunction | ETHICON INC. | 2020-09-14 | GAM | VICRYL VIO 90CM M3 2/0 MH PLUS | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ABDOMINAL CLOSURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, WHILE THE SURGEON WAS CLOSING THE ABDOMEN, THE SUTURE BECAME DETACHED FROM THE NEEDLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE W9440; PMBBQMS0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE SUTURE DETACHED FROM THE NEEDLE DURING THE PROCEDURE? YES THAT IS CORRECT. PLEASE CONFIRM TWO SUTURES WERE DETACHED FROM THE NEEDLE? YES 2. HOW MANY SUTURES IN TOTAL WILL BE COMING BACK SINCE QUANTITY TO BE RETURNED IS 2 AND PER EVENT DESCRIPTION "3 BOXES"? - THE TWO SUTURES WILL BE RETURNED WITH NO NEEDLE. THE THREE BOXES ARE ALL OF THE SAME LOT NUMBER SO WILL BE COLLECTED AND SENT FOR TESTING. TOTAL OF 38 SUTURES. WAS THE NEEDLE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? YES. WAS THERE ANY PATIENT CONSEQUENCE OR INJURY? NO CONSEQUENCE OR INJURY, JUST NEED TO USE ANOTHER SUTURE. WHAT IS THE CONDITION OF THE PATIENT? THE PATIENT LEFT THEATRE / RECOVERY (PACU) IN A SATISFACTORY CONDITION. PROCEDURE DATE? (B)(6) 2020. LOT NUMBER? PMBBQMS0. NOTE: EVENTS REPORTED IN 2210968-2020-06835. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10519020&pc=GAM | 2210968-2020-06778 | 2020-08-01 | Malfunction | ETHICON INC. | 2020-09-10 | GAM | VICRYL RAP UND 75CM M2 USP3/0 SGLE ARMED PSLX PRIME | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE FOLLOWING INFORMATION WAS REQUESTED BUT WAS UNAVAILABLE: IT WAS REPORTED THAT THE SUTURE BROKE DURING KNOTTING IN UNKNOWN SURGERY PLEASE CLARIFY: WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? PROCEDURE NAME AND DATE? EVENT DATE? DEVICE RETURN STATUS/FOLLOW UP? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10508913&pc=GAM | 2210968-2020-06735 | 2020-08-14 | Malfunction | ETHICON INC. | 2020-09-09 | GAM | VICRYL ABS SUR SUT USP SYN | Device Markings/Labelling Problem; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. IT WAS NOTED THAT THE PACKAGE IS ONE AND PRODUCT INSIDE IS ANOTHER. THE PACKET IS OF NW2317 AND THE INNER ZIPPER NUMBER IS VP2317. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY EVENT. REPORTED VP2617 INSIDE NW2317 FOLDER. PHOTOS SHOW VP2317 AND NW2317 FOLDER. PLEASE CONFIRM IF VP2317 WAS FOUND INSIDE NW2317. AS PER THE LATEST UPDATE FROM SALES REP, THE PRODUCT CODE IS NW2317 AND THE INNER ZIPPER NO. IS VP2317 ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IN EVENT DESCRIPTION INDICATES "THE PACKAGE OF SUTURE IS NOT IN A GOOD CONDITION." "PACKAGE IS ONE AND PRODUCT INSIDE IS ANOTHER." "PRODUCT INSIDE IS BROKEN." COULD YOU PLEASE CONFIRM IF THIS 3 ISSUES OCCURRED AT THE SAME TIME? HOW MANY DEVICES WERE INVOLVED? DID THIS ISSUES OCCURRED PRE-OP OR INTRA-OP? PLEASE EXPLAIN. IF THIS OCCURRED INTRA-OP, WHAT WAS THE PROCEDURE NAME? CONFIRM DEVICE'S AVAILABILITY. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10509024&pc=GAM | 2210968-2020-06737 | 2020-08-14 | Malfunction | ETHICON INC. | 2020-09-09 | GAM | VICRYL ABS SUR SUT USP SYN | Device Markings/Labelling Problem; Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC CHOLECYSTECTOMY ON (B)(6) 2020 AND SUTURE WAS USED. PRIOR TO USE, IT WAS NOTED THAT THE PACKAGE IS ONE AND PRODUCT INSIDE IS ANOTHER. THE PACKET IS OF NW2317 AND THE INNER ZIPPER NUMBER IS VP2317. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: PLEASE CLARIFY EVENT. REPORTED VP2617 INSIDE NW2317 FOLDER. PHOTOS SHOW VP2317 AND NW2317 FOLDER. PLEASE CONFIRM IF VP2317 WAS FOUND INSIDE NW2317. AS PER THE LATEST UPDATE FROM SALES REP, THE PRODUCT CODE IS NW2317 AND THE INNER ZIPPER NO. IS VP2317. IN EVENT DESCRIPTION INDICATES "THE PACKAGE OF SUTURE IS NOT IN A GOOD CONDITION." "PACKAGE IS ONE AND PRODUCT INSIDE IS ANOTHER." "PRODUCT INSIDE IS BROKEN." COULD YOU PLEASE CONFIRM IF THIS 3 ISSUES OCCURRED AT THE SAME TIME? HOW MANY DEVICES WERE INVOLVED? DID THIS ISSUES OCCURRED PRE-OP OR INTRA-OP? PLEASE EXPLAIN. IF THIS OCCURRED INTRA-OP, WHAT WAS THE PROCEDURE NAME? CONFIRM DEVICE'S AVAILABILITY. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511605&pc=GAM | 2210968-2020-06746 | 2020-08-13 | Malfunction | ETHICON INC. | 2020-09-09 | GAM | VICRYL CT VIO 75CM M3 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION ON (B)(6) 2020AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE OF THE SUTURE BROKE CLOSE TO THE JUNCTION WIRE/NEEDLE, WHEN THE NEEDLE WAS PLACED ON THE NEEDLE HOLDER, ABOVE THE INSTRUMENTATION TABLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10474075&pc=GAM | 2210968-2020-06552 | 2020-08-20 | Malfunction | ETHICON INC. | 2020-08-31 | GAM | VICRYL RAPIDE UND 4/0 70CM SH-1 PLUS | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE DETACHED FROM THE NEEDLE. IT WAS NOT A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME AND DATE? PATIENT'S CONDITION? A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PPBCTJW0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10465779&pc=GAM | 2210968-2020-06489 | 2020-08-14 | Malfunction | ETHICON INC. | 2020-08-28 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN OB-GYN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE EASILY DURING CONTINUOUS SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10449423&pc=GAM | 2210968-2020-06355 | 2020-08-13 | Malfunction | ETHICON INC. | 2020-08-25 | GAM | VICRYL RAP CT UD 75CM M2 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF WHEN TRYING TO PUT THE NEEDLE BACK ON THE NEEDLE SUPPORT WHILE CLOSING THE SKIN. THERE WAS AN UNKNOWN DELAY IN THE PROCEDURE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10444567&pc=GAM | 2210968-2020-06318 | 2020-08-06 | Malfunction | ETHICON INC. | 2020-08-24 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EXERESIS OF A SENTINEL NODE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE CLOSURE OF THE APONEUROSIS, THE NEEDLE BROKE INTO TWO PART. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE EVENTS OR PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV1086, QCBDHGR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: COULD YOU PLEASE EXPLAIN IF THE NEEDLE FELT INTO THE PATIENT? IF YES COULD YOU PLEASE EXPLAIN. WAS THE NEEDLE RETRIEVED? DID THE PATIENT SUFFER FROM ANY CONSEQUENCE? PLEASE EXPLAIN. DEVICE RETURN FOLLOW UP.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10444660&pc=GAM | 2210968-2020-06317 | 2020-08-03 | Malfunction | ETHICON INC. | 2020-08-24 | GAM | VICRYL RAP CT UD 75CM M2 | Break | No Consequences Or Impact To Patient; Foreign Body In Patient; No Code Available | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT NATURAL CHILDBIRTH ON AN UNKNOWN DATE AND SUTURE WAS USED ON A TEAR OF THE PERINEUM. DURING THE PROCEDURE, WHEN STITCHING THE PERINEUM, THE NEEDLE BROKE. THE SUTURE LINE HAD TO BE REOPENED TO FIND THE NEEDLE AND THE PATIENT WAS RESUTURED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE DATE? PLEASE CONFIRM THAT THE NEEDLE WAS RETRIEVED DURING THE SAME PROCEDURE? WHAT IS THE PATIENT¿S CURRENT STATUS? LOT NUMBER OF THE DEVICE? WHAT WAS USED TO COMPLETE THE PROCEDURE? PATIENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10439206&pc=GAM | 2210968-2020-06276 | 2020-07-06 | Malfunction | ETHICON INC. | 2020-08-21 | GAM | VICRYL CT VIO 75CM M3 | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT MAMMECTOMY PROCEDURE IN JULY 2020 AND SUTURE WAS USED TO DO THE CUTANEOUS AND SUBCUTANEOUS THREAD. ON (B)(6) 2020, 10 DAYS AFTER THE SURGERY, THE CUTANEOUS WOUND DEHISCED AND SUTURE BROKE ON THE LENGTH OF THE SUTURE, THE NODES OF THE SUTURE WERE INTACT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE (B)(4); (B)(4) NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE FOLLOWING INFORMATION WAS REQUESTED BUT UNAVAILABLE: WHAT IS THE CURRENT CONDITION OF THE PATIENT? WHAT IS THE EVENT DAY AND PROCEDURE DATE? PLEASE CONFIRM HOW MANY DEVICES WERE INVOLVED IN THIS SINGLE PROCEDURE? DEVICE RETURN FOLLOW UP. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10439251&pc=GAM | 2210968-2020-06275 | 2020-07-06 | Malfunction | ETHICON INC. | 2020-08-21 | GAM | VICRYL CT VIO 75CM M3 | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT OVARIECTOMY PROCEDURE IN (B)(6) 2020 AND SUTURE WAS USED FOR THE LIGATURE OF THE OVARIAN PEDICLE, FOR THE THREAD OF THE ABDOMINAL WALL AND SUBCUTANEOUS TISSUE AND FOR SKIN. ON (B)(6) 2020, 2 DAYS AFTER THE SURGERY, THE MUSCLE WALL DEHISCED WITH EVENTRATION AND THE SUTURE BROKE ON THE LENGTH OF THE SUTURE, THE NODES OF THE ABDOMINAL WALL SUTURE WAS INTACT. IT WAS REPORTED THERE WAS PRESSURE ON THE SUBCUTANEOUS AND CUTANEOUS SUTURE BUT NO RUPTURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV586; PPBCZJR0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE FOLLOWING INFORMATION WAS REQUESTED BUT UNAVAILABLE: WHAT IS THE CURRENT CONDITION OF THE PATIENT? WHAT IS THE EVENT DAY AND PROCEDURE DATE? PLEASE CONFIRM HOW MANY DEVICES WERE INVOLVED IN THIS SINGLE PROCEDURE? DEVICE RETURN FOLLOW UP. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10439590&pc=GAM | 2210968-2020-06299 | 2020-07-28 | Malfunction | ETHICON INC. | 2020-08-21 | GAM | VICRYL VIO 70CM M2::USP3/0 SGLE ARMED FS-2 | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT REMOVAL OF MELANOMA PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE INITIAL PROCEDURE? REMOVAL OF MELANOMA ON BACK UNDER LOCAL ANESTHESIA (SUBCUTANEOUS SUTURE). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10428043&pc=GAM | 2210968-2020-06221 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-08-19 | GAM | VICRYL RAP CT UD 75CM M3 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN EVENTRATION PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POSTOPERATIVELY, THE SUTURE WAS COMPLETELY REABSORBED AFTER 9 DAYS. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. COULD YOU PLEASE EXPLAIN IF THE PATIENT SUFFER FROM WOUND DEHISCENCE DUE TO THE SUTURE REABSORPTION? ON WHAT TISSUE WAS SUTURE PLACED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10428071&pc=GAM | 2210968-2020-06222 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-08-19 | GAM | VICRYL RAP CT UD 75CM M3 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN EVENTRATION PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POSTOPERATIVELY, THE SUTURE WAS COMPLETELY REABSORBED AFTER 9 DAYS. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. COULD YOU PLEASE EXPLAIN IF THE PATIENT SUFFER FROM WOUND DEHISCENCE DUE TO THE SUTURE REABSORPTION? ON WHAT TISSUE WAS SUTURE PLACED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10428078&pc=GAM | 2210968-2020-06220 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-08-19 | GAM | VICRYL RAP CT UD 75CM M3 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN EVENTRATION PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. POSTOPERATIVELY, THE SUTURE WAS COMPLETELY REABSORBED AFTER 9 DAYS. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. COULD YOU PLEASE EXPLAIN IF THE PATIENT SUFFER FROM WOUND DEHISCENCE DUE TO THE SUTURE REABSORPTION? ON WHAT TISSUE WAS SUTURE PLACED?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10422236&pc=GAM | 2210968-2020-06182 | 2020-08-06 | Malfunction | ETHICON INC. | 2020-08-18 | GAM | VICRYL RAPIDE SUTURE 27"(70CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient; No Code Available | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NORMAL BIRTH DELIVERY ON (B)(6) 2020 AND SUTURE WAS USED ON THE PERINEAL INCISION. DURING PERINEAL SUTURING, THE SUTURE WAS DETACHED FROM THE NEEDLE. AT THE TIME OF SUTURING, THE WHOLE NEEDLE WAS DEEPLY INSERTED INTO THE TISSUE, THEN, IT WAS BURIED IN THE SUBCUTANEOUS TISSUE. THE SURGEON PULLED THE SUTURE TO REMOVE THE NEEDLE FROM THE TISSUE. AT THAT TIME, THE SUTURE WAS DETACHED FROM THE NEEDLE. SO, THE SURGEON HAD IMPRESSION THAT THE NEEDLE WAS DETACHED FROM THE SUTURE EASILY. THE OPERATION TIME WAS EXTENDED WITHIN 30 MINUTES. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW WAS THE NEEDLE RETRIEVED FROM THE PATIENT? NO FURTHER INFORMATION IS AVAILABLE. HOW THE PROCEDURE WAS COMPLETE? NO FURTHER INFORMATION IS AVAILABLE. PATIENT¿S CURRENT STATUS? NO FURTHER INFORMATION IS AVAILABLE. PLEASE PROVIDE PROCEDURE DATE. NO FURTHER INFORMATION IS AVAILABLE. PLEASE PROVIDE LOT NUMBER. NO FURTHER INFORMATION IS AVAILABLE. STATUS OF PRODUCT RETURN / FOLLOW UP. WE REGULARLY CONTACT WITH SALES REP ABOUT THE DEVICE RETURNING. NO FURTHER INFORMATION WILL BE PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IS THE NEEDLE RETAINED IN THE PATIENT? WAS THE PATIENT TREATMENT OR POST-OPERATIVE CARE ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME OF 30 MINUTES? IF YES, PLEASE EXPLAIN. WAS THERE ANY PATIENT IMPACT (IE. ADDITIONAL TREATMENT REQUIRED) AND/OR ADVERSE PATIENT OUTCOME DUE TO THE PROLONGED SURGERY? WHAT IS THE PATIENTS CURRENT STATUS?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398489&pc=GAM | 2210968-2020-06079 | 2020-08-05 | Malfunction | ETHICON INC. | 2020-08-12 | GAM | VICRYL RAP UND 90CM M3 2/0 V-34 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EPISIOTOMY SURGERY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, NEEDLE DISLODGED FROM SUTURE WHILE SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398495&pc=GAM | 2210968-2020-06076 | 2020-08-05 | Malfunction | ETHICON INC. | 2020-08-12 | GAM | VICRYL RAP UND 90CM M3 2/0 V-34 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN EPISIOTOMY SURGERY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, NEEDLE DISLODGED FROM SUTURE WHILE SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE NEEDLE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? EVENT DATE? WAS THERE ANY PATIENT CONSEQUENCE OR INJURY? WHAT IS THE CONDITION OF THE PATIENT? PROCEDURE DATE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398980&pc=GAM | 2210968-2020-06102 | 2020-08-03 | Malfunction | ETHICON INC. | 2020-08-12 | GAM | VICRYL RAPID 3/0 70CM 24MM FS1 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C-SECTION SURGERY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE SNAPPED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS REPORTED "SUTURE SNAPPED DURING OPERATION. TENSILE STRENGTH OF SUTURE NOT LIKE NORMAL BROKE WITH NORMAL TUGGING". WAS THERE ANY ADVERSE CONSEQUENCE ASSOCIATED WITH THE PATIENT? WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN SUTURE BROKE DURING THE PROCEDURE? PLEASE CLARIFY WHAT DO YOU MEAN BY "BROKE WITH NORMAL TUGGING". EVENT DATE? PROCEDURE DATE? A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10390117&pc=GAM | 2210968-2020-05989 | 2020-07-29 | Malfunction | ETHICON INC. | 2020-08-11 | GAM | VICRYL CTD VIO 75CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS FOUND TO BE BROKEN. THE NEEDLE TIP WAS UNABLE TO BE LOCATED TO DATE. XRAY'S HAVE BEEN TAKEN OF THE AREA HOWEVER, AFTER INVESTIGATING AN AREA SHOWING A SIMILAR DENSITY TO THE NEEDLE, IT WAS NOT ABLE TO BE LOCATED. A CT SCAN HAS BEEN BOOKED FOR LATER IN THE WEEK. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE NEEDLE FOUND? WHAT WAS THE SIZE OF THE NEEDLE USED? WAS MORE THAN HALF THE NEEDLE RETAINED IN THE PATIENT? WAS ADDITIONAL DISSECTION REQUIRED IN OTHER ORGANS/TISSUES OTHER THAN THE TARGET TISSUE? WAS THERE ANY CHANGE IN THE PATIENT¿S POST-OPERATIVE CARE DUE TO THE PROLONGED PROCEDURE? A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10390130&pc=GAM | 2210968-2020-06004 | 2020-07-24 | Malfunction | ETHICON INC. | 2020-08-11 | GAM | VICRYL*VLT PLUS 4-0 70CM 1-RB1 | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE SURGERY, THE SUTURE IS DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE INITIAL PROCEDURE? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10390131&pc=GAM | 2210968-2020-06006 | 2020-07-24 | Malfunction | ETHICON INC. | 2020-08-11 | GAM | VICRYL*VLT PLUS 4-0 70CM 1-RB1 | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE SURGERY, THE SUTURE IS DETACHED FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE INITIAL PROCEDURE? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10384696&pc=GAM | 2210968-2020-06002 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-08-10 | GAM | VICRYL CT VIO 70CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE ON AND UNKNOWN DATE AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE DURING THE KNOT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10367686&pc=GAM | 2210968-2020-05877 |  | Malfunction | ETHICON INC. | 2020-08-05 | GAM | VICRYL* UND 3-0 70CM (1)SC-20 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CESAREAN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, AT THE TIME OF THE SUBCUTANEOUS SUTURE, THERE WAS A BREAK IN THE BASE OF THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? WHAT IS CURRENT CONDITION OF THE PATIENT? WHAT WAS USED TO COMPLETE THE PROCEDURE? WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? PROCEDURE DATE? EVENT DATE? DEVICE RETURN STATUS/FOLLOW UP? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364573&pc=GAM | 2210968-2020-05844 | 2020-06-25 | Malfunction | ETHICON INC. | 2020-08-04 | GAM | VICRYL VIO 70CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DIGESTIVE COELIOSCOPIC PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. THE SUTURE BROKE DURING THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # ==> (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS USED TO COMPLETE THE PROCEDURE? WAS THE SURGERY COMPLETED SUCCESSFULLY? WHAT IS THE CURRENT CONDITION OF THE PATIENT? DID THE PATIENT SUFFER ANY CONSEQUENCE? COULD YOU PLEASE CONFIRM DE DEVICE'S AVAILABILITY?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353395&pc=GAM | 2210968-2020-05758 | 2020-07-15 | Malfunction | ETHICON INC. | 2020-07-31 | GAM | VICRYL RAP UND 90CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. THE THREAD PULLED OFF THE NEEDLE, CAUSING DELAY AND ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: COULD YOU PLEASE CLARIFY WHAT DO YOU MEAN BY "THE SUTURE HAD TO BE RE-PERFORMED"? PLEASE EXPLAIN. THE SUTURE HAD TO BE PERFORMED AGAIN (A SECOND TIME BECAUSE THE ISSUE OCCURRED DURING THE FIRST ATTEMPT). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE INITIAL PROCEDURE? HOW MANY MINUTES DOES THE PROCEDURE WAS DELAYED? ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353401&pc=GAM | 2210968-2020-05768 | 2020-07-20 | Malfunction | ETHICON INC. | 2020-07-31 | GAM | VICRYL VIO 90CM M4 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A HYSTEROGRAPHY ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE STRAND PULLED OFF THE NEEDLE AND THE NEEDLE WAS RETRIEVED IMMEDIATELY. THERE WERE NO ADVERSE CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10339485&pc=GAM | 2210968-2020-05671 | 2020-07-20 | Malfunction | ETHICON INC. | 2020-07-29 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN OOPHORECTOMY PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. AFTER FEW TISSUE STITCHES IN THE ABDOMINAL WALL, THE NEEDLE PULLED OFF THE THREAD WITHOUT ANY EXCESSIVE FORCE APPLICATION. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10325636&pc=GAM | 2210968-2020-05566 |  | Malfunction | ETHICON INC. | 2020-07-27 | GAM | VICRYL RAP CT UD 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE LOOSENED FROM THE SUTURE WHILE SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SUTURE WAS IN CONTACT WITH THE PATIENT? HOW THE PROCEDURE WAS COMPLETED. WAS THERE ANY ADVERSE CONSEQUENCE ASSOCIATED WITH THE PATIENT? PLEASE PROVIDE PROCEDURE NAME AND DATE. PLEASE PROVIDE EVENT DATE. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PMBDMJS0, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10325812&pc=GAM | 2210968-2020-05575 | 2020-07-12 | Malfunction | ETHICON INC. | 2020-07-27 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC APPENDECTOMY PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE INTO THREE PARTS WHEN SEWING THE SKIN AT UMBILICAL FORAMEN. THE 1.5 MM OF THE NEEDLE TIP WAS MISSING AND NOT FOUND. NO METAL IMAGE IN UMBILICUS WAS CONFIRMED THROUGH X-RAY. THE PATIENT WAS SENT BACK TO WARD. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WILL THERE BE ANY ADDITIONAL X-RAY OR OTHER TYPE OF PROCEDURE TO LOCATE THE BROKEN NEEDLE? WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? DID THE NEEDLE COME IN TO CONTACT WITH ANOTHER INSTRUMENT DURING THE PROCEDURE? WHAT IS THE PATIENT¿S CURRENT HEALTH STATUS AND CONDITION? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10325892&pc=GAM | 2210968-2020-05580 | 2020-07-16 | Malfunction | ETHICON INC. | 2020-07-27 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT A PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, AFTER A FEW TISSUE STITCHES, THE STRAND BROKE IN ITS LENGTH (AND NOT AT THE SUTURE AREA) WITHOUT ANY EXCESSIVE FORCE APPLICATION. USE OF ANOTHER DEVICE TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER (JV586 ; QABCTWR0), AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS USED TO COMPLETE THE PROCEDURE? WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE DURING THE PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10325962&pc=GAM | 2210968-2020-05583 | 2020-07-15 | Malfunction | ETHICON INC. | 2020-07-27 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT A PROCEDURE ON (B)(6) 2020 AND A SUTURE WAS USED. DURING THE PROCEDURE, AFTER FEW TISSUE STITCHES IN THE ABDOMINAL WALL, THE NEEDLE PULLED OFF THE THREAD WITHOUT ANY EXCESSIVE FORCE APPLICATION. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER (JV452; PMBDKGR0), AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10324033&pc=GAM | 2210968-2020-05559 |  | Malfunction | ETHICON INC. | 2020-07-24 | GAM | VICRYL RAP UND 90CM M4::USP1 S/A CTX PLUS | Appropriate Term/Code Not Available | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A NORMAL BIRTH DELIVERY ON UNKNOWN DATE AND THE SUTURE WAS USED FOR PERINEUM BY CONTINUOUS SUTURING. DURING DELIVERY, IT WAS FOUND THAT THE LENGTH OF THE SUTURE HAD BEEN SHORTER THAN THE SPECIFIED LENGTH OF THE SUTURE. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED (B)(4) NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308705&pc=GAM | 2210968-2020-05464 | 2020-06-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE DOG UNDERWENT AN OVARIECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED TO PERFORM OVARIAN PEDICLE LIGATURES, ABDOMINAL WALL SUTURES AND CUTANEOUS SUTURES. DURING THE PROCEDURE, AFTER A FEW TISSUE PASSAGES, THE NEEDLE PULLED OFF, WITHOUT AN EXCESSIVE STRAIN. SURGERY WAS COMPLETED, USING A NEW WIRE, WITH NO IMPACT ON THE ANIMAL. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308717&pc=GAM | 2210968-2020-05465 | 2020-06-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE DOG UNDERWENT AN OVARIECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED TO PERFORM OVARIAN PEDICLE LIGATURES, ABDOMINAL WALL SUTURES AND CUTANEOUS SUTURES. DURING THE PROCEDURE, AFTER A FEW TISSUE PASSAGES, THE NEEDLE PULLED OFF, WITHOUT AN EXCESSIVE STRAIN. SURGERY WAS COMPLETED, USING A NEW WIRE, WITH NO IMPACT ON THE ANIMAL. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308719&pc=GAM | 2210968-2020-05466 | 2020-06-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE DOG UNDERWENT AN OVARIECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED TO PERFORM OVARIAN PEDICLE LIGATURES, ABDOMINAL WALL SUTURES AND CUTANEOUS SUTURES. DURING THE PROCEDURE, AFTER A FEW TISSUE PASSAGES, THE NEEDLE PULLED OFF, WITHOUT AN EXCESSIVE STRAIN. SURGERY WAS COMPLETED, USING A NEW WIRE, WITH NO IMPACT ON THE ANIMAL. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308722&pc=GAM | 2210968-2020-05467 | 2020-06-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE DOG UNDERWENT AN OVARIECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED TO PERFORM OVARIAN PEDICLE LIGATURES, ABDOMINAL WALL SUTURES AND CUTANEOUS SUTURES. DURING THE PROCEDURE, AFTER A FEW TISSUE PASSAGES, THE NEEDLE PULLED OFF, WITHOUT AN EXCESSIVE STRAIN. SURGERY WAS COMPLETED, USING A NEW WIRE, WITH NO IMPACT ON THE ANIMAL. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308723&pc=GAM | 2210968-2020-05469 | 2020-06-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE DOG UNDERWENT AN OVARIECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED TO PERFORM OVARIAN PEDICLE LIGATURES, ABDOMINAL WALL SUTURES AND CUTANEOUS SUTURES. DURING THE PROCEDURE, AFTER A FEW TISSUE PASSAGES, THE NEEDLE PULLED OFF, WITHOUT AN EXCESSIVE STRAIN. SURGERY WAS COMPLETED, USING A NEW WIRE, WITH NO IMPACT ON THE ANIMAL. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308725&pc=GAM | 2210968-2020-05468 | 2020-06-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE DOG UNDERWENT AN OVARIECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED TO PERFORM OVARIAN PEDICLE LIGATURES, ABDOMINAL WALL SUTURES AND CUTANEOUS SUTURES. DURING THE PROCEDURE, AFTER A FEW TISSUE PASSAGES, THE NEEDLE PULLED OFF, WITHOUT AN EXCESSIVE STRAIN. SURGERY WAS COMPLETED, USING A NEW WIRE, WITH NO IMPACT ON THE ANIMAL. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308734&pc=GAM | 2210968-2020-05470 | 2020-06-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A FEMALE DOG UNDERWENT AN OVARIECTOMY PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED TO PERFORM OVARIAN PEDICLE LIGATURES, ABDOMINAL WALL SUTURES AND CUTANEOUS SUTURES. DURING THE PROCEDURE, AFTER A FEW TISSUE PASSAGES, THE NEEDLE PULLED OFF, WITHOUT AN EXCESSIVE STRAIN. SURGERY WAS COMPLETED, USING A NEW WIRE, WITH NO IMPACT ON THE ANIMAL. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311010&pc=GAM | 2210968-2020-05488 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311036&pc=GAM | 2210968-2020-05485 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311125&pc=GAM | 2210968-2020-05486 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311132&pc=GAM | 2210968-2020-05487 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486 AND 2210968-2020-05488 IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311229&pc=GAM | 2210968-2020-05482 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488 IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311236&pc=GAM | 2210968-2020-05478 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311238&pc=GAM | 2210968-2020-05476 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311246&pc=GAM | 2210968-2020-05483 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311284&pc=GAM | 2210968-2020-05475 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311292&pc=GAM | 2210968-2020-05479 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311367&pc=GAM | 2210968-2020-05477 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311374&pc=GAM | 2210968-2020-05481 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311402&pc=GAM | 2210968-2020-05480 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-07-22 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NOTE: EVENTS REPORTED IN 2210968-2020-04851, 2210968-2020-05475, 2210968-2020-05476, 2210968-2020-05477, 2210968-2020-05478, 2210968-2020-05479, 2210968-2020-05480, 2210968-2020-05481, 2210968-2020-05482, 2210968-2020-05483, 2210968-2020-05485, 2210968-2020-05486, 2210968-2020-05487 AND 2210968-2020-05488. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10305824&pc=GAM | 2210968-2020-05446 |  | Malfunction | ETHICON INC. | 2020-07-21 | GAM | VICRYL CT VIO 75CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT A CAT UNDERWENT AN OVARIECTOMY IN (B)(6) 2020 AND THE SUTURE WAS USED. AFTER 2 OR 3 DAYS FOLLOWING THE SURGERY, THE BREAKAGE OF THE SUTURE OCCURRED. THE SUTURE BROKE ON THE WIRE LENGTH; THE NODES OF THE CUTANEOUS THREAD WERE INTACT. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY FOR EACH FILE/CAT HOW MANY VICRYL SUTURE DEVICES BROKE POST-OP ON ONE CAT? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10305841&pc=GAM | 2210968-2020-05447 |  | Malfunction | ETHICON INC. | 2020-07-21 | GAM | VICRYL CT VIO 75CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT A CAT UNDERWENT AN OVARIECTOMY IN (B)(6) 2020 AND THE SUTURE WAS USED. AFTER 2 OR 3 DAYS FOLLOWING THE SURGERY, THE BREAKAGE OF THE SUTURE OCCURRED. THE SUTURE BROKE ON THE WIRE LENGTH; THE NODES OF THE CUTANEOUS THREAD WERE INTACT. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY FOR EACH FILE/CAT HOW MANY VICRYL SUTURE DEVICES BROKE POST-OP ON ONE CAT? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10305848&pc=GAM | 2210968-2020-05448 |  | Malfunction | ETHICON INC. | 2020-07-21 | GAM | VICRYL CT VIO 75CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT A CAT UNDERWENT AN OVARIECTOMY IN (B)(6) 2020 AND THE SUTURE WAS USED. AFTER 2 OR 3 DAYS FOLLOWING THE SURGERY, THE BREAKAGE OF THE SUTURE OCCURRED. THE SUTURE BROKE ON THE WIRE LENGTH; THE NODES OF THE CUTANEOUS THREAD WERE INTACT. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY FOR EACH FILE/CAT HOW MANY VICRYL SUTURE DEVICES BROKE POST-OP ON ONE CAT? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10305872&pc=GAM | 2210968-2020-05449 |  | Malfunction | ETHICON INC. | 2020-07-21 | GAM | VICRYL CT VIO 75CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT A CAT UNDERWENT AN OVARIECTOMY IN (B)(6) 2020 AND THE SUTURE WAS USED. AFTER 2 OR 3 DAYS FOLLOWING THE SURGERY, THE BREAKAGE OF THE SUTURE OCCURRED. THE SUTURE BROKE ON THE WIRE LENGTH; THE NODES OF THE CUTANEOUS THREAD WERE INTACT. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY FOR EACH FILE/CAT HOW MANY VICRYL SUTURE DEVICES BROKE POST-OP ON ONE CAT? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10305954&pc=GAM | 2210968-2020-05452 | 2020-07-01 | Malfunction | ETHICON INC. | 2020-07-21 | GAM | VICRYL VIO 45CM M1.5::USP4/0 SGLE ARM FS-2 | Break | Discomfort |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A SURGERY FOR LIP TRAUMA ON (B)(6) 2020 AND THE SUTURE WAS USED AS A SINGLE SUTURE. SHORTLY AFTER HAVING LEFT THE CLINIC THE PATIENT RETURNED AS THE SUTURE HAS COME OUT, BROKE. A NEW ANESTHETIC AND A NEW SUTURE WERE ADMINISTERED, THIS TIME WITH EXTRA ATTENTION TO TIGHTENING. SOON AFTER THE PATIENT RETURNS HOME THEY CALL THE CLINIC BECAUSE THE SUTURE HAS COME OUT AGAIN. PATIENT WAS REQUESTED TO CONTACT PRIMARY CARE FOR POSSIBLE REFERRAL TO THE ER, SINCE THE INJURY IS AT THE TRANSITION BETWEEN THE INNER AND OUTER LIP. THE DECISION WAS A SPONTANEOUS HEALING WITHOUT SUTURE, DESPITE THAT A FLAP CAN EASILY BE PULLED AWAY FROM THE WOUND. PATIENT IS THEREBY IMPACTED BY REPEATED TREATMENTS WITHOUT THE DESIRED RESULT. IT IS UNKNOWN WHETHER LACK OF SUTURE WILL HAVE LONG TERM CONSEQUENCE. THE PATIENT EXPOSED TO NEEDLESS DISCOMFORT. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT TYPE OF THE LIP TRAUMA WAS REPAIRED INITIALLY? WHERE WAS THE SUTURE APPLIED ON THE LIP? WHAT WAS THE TISSUE CONDITION, I.E., NORMAL OR THIN, CALCIFIED, FRAGILE, DISEASED? HOW WAS THE SUTURE PLACED, INTERRUPTED OR CONTINUOUS? HOW THE SUTURE WAS INITIALLY TIED? WHAT SUTURE TECHNIQUE WAS USED DURING INITIAL SUTURING AND RE-DO SURGERY? DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE OR DURING. THE INITIAL SUTURE PLACEMENT? WAS THERE ANY PRECIPITATING STRESS FACTOR FOR THE SUTURE BREAKAGE 1ST TIME? IN WHAT PLACE THE SUTURE BROKE FIRST TIME? WAS THERE ANY PRECIPITATING STRESS FACTOR FOR THE SUTURE BREAKAGE 2ND TIME? WHERE WAS THE REPLACED SUTURE BROKE? PLEASE SPECIFY WHAT WAS A DEFECT THERE? HOW WAS THE SKIN FLAP MANAGED? PLEASE DESCRIBE ¿REPEATED TREATMENTS¿ AFTER 2ND TIME BREAKAGE? WHAT IS PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THESE EVENTS? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS? WHAT IS THE PATIENT CURRENT STATUS? ADDITIONALLY, THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. IT WAS REPORTED THAT ¿SEVERAL SUTURE THREADS WITH THE SAME LOT NUMBERS ARE PRONE TO BE DEFECTIVE¿. WERE THESE EVENTS REPORTED PREVIOUSLY? IF YES, PLEASE PROVIDE PC NUMBERS. IF NOT, PLEASE CREATE ADDITIONAL FILES TO CAPTURE EACH ADDITIONAL EVENT FOR EACH ONE PATIENT INVOLVED? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADVERSE EVENT REGARDING 1 ST SUTURE BREAKAGE POST-OP AND RE-SUTURING WAS SUBMITTED VIA 2210968-2020-05451. DEVICE EVENT REGARDING 2ND SUTURE BREAKAGE POST-OP AND DISCOMFORT WAS SUBMITTED VIA 2210968-2020-05452.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10299423&pc=GAM | 2210968-2020-05408 | 2020-07-06 | Malfunction | ETHICON INC. | 2020-07-20 | GAM | VICRYL RAP UND1/5-0 45CM 1A.P1 | Contamination /Decontamination Problem | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. PRIOR TO USE, THE NURSE NOTED THAT A STRAND OF HAIR WAS FOUND ON THE PAPER SUTURE PACKAGING. THE SUTURE INSIDE WAS NOT USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A PHOTO UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PHOTO EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10290328&pc=GAM | 2210968-2020-05369 |  | Malfunction | ETHICON INC. | 2020-07-17 | GAM | VICRYL VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE CUSTOMER NOTICED A BOX OF THE DEVICE WITH THE TAG ¿NOT FOR HUMAN USE¿; THE ISSUE WAS REGARDING THE LABEL ON THE BOX. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NON-CONFORMANCE WHICH WAS RAISED TO ADDRESS THE EVENT REPORTED. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. IT WAS REPORTED PACKAGING LABELING IDENTIFICATION. A PICTURE WAS RECEIVED FOR EVALUATION. UPON VISUAL INSPECTION A BOX OF PRODUCT WITH A LEGEND OF "NOT FOR HUMAN USE"COULD BE OBSERVED. ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10284101&pc=GAM | 2210968-2020-05355 |  | Malfunction | ETHICON INC. | 2020-07-16 | GAM | VICRYL VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CESAREAN SECTION PROCEDURE ON AN UNKNOWN DATE; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE IN THE MIDDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT IS THE EVENT DAY AND PROCEDURE DATE? UNKNOWN. DOES A PIECE OF THE NEEDLE REMAIN IN THE PATIENT¿S? IF YES, WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? NO, REMAINING PARTS WERE REPORTED. WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? NO. WHAT IS CURRENT CONDITION OF THE PATIENT? UNKNOWN. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PPBBBCR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10284648&pc=GAM | 2210968-2020-05358 | 2020-07-04 | Malfunction | ETHICON INC. | 2020-07-16 | GAM | VICRYL VIO 90CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE IN THE UTERUS. THERE WERE FRAGMENTS GENERATED BUT WERE REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH AN UNSPECIFIED DELAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. HOW MANY MINUTES WAS THE PROCEDURE DELAYED DUE TO THE NEEDLE BREAKAGE? A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PMBCHTQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271906&pc=GAM | 2210968-2020-05271 | 2020-06-25 | Malfunction | ETHICON INC. | 2020-07-14 | GAM | VICRYL VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: IS THE NEEDLE PIECE RETAINED IN THE PATIENT? IF YES, PLEASE EXPLAIN. HOW WAS THE NEEDLE RETRIEVED FROM THE PATIENT? WAS THERE ANY ADVERSE CONSEQUENCE ASSOCIATED WITH THE PATIENT? NO. PLEASE PROVIDE PROCEDURE DATE: (B)(6) 2020. PLEASE PROVIDE LOT NUMBER: PPBDJ7Q0. STATUS OF PRODUCT RETURN / FOLLOW UP: THE PRODUCT IS AVAILABLE BUT NOW IT STILL AT THE HOSPITAL SITE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10272113&pc=GAM | 2210968-2020-05274 | 2020-06-01 | Malfunction | ETHICON INC. | 2020-07-14 | GAM | VICRYL CT VIO 75CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SITE IMPLANTABLE ENDOVENUS IMPLANTATION PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE SPONTANEOUSLY PULLED OFF DURING ITS USE ON THE NEEDLE HOLDER WITHOUT AN ABNORMAL TENSION ON THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE EVENT DAY AND PROCEDURE DATE? THE PRECISE DATE IS NOT KNOWN. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PJBBMDR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10263342&pc=GAM | 2210968-2020-05228 | 2020-06-30 | Malfunction | ETHICON INC. | 2020-07-13 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROSTATECTOMY PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS OBSERVED TO BE MISSING, POSSIBLY ALREADY DETACHED IN THE SACHET OR DETACHED AT TIME OF ARMING THE NEEDLE HOLDER. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PMU043, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10233091&pc=GAM | 2210968-2020-05072 | 2020-06-18 | Malfunction | ETHICON INC. | 2020-07-06 | GAM | VICRYL CT VIO 75CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SITE IMPLANTABLE ENDOVENUS IMPLANTATION PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF DURING ITS USE ON THE NEEDLE HOLDER WITHOUT AN ABNORMAL TENSION ON THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE INITIAL PROCEDURE? PROCEDURE: A SITE IMPLANTABLE ENDOVENUS IMPLANTATION. THE CUSTOMER REPORTS THAT A NEEDLE SPONTANEOUSLY PULLED OFF DURING ITS USE ON THE NEEDLE HOLDER WITHOUT AN ABNORMAL TENSION ON THE WIRE. WHEN DID THE EVENT OCCURRED? PRE-OP OR INTRA-OP. WHAT IS THE EVENT DAY AND PROCEDURE DATE? DATE OF THE PROCEDURE: 15 DAYS AGO. COULD YOU PLEASE CONFIRM IF DID THE PATIENT SUFFER FROM INJURY OR CONSEQUENCE DUE TO THE PULL OFF SUTURE NEEDLE? NO PATIENT CONSEQUENCE BECAUSE THE INCIDENT OCCURRED BEFORE THE PROCEDURE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10233095&pc=GAM | 2210968-2020-05074 | 2020-06-18 | Malfunction | ETHICON INC. | 2020-07-06 | GAM | VICRYL CT VIO 75CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SITE IMPLANTABLE ENDOVENUS IMPLANTATION PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF DURING ITS USE ON THE NEEDLE HOLDER WITHOUT AN ABNORMAL TENSION ON THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE INITIAL PROCEDURE? PROCEDURE : A SITE IMPLANTABLE ENDOVENUS IMPLANTATION. THE CUSTOMER REPORTS THAT A NEEDLE SPONTANEOUSLY PULLED OFF DURING ITS USE ON THE NEEDLE HOLDER WITHOUT AN ABNORMAL TENSION ON THE WIRE. WHEN DID THE EVENT OCCURRED? PRE-OP OR INTRA-OP. WHAT IS THE EVENT DAY AND PROCEDURE DATE? DATE OF THE PROCEDURE: 15 DAYS AGO. COULD YOU PLEASE CONFIRM IF DID THE PATIENT SUFFER FROM INJURY OR CONSEQUENCE DUE TO. THE PULL OFF SUTURE NEEDLE? NO PATIENT CONSEQUENCE BECAUSE THE INCIDENT OCCURRED BEFORE THE PROCEDURE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10233098&pc=GAM | 2210968-2020-05073 | 2020-06-18 | Malfunction | ETHICON INC. | 2020-07-06 | GAM | VICRYL CT VIO 75CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SITE IMPLANTABLE ENDOVENOUS IMPLANTATION PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF DURING ITS USE ON THE NEEDLE HOLDER WITHOUT AN ABNORMAL TENSION ON THE WIRE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE INITIAL PROCEDURE? PROCEDURE : A SITE IMPLANTABLE ENDOVENUS IMPLANTATION. THE CUSTOMER REPORTS THAT A NEEDLE SPONTANEOUSLY PULLED OFF DURING ITS USE ON THE NEEDLE HOLDER WITHOUT AN ABNORMAL TENSION ON THE WIRE. WHEN DID THE EVENT OCCURRED? PRE-OP OR INTRA-OP. WHAT IS THE EVENT DAY AND PROCEDURE DATE? DATE OF THE PROCEDURE: 15 DAYS AGO. COULD YOU PLEASE CONFIRM IF DID THE PATIENT SUFFER FROM INJURY OR CONSEQUENCE DUE TO THE PULL OFF SUTURE NEEDLE? NO PATIENT CONSEQUENCE BECAUSE THE INCIDENT OCCURRED BEFORE THE PROCEDURE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198297&pc=GAM | 2210968-2020-04832 |  | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT UPON DELIVERY THE CUSTOMER SAW THAT ON THE BOX WAS NOTED "NOT FOR HUMAN USE." HOWEVER THE CORRECT ITEM NUMBER WAS STATED ON THE PACKAGE. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: A PHOTO WAS RECEIVED FOR EVALUATION. UPON VISUAL INSPECTION BOXES OF PRODUCT CODE V371, LOT # PHBBJHR0 WITH A LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS WERE SUBMITTED VIA: 2210968-2020-04826, 2210968-2020-04827, 2210968-2020-04828, 2210968-2020-04829, 2210968-2020-04830, 2210968-2020-04832, 2210968-2020-04834. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198316&pc=GAM | 2210968-2020-04828 |  | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT UPON DELIVERY THE CUSTOMER SAW THAT ON THE BOX WAS NOTED "NOT FOR HUMAN USE." HOWEVER THE CORRECT ITEM NUMBER WAS STATED ON THE PACKAGE. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: A PHOTO WAS RECEIVED FOR EVALUATION. UPON VISUAL INSPECTION BOXES OF PRODUCT CODE V371, LOT # PHBBJHR0 WITH A LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS WERE SUBMITTED VIA: 2210968-2020-04826, 2210968-2020-04827, 2210968-2020-04828, 2210968-2020-04829, 2210968-2020-04830, 2210968-2020-04832, 2210968-2020-04834. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198325&pc=GAM | 2210968-2020-04841 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198328&pc=GAM | 2210968-2020-04840 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198333&pc=GAM | 2210968-2020-04838 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198335&pc=GAM | 2210968-2020-04837 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198341&pc=GAM | 2210968-2020-04836 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198345&pc=GAM | 2210968-2020-04835 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198346&pc=GAM | 2210968-2020-04833 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198352&pc=GAM | 2210968-2020-04831 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198383&pc=GAM | 2210968-2020-04839 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10200266&pc=GAM | 2210968-2020-04843 | 2020-06-15 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL RAP UND 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF THE THREAD. THE NEEDLE STAYED STUCK ON THE TISSUE OF THE PATIENT. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10200865&pc=GAM | 2210968-2020-04851 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE DEVICE HAD THE TAG "NOT FOR HUMAN USE" HAS BEEN USED ON THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: -WHAT DO YOU MEAN BY "A DEVICE WITH THE TAG "NOT FOR HUMAN USE" HAS BEEN USED ON A PATIENT."? THE CUSTOMER RECEIVED SEVERAL BOXES OF THE REFERENCE V372H / LOT PHBBCJQ0 WITH THE LABEL "NOT FOR HUMAN USE". ONE SUTURE OF THESES BOXES HAS BEEN USED ON A PATIENT. WAS THE ETHICON SUTURE DEVICE BRAND NEW? YES. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? HOW MANY BOXES DID THE CUSTOMER RECEIVE THAT WERE LABELED "NOT FOR HUMAN USE"? IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10200969&pc=GAM | 2210968-2020-04847 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: EVENTS REPORTED VIA MW# 2210968-2020-04846, 2210968-2020-04847, 2210968-2020-04848, 2210968-2020-04849.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10200992&pc=GAM | 2210968-2020-04848 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: EVENTS REPORTED VIA MW# 2210968-2020-04846, 2210968-2020-04847, 2210968-2020-04848, 2210968-2020-04849.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201031&pc=GAM | 2210968-2020-04846 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: EVENTS REPORTED VIA MW# 2210968-2020-04846, 2210968-2020-04847, 2210968-2020-04848, 2210968-2020-04849.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201040&pc=GAM | 2210968-2020-04849 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PACKAGING WAS MARKED 'NOT FOR HUMAN USE'. THERE WAS NO PATIENT INVOLVEMENT REPORTED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: EVENTS REPORTED VIA MW# 2210968-2020-04846, 2210968-2020-04847, 2210968-2020-04848, 2210968-2020-04849.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201492&pc=GAM | 2210968-2020-04872 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE BOX OF THE SUTURE PRODUCT HAD THE LABEL "NOT FOR HUMAN USE". IT WAS NOT REPORTED IF THERE WAS ANY PATIENT INVOLVEMENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? DID THIS ISSUE OCCUR IN ONE PATIENT PROCEDURE? INVESTIGATION SUMMARY: ONE PICTURE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF ONE PICTURE, A SEALED BOX OF PRODUCT CODE V372H WITH LOT PHBBCJQ0 AND A PRINTED THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE TO ADDRESS THE EVENT REPORTED UNDER (B)(4). NOTE: EVENTS REPORTED VIA MW # 2210968-2020-04870, 2210968-2020-04871, 2210968-2020-04873, 2210968-2020-04874, 2210968-2020-04875, 2210968-2020-04876.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201494&pc=GAM | 2210968-2020-04871 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE BOX OF THE SUTURE PRODUCT HAD THE LABEL "NOT FOR HUMAN USE". IT WAS NOT REPORTED IF THERE WAS ANY PATIENT INVOLVEMENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? DID THIS ISSUE OCCUR IN ONE PATIENT PROCEDURE? INVESTIGATION SUMMARY: ONE PICTURE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF ONE PICTURE, A SEALED BOX OF PRODUCT CODE V372H WITH LOT PHBBCJQ0 AND A PRINTED THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE TO ADDRESS THE EVENT REPORTED UNDER (B)(4).. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: EVENTS REPORTED VIA MW # 2210968-2020-04870, 2210968-2020-04871, 2210968-2020-04872, 2210968-2020-04873, 2210968-2020-04874, 2210968-2020-04875, 2210968-2020-04876.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201502&pc=GAM | 2210968-2020-04873 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE BOX OF THE SUTURE PRODUCT HAD THE LABEL "NOT FOR HUMAN USE". IT WAS NOT REPORTED IF THERE WAS ANY PATIENT INVOLVEMENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? DID THIS ISSUE OCCUR IN ONE PATIENT PROCEDURE? INVESTIGATION SUMMARY: ONE PICTURE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF ONE PICTURE, A SEALED BOX OF PRODUCT CODE V372H WITH LOT PHBBCJQ0 AND A PRINTED THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE TO ADDRESS THE EVENT REPORTED UNDER (B)(4). NOTE: EVENTS REPORTED VIA MW # 2210968-2020-04870, 2210968-2020-04871, 2210968-2020-04872, 2210968-2020-04874, 2210968-2020-04875, 2210968-2020-04876
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201504&pc=GAM | 2210968-2020-04870 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE BOX OF THE PRODUCT HAD THE LABEL "NOT FOR HUMAN USE". A PROCEDURE WAS IN PROGRESS WHEN THE PROBLEM WAS DISCOVERED AND THE SURGEON IMMEDIATELY REMOVED THE IMPACTED SUTURE PRESENT IN THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? DID THIS ISSUE OCCUR IN ONE PATIENT PROCEDURE? INVESTIGATION SUMMARY: ONE PICTURE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF ONE PICTURE, A SEALED BOX OF PRODUCT CODE V372H WITH LOT PHBBCJQ0 AND A PRINTED THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE TO ADDRESS THE EVENT REPORTED UNDER (B)(4).. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: EVENTS REPORTED VIA MW # 2210968-2020-04870, 2210968-2020-04871, 2210968-2020-04872, 2210968-2020-04873, 2210968-2020-04874, 2210968-2020-04875, 2210968-2020-04876.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201516&pc=GAM | 2210968-2020-04875 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE BOX OF THE SUTURE PRODUCT HAD THE LABEL "NOT FOR HUMAN USE". IT WAS NOT REPORTED IF THERE WAS ANY PATIENT INVOLVEMENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? DID THIS ISSUE OCCUR IN ONE PATIENT PROCEDURE? INVESTIGATION SUMMARY: ONE PICTURE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF ONE PICTURE, A SEALED BOX OF PRODUCT CODE V372H WITH LOT PHBBCJQ0 AND A PRINTED THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE TO ADDRESS THE EVENT REPORTED UNDER (B)(4).. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: EVENTS REPORTED VIA MW # 2210968-2020-04870, 2210968-2020-04871, 2210968-2020-04872, 2210968-2020-04873, 2210968-2020-04874, 2210968-2020-04875, 2210968-2020-04876
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201520&pc=GAM | 2210968-2020-04874 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE BOX OF THE SUTURE PRODUCT HAD THE LABEL "NOT FOR HUMAN USE". IT WAS NOT REPORTED IF THERE WAS ANY PATIENT INVOLVEMENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? DID THIS ISSUE OCCUR IN ONE PATIENT PROCEDURE? INVESTIGATION SUMMARY: ONE PICTURE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF ONE PICTURE, A SEALED BOX OF PRODUCT CODE V372H WITH LOT PHBBCJQ0 AND A PRINTED THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE TO ADDRESS THE EVENT REPORTED UNDER (B)(4). . IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: EVENTS REPORTED VIA MW # 2210968-2020-04870, 2210968-2020-04871, 2210968-2020-04872, 2210968-2020-04873, 2210968-2020-04874, 2210968-2020-04875, 2210968-2020-04876
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201521&pc=GAM | 2210968-2020-04876 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE BOX OF THE SUTURE PRODUCT HAD THE LABEL "NOT FOR HUMAN USE". IT WAS NOT REPORTED IF THERE WAS ANY PATIENT INVOLVEMENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? DID THIS ISSUE OCCUR IN ONE PATIENT PROCEDURE? INVESTIGATION SUMMARY: ONE PICTURE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF ONE PICTURE, A SEALED BOX OF PRODUCT CODE V372H WITH LOT PHBBCJQ0 AND A PRINTED THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED NONCONFORMANCE TO ADDRESS THE EVENT REPORTED UNDER (B)(4).. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: EVENTS REPORTED VIA MW # 2210968-2020-04870, 2210968-2020-04871, 2210968-2020-04872, 2210968-2020-04873, 2210968-2020-04874, 2210968-2020-04875, 2210968-2020-04876
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201539&pc=GAM | 2210968-2020-04855 | 2019-07-22 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE CLIENT RECEIVED A BOX OF SUTURE WITH THE NOT FOR HUMAN USE PRINTED ON THE BOX. THE BOX HAS AN ANNOTATION PRINTED NEXT TO THE DATAMATRIX THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: ONLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, AN UNOPENED BOX OF PRODUCT CODE V372H WITH LOT# PHBBCJQ0 AND PRINTED AT THE SECONDARY PACKAGING THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA: 2210968-2020-04855, 2210968-2020-04859, 2210968-2020-04861, 2210968-2020-04867, 2210968-2020-04869.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201558&pc=GAM | 2210968-2020-04856 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866 AND 2210968-2020-04868. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841 . IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201607&pc=GAM | 2210968-2020-04859 | 2019-07-22 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE CLIENT RECEIVED A BOX OF SUTURE WITH THE NOT FOR HUMAN USE PRINTED ON THE BOX. THE BOX HAS AN ANNOTATION PRINTED NEXT TO THE DATAMATRIX. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: ONLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, AN UNOPENED BOX OF PRODUCT CODE V372H WITH LOT# PHBBCJQ0 AND PRINTED AT THE SECONDARY PACKAGING THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA: 2210968-2020-04855, 2210968-2020-04859, 2210968-2020-04861, 2210968-2020-04867, 2210968-2020-04869.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201739&pc=GAM | 2210968-2020-04857 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866 AND 2210968-2020-04868. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201756&pc=GAM | 2210968-2020-04858 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866 AND 2210968-2020-04868. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10201841&pc=GAM | 2210968-2020-04860 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866 AND 2210968-2020-04868. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841 . IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10202143&pc=GAM | 2210968-2020-04862 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866 AND 2210968-2020-04868. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10202285&pc=GAM | 2210968-2020-04863 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866 AND 2210968-2020-04868. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10202324&pc=GAM | 2210968-2020-04864 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866 AND 2210968-2020-04868. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10202409&pc=GAM | 2210968-2020-04865 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866 AND 2210968-2020-04868. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10202496&pc=GAM | 2210968-2020-04866 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SUMMARY: SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866 AND 2210968-2020-04868. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10202532&pc=GAM | 2210968-2020-04861 | 2019-07-22 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE CLIENT RECEIVED A BOX OF SUTURE WITH THE NOT FOR HUMAN USE PRINTED ON THE BOX. THE BOX HAS AN ANNOTATION PRINTED NEXT TO THE DATAMATRIX. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: ONLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, AN UNOPENED BOX OF PRODUCT CODE V372H WITH LOT# PHBBCJQ0 AND PRINTED AT THE SECONDARY PACKAGING THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA: 2210968-2020-04855, 2210968-2020-04859, 2210968-2020-04861, 2210968-2020-04867, 2210968-2020-04869.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10202557&pc=GAM | 2210968-2020-04868 | 2019-08-27 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: DURING THE EVALUATION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE SUTURE PRODUCT WAS OBSERVED. THERE WAS NO PATIENT INVOLVEMENT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. SEALED BOX OF PRODUCT CODE V372, LOT PHBBCJQ0 RECEIVED FOR ANALYSIS. DURING THE VISUAL INSPECTION, A LEGEND OF "NOT FOR HUMAN USE" AT THE ARTWORK OF THE PRODUCT COULD BE OBSERVED, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. HOWEVER, THE SAMPLES AT THE INSIDE WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO THE REQUIREMENTS OF THE FINISHED GOODS PRODUCT. 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04856, 2210968-2020-04857, 2210968-2020-04858, 2210968-2020-04860, 2210968-2020-04862, 2210968-2020-04863, 2210968-2020-04864, 2210968-2020-04865, 2210968-2020-04866. AND ADDITIONAL 10 EVENTS WERE SUBMITTED VIA 2210968-2020-04831, 2210968-2019-89974, 2210968-2020-04833, 2210968-2020-04835, 2210968-2020-04836, 2210968-2020-04837, 2210968-2020-04838, 2210968-2020-04839, 2210968-2020-04840 AND 2210968-2020-04841
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10202691&pc=GAM | 2210968-2020-04867 | 2019-07-22 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE CLIENT RECEIVED A BOX OF SUTURE WITH THE NOT FOR HUMAN USE PRINTED ON THE BOX. THE BOX HAS AN ANNOTATION PRINTED NEXT TO THE DATAMATRIX. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: ONLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, AN UNOPENED BOX OF PRODUCT CODE V372H WITH LOT# PHBBCJQ0 AND PRINTED AT THE SECONDARY PACKAGING THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA: 2210968-2020-04855, 2210968-2020-04859, 2210968-2020-04861, 2210968-2020-04867, 2210968-2020-04869.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10203032&pc=GAM | 2210968-2020-04869 | 2019-07-22 | Malfunction | ETHICON INC. | 2020-06-26 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE CLIENT RECEIVED A BOX OF SUTURE WITH THE NOT FOR HUMAN USE PRINTED ON THE BOX. THE BOX HAS AN ANNOTATION PRINTED NEXT TO THE DATAMATRIX. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: ONLY A PICTURE OF THE SAMPLE WAS RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, AN UNOPENED BOX OF PRODUCT CODE V372H WITH LOT# PHBBCJQ0 AND PRINTED AT THE SECONDARY PACKAGING THE LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. EVENTS WERE SUBMITTED VIA: 2210968-2020-04855, 2210968-2020-04859, 2210968-2020-04861, 2210968-2020-04867, 2210968-2020-04869.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197991&pc=GAM | 2210968-2020-04824 |  | Malfunction | ETHICON INC. | 2020-06-25 | GAM | VICRYL CT VIO 90CM M5 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY IN 2019 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE CUSTOMER RECEIVED A BOX OF THE PRODUCT WITH THE LABEL "NOT FOR HUMAN USE" ON IT. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198266&pc=GAM | 2210968-2020-04834 |  | Malfunction | ETHICON INC. | 2020-06-25 | GAM | VICRYL VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT UPON DELIVERY THE CUSTOMER SAW THAT ON THE BOX WAS NOTED "NOT FOR HUMAN USE." HOWEVER THE CORRECT ITEM NUMBER WAS STATED ON THE PACKAGE. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: A PHOTO WAS RECEIVED FOR EVALUATION. UPON VISUAL INSPECTION BOXES OF PRODUCT CODE V371, LOT # PHBBJHR0 WITH A LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS WERE SUBMITTED VIA: 2210968-2020-04826, 2210968-2020-04827, 2210968-2020-04828, 2210968-2020-04829, 2210968-2020-04830, 2210968-2020-04832, 2210968-2020-04834. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198272&pc=GAM | 2210968-2020-04826 |  | Malfunction | ETHICON INC. | 2020-06-25 | GAM | VICRYL VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT UPON DELIVERY THE CUSTOMER SAW THAT ON THE BOX WAS NOTED "NOT FOR HUMAN USE." HOWEVER THE CORRECT ITEM NUMBER WAS STATED ON THE PACKAGE. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: A PHOTO WAS RECEIVED FOR EVALUATION. UPON VISUAL INSPECTION BOXES OF PRODUCT CODE V371, LOT # PHBBJHR0 WITH A LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS WERE SUBMITTED VIA: 2210968-2020-04826, 2210968-2020-04827, 2210968-2020-04828, 2210968-2020-04829, 2210968-2020-04830, 2210968-2020-04832, 2210968-2020-04834. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198282&pc=GAM | 2210968-2020-04829 |  | Malfunction | ETHICON INC. | 2020-06-25 | GAM | VICRYL VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT UPON DELIVERY THE CUSTOMER SAW THAT ON THE BOX WAS NOTED "NOT FOR HUMAN USE." HOWEVER THE CORRECT ITEM NUMBER WAS STATED ON THE PACKAGE. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: A PHOTO WAS RECEIVED FOR EVALUATION. UPON VISUAL INSPECTION BOXES OF PRODUCT CODE V371, LOT # PHBBJHR0 WITH A LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS WERE SUBMITTED VIA: 2210968-2020-04826, 2210968-2020-04827, 2210968-2020-04828, 2210968-2020-04829, 2210968-2020-04830, 2210968-2020-04832, 2210968-2020-04834. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198287&pc=GAM | 2210968-2020-04827 |  | Malfunction | ETHICON INC. | 2020-06-25 | GAM | VICRYL VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT UPON DELIVERY THE CUSTOMER SAW THAT ON THE BOX WAS NOTED "NOT FOR HUMAN USE." HOWEVER THE CORRECT ITEM NUMBER WAS STATED ON THE PACKAGE. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: A PHOTO WAS RECEIVED FOR EVALUATION. UPON VISUAL INSPECTION BOXES OF PRODUCT CODE V371, LOT # PHBBJHR0 WITH A LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS WERE SUBMITTED VIA: 2210968-2020-04826, 2210968-2020-04827, 2210968-2020-04828, 2210968-2020-04829, 2210968-2020-04830, 2210968-2020-04832, 2210968-2020-04834. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10198291&pc=GAM | 2210968-2020-04830 |  | Malfunction | ETHICON INC. | 2020-06-25 | GAM | VICRYL VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT UPON DELIVERY THE CUSTOMER SAW THAT ON THE BOX WAS NOTED "NOT FOR HUMAN USE." HOWEVER THE CORRECT ITEM NUMBER WAS STATED ON THE PACKAGE. THERE WAS NO PATIENT INVOLVEMENT REPORTED. Manufacturer Narrative: (B)(4). EVALUATION SUMMARY: A PHOTO WAS RECEIVED FOR EVALUATION. UPON VISUAL INSPECTION BOXES OF PRODUCT CODE V371, LOT # PHBBJHR0 WITH A LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED. THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE RELATED TO THIS ISSUE. THE NECESSARY ACTIONS HAVE BEEN TAKEN AND DOCUMENTED IN THE APPROPRIATE QUALITY SYSTEM. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS WERE SUBMITTED VIA: 2210968-2020-04826, 2210968-2020-04827, 2210968-2020-04828, 2210968-2020-04829, 2210968-2020-04830, 2210968-2020-04832, 2210968-2020-04834. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10185529&pc=GAM | 2210968-2020-04697 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES.¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04699, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04705, 2210968-2020-04774.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10185569&pc=GAM | 2210968-2020-04700 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES.¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04699, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04705, 2210968-2020-04774.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10185614&pc=GAM | 2210968-2020-04701 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES.¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, , 2210968-2020-04700, 2210968-2020-04699, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04705, 2210968-2020-04774.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10185650&pc=GAM | 2210968-2020-04702 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04699, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04705, 2210968-2020-04774. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10185721&pc=GAM | 2210968-2020-04699 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES.¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04705, 2210968-2020-04774.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10185992&pc=GAM | 2210968-2020-04774 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04705, 2210968-2020-04699. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10186822&pc=GAM | 2210968-2020-04707 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04706, 2210968-2020-04699, 2210968-2020-04774, 2210968-2020-04705.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10186823&pc=GAM | 2210968-2020-04706 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER (B)(4) AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04699, 2210968-2020-04774, 2210968-2020-04705. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10186824&pc=GAM | 2210968-2020-04708 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES.¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04699, 2210968-2020-04774, 2210968-2020-04705.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10186825&pc=GAM | 2210968-2020-04710 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04699, 2210968-2020-04774, 2210968-2020-04705.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10186826&pc=GAM | 2210968-2020-04773 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER (B)(4) AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04699, 2210968-2020-04774, 2210968-2020-04705. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10186829&pc=GAM | 2210968-2020-04772 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER (B)(4), AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04699, 2210968-2020-04774, 2210968-2020-04705 IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10186834&pc=GAM | 2210968-2020-04711 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES.¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04699, 2210968-2020-04774, 2210968-2020-04705.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10186837&pc=GAM | 2210968-2020-04705 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER (B)(4) AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04709, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04699, 2210968-2020-04774. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10186841&pc=GAM | 2210968-2020-04709 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-23 | GAM | VICRYL VIO 75CM M2 | Break; Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT ANIMALS UNDERWENT AN OVARIECTOMY ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AND THE NEEDLE PULLED OFF. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER JV452; QABBTLR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IT WAS REPORTED THAT ¿A TOTAL OF 15 VICRYL JV452 FROM THE SAME LOT NUMBER WERE DEFECTIVE DURING SEVERAL SURGERIES.¿ HOW MANY PROCEDURES THAT HAD ¿PULLED OFF AND BREAKAGE OF THE WIRES¿ ISSUES? WERE THESE CASES PREVIOUSLY REPORTED TO ETHICON INDIVIDUALLY? IF YES, PLEASE PROVIDE THE FILES NUMBERS. THE EXACT NUMBER OF PROCEDURES IS UNKNOWN. NOTE: EVENTS REPORTED IN: 2210968-2020-04702, 2210968-2020-04701, 2210968-2020-04700, 2210968-2020-04711, 2210968-2020-04772, 2210968-2020-04710, 2210968-2020-04773. 2210968-2020-04708, 2210968-2020-04707, 2210968-2020-04706, 2210968-2020-04699, 2210968-2020-04774, 2210968-2020-04705.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10182438&pc=GAM | 2210968-2020-04730 | 2020-06-10 | Malfunction | ETHICON INC. | 2020-06-22 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAP COLON ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGEON USED A SUTURE FOR HEMOSTASIS. AFTER THE FIRST KNOT, IN TIGHTENING THE SECOND, THERE WAS AN ACCIDENTAL DETACHMENT OF THE NEEDLE FROM THE THREAD. A DIFFERENT TYPE SUTURE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. THE SURGERY WAS DELAYED 10 MINUTES. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE (B)(4) NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10173527&pc=GAM | 2210968-2020-04649 |  | Malfunction | ETHICON INC. | 2020-06-19 | GAM | VICRYL RAP UND 45CM 3-0 PS-22 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE THE SUTURE SNAPPED. A LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THE PROCEDURE HAD A FEW MINUTES DELAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: HOW LONG WAS THE DELAY? MINUTES, EXACT UNKNOWN. WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME? NO. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN, NO. PATIENT¿S CURRENT STATUS? UNKNOWN. NAME AND DATE OF PROCEDURE, UNKNOWN. PLEASE ALSO PROVIDE US WITH AN UPDATE WITH REGARDS TO THE PRODUCT RETURN. HAVE YOU ALREADY RETURNED THE PRODUCT FOR ANALYSIS? COLLECTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10149808&pc=GAM | 2210968-2020-04550 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-12 | GAM | VICRYL CT VIO 70CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE IN MAY 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE (B)(4); NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS REPORTED: "NEEDLE BREAK" PHOTO NA1 SHOWS A PICTURE OF A SUTURE BREAKAGE PLEASE CLARIFY:IS IT SUTURE BREAKAGE OR NEEDLE BREAKAGE? EVENT DAY? PROCEDURE DATE? PROCEDURE NAME? WHAT WAS USED TO COMPLETE THE PROCEDURE? WERE THERE ANY ADVERSE PATIENT CONSEQUENCES? INVESTIGATION SUMMARY: THREE PICTURES WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE SIDE OF THE BOX, ONE UNOPENED FOIL AND A NEEDLE-SUTURE PIECE OF PRODUCT CODE (B)(4), LOT PGBDCWQ0 WAS NOTED. NO NEEDLE BREAKAGE COULD BE OBSERVED. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10146792&pc=GAM | 2210968-2020-04542 |  | Malfunction | ETHICON INC. | 2020-06-11 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC SURGERY IN 2020 AND THE SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141279&pc=GAM | 2210968-2020-04505 |  | Malfunction | ETHICON INC. | 2020-06-10 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC SURGERY IN 2020 AND THE SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141305&pc=GAM | 2210968-2020-04506 |  | Malfunction | ETHICON INC. | 2020-06-10 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC SURGERY IN 2020 AND THE SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141313&pc=GAM | 2210968-2020-04507 |  | Malfunction | ETHICON INC. | 2020-06-10 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141331&pc=GAM | 2210968-2020-04508 |  | Malfunction | ETHICON INC. | 2020-06-10 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC SURGERY IN 2020 AND THE SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141353&pc=GAM | 2210968-2020-04509 |  | Malfunction | ETHICON INC. | 2020-06-10 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC SURGERY IN 2020 AND THE SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141359&pc=GAM | 2210968-2020-04510 |  | Malfunction | ETHICON INC. | 2020-06-10 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC SURGERY IN 2020 AND THE SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141368&pc=GAM | 2210968-2020-04511 |  | Malfunction | ETHICON INC. | 2020-06-10 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC SURGERY IN 2020 AND THE SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141376&pc=GAM | 2210968-2020-04512 |  | Malfunction | ETHICON INC. | 2020-06-10 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC SURGERY IN 2020 AND THE SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10134700&pc=GAM | 2210968-2020-04448 | 2020-05-01 | Malfunction | ETHICON INC. | 2020-06-09 | GAM | VICRYL VIO 70CM M3.5::USP0 SGLE ARMED CT-1 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A UNDERWENT A CESAREAN SECTION IN (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, AS THE SURGEON WAS CLOSING THE UTERUS, THE NEEDLE SEPARATED FROM THE SUTURE, LODGING THE NEEDLE INTO THE UTERUS. THE NEEDLE WAS ABLE TO BE REMOVED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS MEDWATCH REPORT IS IN RESPONSE TO RECEIPT OF FACILITY REPORT # 3602300000-2020-8001.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135350&pc=GAM | 2210968-2020-04456 | 2020-06-01 | Malfunction | ETHICON INC. | 2020-06-09 | GAM | VICRYL VIO 70CM M2 | Material Separation | No Consequences Or Impact To Patient; Not Applicable; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT LAPAROSCOPIC PROCEDURE GASTRIC BYPASS PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. WHILE TIGHTENING THE KNOT, THE NEEDLE CAME OFF, SO THE SURGEON HAD TO USE A NEW SUTURE TO CONTINUE THE PROCEDURE. THE PROCEDURE WAS DELAYED 10 MINUTES. ANOTHER SUTURE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10128261&pc=GAM | 2210968-2020-04357 | 2020-01-22 | Malfunction | ETHICON INC. | 2020-06-08 | GAM | VICRYL VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A C SECTION ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SURGEON NOTICED THE TIP OF THE NEEDLE WAS BROKEN OFF. DURING THE X-RAY, THE NEEDLE WAS NOT FOUND IN THE BODY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE LOT NUMBER? DO YOU SEARCH THE TIP NEEDLE IN THE PACKAGE OR IF IT FALL IN THE FLOOR? PLEASE CONFIRM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10115200&pc=GAM | 2210968-2020-04293 | 2020-05-27 | Malfunction | ETHICON INC. | 2020-06-03 | GAM | VICRYL RAP UND 45CM 4/0 FS-2S | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN EAR, NOSE AND THROAT PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE SEPARATED FROM SUTURE AFTER THE FIRST STITCH. THE PROCEDURE WAS COMPLETED WITH SAME LIKE PRODUCT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V2920H;MJ8CJMT0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY PATIENT CONSEQUENCE OR INJURY DUE TO ISSUE? NO WHAT IS THE CONDITION OF THE PATIENT? NO INFORMATION WAS MADE AVAILABLE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105249&pc=GAM | 2210968-2020-04278 | 2020-05-21 | Malfunction | ETHICON INC. | 2020-06-01 | GAM | VICRYL*RAPID UND 2-0 90CM 1CT1 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10099814&pc=GAM | 2210968-2020-04253 | 2020-05-19 | Malfunction | ETHICON INC. | 2020-05-29 | GAM | VICRYL VIO 70CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE BROKE INTO TWO PARTS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? UNKNOWN, WAS PROCEDURE SUCCESSFULLY COMPLETED? UNKNOWN, WERE FRAGMENTS GENERATED? UNKNOWN, IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES- UNKNOWN, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IS THE NEEDLE PIECE RETAINED IN THE PATIENT? IF YES, PLEASE EXPLAIN. HOW WAS THE NEEDLE RETRIEVED FROM THE PATIENT? WAS THERE ANY ADVERSE CONSEQUENCE ASSOCIATED WITH THE PATIENT? WAS PROCEDURE SUCCESSFULLY COMPLETED? PLEASE PROVIDE LOT NUMBER. NAME OF THE PROCEDURE? PLEASE PROVIDE PROCEDURE DATE. PLEASE PROVIDE EVENT DATE. STATUS OF DEVICE RETURN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10099833&pc=GEA | 2210968-2020-04255 |  | Malfunction | ETHICON INC. | 2020-05-29 | GEA | VICRYL*VLT 0 45CM - ENDOLOOP | Break | No Consequences Or Impact To Patient | K925914 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THEY OPENED THE SUTURE AND IT BROKE VERY EASILY. THEY WERE PREVIOUSLY USING THIS SUTURE AND THIS IS THE FIRST TIME THAT IT FAILED. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE INITIAL PROCEDURE? WHAT IS THE PROCEDURE DATE AND EVENT DAY? COULD YOU PLEASE CONFIRM WHAT WAS BROKE? SUTURE OR NEEDLE? ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE DEVICE USED IN A SURGICAL PROCEDURE? (NO, YES AND WHICH) IF YES: DATE OF THE SURGICAL PROCEDURE. PATIENT DEMOGRAPHICS (INITIALS STARTING WITH LAST NAME, AGE, GENDER). WHAT ACTION DID THE SURGEON TAKE TO CORRECT / CORRECT THE PROBLEM PRESENTED WITH THE MEDICAL DEVICE? WAS ANOTHER MEDICAL INTERVENTION REQUIRED DUE TO THE PROBLEM PRESENTED WITH THE DEVICE? (EG, X-RAYS, ADDITIONAL PROCEDURES, ADDITIONAL CLINICAL TREATMENT, ADDITIONAL MEDICATIONS, ADDITIONAL REVIEW). WAS THERE DAMAGE TO THE PATIENT DUE TO THE PROBLEM PRESENTED WITH THE DEVICE? (NO; YES AND DESCRIBE IN DETAIL FOR EXAMPLE DEATH, PERMANENT DEFICIENCY OF A BODILY FUNCTION, PERMANENT DAMAGE TO A BODILY STRUCTURE, PROLONGED HOSPITALIZATION, BLEEDING REQUIRING TRANSFUSION, PHYSICAL INJURY, ETC.). WHAT IS THE CURRENT CONDITION OF THE PATIENT? (FOR EXAMPLE: DISCHARGED WITHOUT CONSEQUENCES DERIVED FROM THE PROBLEM WITH THE DEVICE, IN RECOVERY FROM SURGERY, CONTINUES HOSPITALIZED DUE TO ... ETC.). INITIALS OF THE SURGEON WHO OPERATED THE MEDICAL DEVICE (STARTING WITH LAST NAME) FULL NAME AND ADDRESS OF THE REPORTING HOSPITAL / CLINIC. WILL THE PRODUCT BE ABLE TO BE RECOVERED FOR ANALYSIS? (NO AND REASON; YES AND STATUS OF RETURN).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10099834&pc=GEA | 2210968-2020-04254 |  | Malfunction | ETHICON INC. | 2020-05-29 | GEA | VICRYL*VLT 0 45CM - ENDOLOOP | Break | No Consequences Or Impact To Patient | K925914 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THEY OPENED THE SUTURE AND IT BROKE VERY EASILY. THEY WERE PREVIOUSLY USING THIS SUTURE AND THIS IS THE FIRST TIME THAT IT FAILED. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE INITIAL PROCEDURE? WHAT IS THE PROCEDURE DATE AND EVENT DAY? COULD YOU PLEASE CONFIRM WHAT WAS BROKE? SUTURE OR NEEDLE? ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WAS THE DEVICE USED IN A SURGICAL PROCEDURE? (NO, YES AND WHICH) IF YES: DATE OF THE SURGICAL PROCEDURE. PATIENT DEMOGRAPHICS (INITIALS STARTING WITH LAST NAME, AGE, GENDER) WHAT ACTION DID THE SURGEON TAKE TO CORRECT / CORRECT THE PROBLEM PRESENTED WITH THE MEDICAL DEVICE? WAS ANOTHER MEDICAL INTERVENTION REQUIRED DUE TO THE PROBLEM PRESENTED WITH THE DEVICE? (EG, X-RAYS, ADDITIONAL PROCEDURES, ADDITIONAL CLINICAL TREATMENT, ADDITIONAL MEDICATIONS, ADDITIONAL REVIEW). WAS THERE DAMAGE TO THE PATIENT DUE TO THE PROBLEM PRESENTED WITH THE DEVICE? (NO; YES AND DESCRIBE IN DETAIL FOR EXAMPLE DEATH, PERMANENT DEFICIENCY OF A BODILY FUNCTION, PERMANENT DAMAGE TO A BODILY STRUCTURE, PROLONGED HOSPITALIZATION, BLEEDING REQUIRING TRANSFUSION, PHYSICAL INJURY, ETC.). WHAT IS THE CURRENT CONDITION OF THE PATIENT? (FOR EXAMPLE: DISCHARGED WITHOUT CONSEQUENCES DERIVED FROM THE PROBLEM WITH THE DEVICE, IN RECOVERY FROM SURGERY, CONTINUES HOSPITALIZED DUE TO ... ETC.). INITIALS OF THE SURGEON WHO OPERATED THE MEDICAL DEVICE (STARTING WITH LAST NAME) FULL NAME AND ADDRESS OF THE REPORTING HOSPITAL / CLINIC. WILL THE PRODUCT BE ABLE TO BE RECOVERED FOR ANALYSIS? (NO AND REASON; YES AND STATUS OF RETURN).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10095360&pc=GAM | 2210968-2020-04207 | 2020-05-19 | Malfunction | ETHICON INC. | 2020-05-28 | GAM | VICRYL*RAPIDE*UND 0 90CM CT | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CHILDBIRTH PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. DURING CHILDBIRTH, THE SUTURE WAS DETACHED FROM THE NEEDLE DURING CONTINUOUS SUTURING ON THE PERINEUM. IT WAS NOT CONTROL RELEASE NEEDLE. THIS ISSUE OCCURRED SEVERAL TIMES IN A ROW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS WERE SUBMITTED VIA 2210968-2020-04207 AND 2210968-2020-04208.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10095369&pc=GAM | 2210968-2020-04208 | 2020-05-19 | Malfunction | ETHICON INC. | 2020-05-28 | GAM | VICRYL*RAPIDE*UND 0 90CM CT | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CHILDBIRTH PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. DURING CHILDBIRTH, THE SUTURE WAS DETACHED FROM THE NEEDLE DURING CONTINUOUS SUTURING ON THE PERINEUM. IT WAS NOT CONTROL RELEASE NEEDLE. THIS ISSUE OCCURRED SEVERAL TIMES IN A ROW. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS WERE SUBMITTED VIA 2210968-2020-04207 AND 2210968-2020-04208.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10096234&pc=GAM | 2210968-2020-04157 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-05-28 | GAM | VICRYL RAP UND90CM 2/0 CT-1 PLUS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE ON AN UNKNOWN DATE AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PCFX1 ADDITIONAL INFORMATION REQUESTED: PROCEDURE DATE, LOT NUMBER.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10099645&pc=GAM | 2210968-2020-04251 |  | Malfunction | ETHICON INC. | 2020-05-28 | GAM | VICRYL RAPIDE VIO 30CM M0.4 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BROKEN DURING USE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE ALL 4 SUTURES INVOLVED IN THIS ONE PATIENT SINGLE PROCEDURE? PLEASE PROVIDE LOT NUMBER? PROCEDURE/EVENT DATE? DEVICE RETURN STATUS/FOLLOW UP EVENTS WERE SUBMITTED VIA 2210968-2020-04248, 2210968-2020-04249, 2210968-2020-04250 AND 2210968-2020-04251.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10099646&pc=GAM | 2210968-2020-04250 |  | Malfunction | ETHICON INC. | 2020-05-28 | GAM | VICRYL RAPIDE VIO 30CM M0.4 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BROKEN DURING USE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE ALL 4 SUTURES INVOLVED IN THIS ONE PATIENT SINGLE PROCEDURE? PLEASE PROVIDE LOT NUMBER? PROCEDURE/EVENT DATE? DEVICE RETURN STATUS/FOLLOW UP. EVENTS WERE SUBMITTED VIA 2210968-2020-04248, 2210968-2020-04249, 2210968-2020-04250 AND 2210968-2020-04251.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10099648&pc=GAM | 2210968-2020-04249 |  | Malfunction | ETHICON INC. | 2020-05-28 | GAM | VICRYL RAPIDE VIO 30CM M0.4 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BROKEN DURING USE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE ALL 4 SUTURES INVOLVED IN THIS ONE PATIENT SINGLE PROCEDURE? PLEASE PROVIDE LOT NUMBER? PROCEDURE/EVENT DATE? DEVICE RETURN STATUS/FOLLOW UP. EVENTS WERE SUBMITTED VIA 2210968-2020-04248, 2210968-2020-04249, 2210968-2020-04250 AND 2210968-2020-04251.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10099655&pc=GAM | 2210968-2020-04248 |  | Malfunction | ETHICON INC. | 2020-05-28 | GAM | VICRYL RAPIDE VIO 30CM M0.4 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BROKEN DURING USE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE ALL 4 SUTURES INVOLVED IN THIS ONE PATIENT SINGLE PROCEDURE? PLEASE PROVIDE LOT NUMBER? PROCEDURE/EVENT DATE? DEVICE RETURN STATUS/FOLLOW UP. EVENTS WERE SUBMITTED VIA 2210968-2020-04248, 2210968-2020-04249, 2210968-2020-04250 AND 2210968-2020-04251.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10088043&pc=GAM | 2210968-2020-04109 | 2020-05-14 | Malfunction | ETHICON INC. | 2020-05-26 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A SPINE SURGERY ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE SURGERY THE SUTURE BROKE OFF. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THE DEVICE HISTORY RECORDS WERE REVIEWED AND THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE PROCEDURE DATE? IS THERE A PHOTO AVAILABLE, THAT YOU CAN PROVIDE US? WHAT WAS USED TO COMPLETE THE SURGERY? WAS THE SURGERY COMPLETED SUCCESSFULLY?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10088251&pc=GAM | 2210968-2020-04110 | 2020-04-30 | Malfunction | ETHICON INC. | 2020-05-26 | GAM | VICRYL*VLT PLUS 0 70CM 1-CT | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE WITH LITTLE PRESSURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE INITIAL PROCEDURE? WHAT WAS USED TO COMPLETE THE PROCEDURE? WAS THE SURGERY COMPLETED SUCCESSFULLY? WAS THERE ANY PATIENT CONSEQUENCE? TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10073169&pc=GAM | 2210968-2020-04028 | 2020-04-09 | Malfunction | ETHICON INC. | 2020-05-19 | GAM | VICRYL RAP UND 75CM M2 USP3/0 SGLE ARMED PSLX PRIME | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A REMOVAL OF INTERNAL FIXATION SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS ALSO REPORTED THAT THE SUTURE BROKE WHEN DRAWING THE SUTURE AFTER THE NEEDLE HOLD CLAMPED THE NEEDLE. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE SURGERY. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10073425&pc=FTL | 2210968-2020-04034 | 2020-04-13 | Malfunction | ETHICON INC. | 2020-05-19 | FTL | VICRYL MESH 26.5X21.5CM | Packaging Problem | No Consequences Or Impact To Patient | K810428 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND THE MESH WAS USED. IT WAS REPORTED THAT THERE WAS A WELDING DEFECT OF THE DEVICE SO IT WAS IMPOSSIBLE TO USE IT BECAUSE THE PRODUCT WAS UNSTERILIZED. IT WAS REPORTED THAT THE INDIVIDUAL STERILITY PROTECTOR IS POORLY SEALED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS MEANT BY ¿THE PRODUCT IS UNSTERILIZED?¿ WAS THE PACKAGING COMPROMISED? WHAT IS MEANT BY THE ¿WELDING DEFECT?¿ ARE THERE ANY PHOTOS OF THE PRODUCT INVOLVED? WAS THE PROCEDURE COMPLETED SUCCESSFULLY WITH ANOTHER PRODUCT? THE INDIVIDUAL STERILITY PROTECTOR IS POORLY SEALED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10065535&pc=GAM | 2210968-2020-03979 | 2020-04-22 | Malfunction | ETHICON INC. | 2020-05-18 | GAM | VICRYL RAP UND 75CM M2 USP3/0 SGLE ARMED PSLX PRIME | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE WHEN DRAWING THE SUTURE AFTER CLAMPING THE NEEDLE. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10059881&pc=GAM | 2210968-2020-03818 | 2020-04-20 | Malfunction | ETHICON INC. | 2020-05-14 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | Foreign Body In Patient; Not Applicable | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NATURAL BIRTH SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. DURING THE SURGERY, THE NEEDLE BROKE WHEN SEWING THE PERINEAL WOUND. ABOUT 1.5CM OF NEEDLE REMAINED IN MUSCULARIS PERINEUM OF THE PATIENT. THE REMAINING PART WAS NOT FOUND IMMEDIATELY. THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL AND THE REMAINING NEEDLE WAS RETRIEVED AT LAST. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE SURGERY. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DATE OF NATURAL BIRTH SURGERY? WHAT WAS THE CONDITION OF THE TISSUE (HEALTHY, DISEASED) SUTURED? WHAT INSTRUMENTS WERE USED TO GRASP THE NEEDLE? WHERE WAS THE NEEDLE GRASPED DURING USE? WERE X-RAYS PERFORMED TO LOCALIZE THE NEEDLE FRAGMENT? WHEN WAS THE SURGERY FOR REMOVAL OF THE REMAINING NEEDLE PERFORMED? WERE ALL PIECES OF BROKEN NEEDLE RETRIEVED? IF APPLICABLE, WILL PRODUCT BE RETURNED, RETURN DATE, TRACKING INFORMATION? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? WHAT IS THE CURRENT CONDITION OF THE PATIENT? THE PATIENT DEMOGRAPHIC INFO: AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE AND OTHER RELEVANT PATIENT HISTORY?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10054535&pc=GAM | 2210968-2020-03916 | 2020-04-28 | Malfunction | ETHICON INC. | 2020-05-13 | GAM | VICRYL* VLT PLUS 2-0 90CM 1CT1 | Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. PRIOR TO USE ON A PATIENT, IT WAS NOTED THAT THE PRODUCT HAD THE APPEARANCE OF HIGH TEMPERATURE DETERIORATION, A POROUS PLASTIC PART, AND DISCOLORATION OF THE FILAMENT. IT WAS STATED THAT THE EVENT OCCURRED DURING STERILE PROCESSING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AM3603 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE PHOTO UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PHOTO EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IN EVENT DESCRIPTION MENTION THAT THE EVENT OCCURRED DURING STERILE PROCESSING, COULD YOU PLEASE CONFIRM IF THE DEVICE WAS EXPOSED TO HIGH TEMPERATURES? THE SUTURES ARE STORED IN A SPECIAL PLACE WITH A TEMPERATURE NOT HIGHER THAN 21 °". ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHERE WAS THE ISSUE FOUND (I.E., AT A HOSPITAL, AT A DISTRIBUTION WAREHOUSE, ETC)? WAS THE ISSUE NOTICED PRIOR TO USE ON A PATIENT? WAS THE DEVICE RE-STERILIZED PRIOR TO USE? DEVICE RETURN FOLLOW UP.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10047659&pc=GAM | 2210968-2020-03863 | 2020-02-11 | Malfunction | ETHICON INC. | 2020-05-12 | GAM | VICRYL CT VIO 90CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE WITHOUT ANY SPECIFIC STRESS IN THE PATIENT'S ABDOMEN. IT WAS DIFFICULT TO RETRIEVE IT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT MLBDSXQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE INITIAL PROCEDURE? DID X-RAY TAKEN? WHAT IS THE CURRENT CONDITION OF THE PATIENT? COULD YOU PLEASE CONFIRM THE AVAILABILITY OF THE DEVICE? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10016088&pc=GAM | 2210968-2020-03656 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-04-30 | GAM | VICRYL VIO 90CM 0 CTX | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE AND THREAD SEPARATED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE, THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10003484&pc=GAM | 2210968-2020-03551 |  | Malfunction | ETHICON INC. | 2020-04-27 | GAM | VICRYL RAP UND 75CM 3-0 PC-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON UNKNOWN DATE AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BREAKING AWAY FROM THE NEEDLE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHEN DID THE ISSUE OCCUR; IN THE PACKAGE/ DURING DISPENSING (BEFORE USE ON PATIENT/DURING ACTUAL SUTURING (USE ON PATIENT)? HOW MANY DEVICES WERE INVOLVED IN THIS SINGLE PROCEDURE? PLEASE CLARIFY DID THE ISSUE OCCURRED ON TWO SEPARATE PROCEDURES? PLEASE EXPLAIN NAME OF THE PROCEDURE? DATE THE EVENT OCCURRED. LOT NUMBER FOR BOTH PROCEDURES INVOLVED. WHAT WAS THE DATE OF THE INITIAL PROCEDURE? DEVICE RETURN STATUS/FOLLOW UP? THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10003485&pc=GAM | 2210968-2020-03550 |  | Malfunction | ETHICON INC. | 2020-04-27 | GAM | VICRYL RAP UND 75CM 3-0 PC-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON UNKNOWN DATE AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BREAKING AWAY FROM THE NEEDLE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT ATTEMPTS ARE BEING MADE TO CLARIFY THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHEN DID THE ISSUE OCCUR; IN THE PACKAGE/ DURING DISPENSING (BEFORE USE ON PATIENT/DURING ACTUAL SUTURING (USE ON PATIENT)? *HOW MANY DEVICES WERE INVOLVED IN THIS SINGLE PROCEDURE? PLEASE CLARIFY DID THE ISSUE OCCURRED ON TWO SEPARATE PROCEDURES? PLEASE EXPLAIN NAME OF THE PROCEDURE? DATE THE EVENT OCCURRED. LOT NUMBER FOR BOTH PROCEDURES INVOLVED. WHAT WAS THE DATE OF THE INITIAL PROCEDURE? DEVICE RETURN STATUS/FOLLOW UP? THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9990374&pc=GAM | 2210968-2020-03407 | 2020-02-02 | Malfunction | ETHICON INC. | 2020-04-23 | GAM | VICRYL CT VIO 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE NAME OF THE PROCEDURE? WHAT IS THE CURRENT CONDITION OF THE PATIENT? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9995140&pc=GAM | MW5094254 | 2020-04-16 | Malfunction | ETHICON, INC. | 2020-04-23 | GAM | VICRYL PLUS SUTURE NEEDLE | Break | No Known Impact Or Consequence To Patient |  |  | 1 | Event Description: DURING LAPAROSCOPIC PROCEDURE A FASCIA CLOSING SUTURE NEEDLE BROKE. THE BROKEN NEEDLE WAS RETRIEVED FROM THE PATIENT. X-RAY SHOWED NO RETAINED METALLIC FOREIGN BODY. FDA SAFETY REPORT ID# (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9982337&pc=GAM | 2210968-2020-03292 | 2020-03-26 | Malfunction | ETHICON INC. | 2020-04-21 | GAM | VICRYL RAPIDE 0 1X90 V-34 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF THE NEEDLE. THE SEAM WAS STOPPED AND THE THREAD CUT WAS REMOVED. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AM7562, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9982442&pc=GAM | 2210968-2020-03294 |  | Malfunction | ETHICON INC. | 2020-04-21 | GAM | VICRYL CT VIO 70CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NEUROSURGERY PROCEDURE ON AN UNKNOWN DATE; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT (B)(4) , AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IN EVENT DESCRIPTION INDICATE THERE IS NOT MUCH MORE INFORMATION BUT COULD YOU PLEASE PROVIDE THIS INFORMATION? WAS THE SURGERY DELAYED DUE TO THE ISSUE? WHAT IS THE CURRENT CONDITION OF THE PATIENT? WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? WAS THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? DOES A PIECE OF THE NEEDLE REMAIN IN THE PATIENT¿S TISSUE? WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? COULD YOU PLEASE CONFIRM THE AVAILABILITY OF THE NEEDLE? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983618&pc=GAM | 2210968-2020-03304 | 2020-01-23 | Malfunction | ETHICON INC. | 2020-04-21 | GAM | VICRYL CT VIO 75CM M4 | Break; Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE FIRST POINT THE NEEDLE WAS PASSED, BUT THE THREAD WAS BROKEN WHEN THE NEEDLE WAS PULLED. THE ISSUE OCCURRED TWICE DURING THE PROCEDURE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE NAME OF THE INITIAL PROCEDURE? COULD YOU PLEASE CONFIRM IF THE NEEDLE GOT SEPARATED FROM THE NEEDLE OR IF THE SUTURE BROKE? OR IF BOTH EVENTS OCCURRED? IT WAS REPORTED THAT THE ISSUE OCCURRED TWICE DURING THE PROCEDURE. DID THIS EVENT OCCUR WITH ONE OR TWO SUTURE/NEEDLE COMBINATIONS? HAS THE DEVICE BEEN SHIPPED FOR EVALUATION? IF SO, PLEASE PROVIDE ANY AVAILABLE TRACKING INFORMATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9984398&pc=GAM | 2210968-2020-03329 |  | Malfunction | ETHICON INC. | 2020-04-21 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Hematoma; Blood Loss | K022269 |  | 1 | Event Description: TITLE: EFFICACY AND SAFETY OF SUBCUTANEOUS FELLOW¿S STITCH USING ¿FISHERMAN¿S KNOT¿ TECHNIQUE TO ACHIEVE LARGE CALIBER (> 10 FRENCH) VENOUS HEMOSTASIS AMONG PATIENTS UNDERGOING INTERVENTION INVOLVING VENOUS ACCESS, VARIOUS TECHNIQUES HAVE BEEN IMPLEMENTED TO ACHIEVE HEMOSTASIS IN ORDER TO REDUCE LOCAL ACCESS SITE COMPLICATIONS, TO DECREASE LENGTH OF STAY AND TO FACILITATE EARLY AMBULATION. THE AUTHORS AIMED TO ASSESS THE EFFICACY AND SAFETY OF FELLOW¿S STITCH USING ¿FISHERMAN¿S KNOT¿ (FIGURE OF Z (FOZ)) TECHNIQUE WHEN COMPARED WITH CONVENTIONAL MANUAL COMPRESSION FOR IMMEDIATE CLOSURE OF LARGE VENOUS SHEATH. BETWEEN NOVEMBER 2012 AND MARCH 2019, A TOTAL OF 949 PATIENTS (MEAN AGE: 13.1 ± 8.2 YEARS; 574 MALE AND 375 FEMALE PATIENTS) UNDERWENT VARIOUS INTERVENTIONS WHICH INVOLVED VENOUS ACCESS REQUIRING HEMOSTASIS. IN GROUP I (THE FELLOW¿S STITCH GROUP), A VICRYL 0 SUTURE (ETHICON) ON A LARGE-CURVED CUTTING NEEDLE WAS PASSED LITTLE DISTAL (5 - 10 MM) TO ENTRY POINT OF SHEATH INTO SKIN BELOW THE SHEATH AND ADVANCED THROUGH THE SUBCUTANEOUS TISSUE HOLDING ITS GENEROUS AMOUNT BUT WITHOUT GOING TOO DEEP TO LIGATE THE FEMORAL VEIN. REPORTED COMPLICATIONS INCLUDED HEMATOMA (N-1), RE-BLEEDING (N-5), STITCH FAILURE AS A RESULT OF SNAPPING OF SUTURE DURING KNOTTING (N-5). WITH THE INCREASING USE OF ANTICOAGULATION AND ANTIPLATELET AGENTS DURING VARIOUS CARDIAC CATHETERIZATION PROCEDURES, THE RATE OF LOCAL ACCESS SITE COMPLICATIONS HAS INCREASED. MANUAL COMPRESSION HAS BEEN THE GOLD STANDARD TO ACHIEVE HEMOSTASIS. SUBCUTANEOUS TEMPORARY ¿FELLOW¿S STITCH¿ IS A SIMPLE, CLINICALLY COST EFFECTIVE, EFFICACIOUS AND SAFE TECHNIQUE TO ACHIEVE IMMEDIATE POST-PROCEDURE VENOUS HEMOSTASIS FOR SHEATH SIZES AT LEAST UP TO 14 FR AMONG WIDE VARIETY OF POPULATION GROUP." Manufacturer Narrative: (B)(4). THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE ADDITIONAL INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ETHICON PRODUCT (:VICRYL SUTURE) INVOLVED CAUSED AND/OR CONTRIBUTED TO THE ADVERSE EVENTS DESCRIBED IN THE ARTICLE ( HEMATOMA , RE-BLEEDING)? WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE COMPLAINT REFERENCE NUMBERS. PATIENT DEMOGRAPHICS THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: CARDIOL RES. 2019; 10(5): 303-308. DOI: HTTPS://DOI.ORG/10.14740/CR931. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9970997&pc=GAM | 2210968-2020-03139 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE OVARIAN PEDICLE LIGATURES, DURING THE TIGHTENING OF THE NODE, THE SUTURE BROKE OVER THE ENTIRE LENGTH WITHOUT AN EXCESSIVE STRAIN. THE NODE OF THE LIGATION IS INTACT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE JV586, LOT QABBKXR0 COULD BE OBSERVED.. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS SUBMITTED VIA 2210968-2020-03136, 2210968-2020-03137, 2210968-2020-03138 AND 2210968-2020-03139.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9970998&pc=GAM | 2210968-2020-03137 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE OVARIAN PEDICLE LIGATURES, DURING THE TIGHTENING OF THE NODE, THE SUTURE BROKE OVER THE ENTIRE LENGTH WITHOUT AN EXCESSIVE STRAIN. THE NODE OF THE LIGATION IS INTACT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE JV586, LOT QABBKXR0 COULD BE OBSERVED.. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT EVENTS SUBMITTED VIA 2210968-2020-03136, 2210968-2020-03138 AND 2210968-2020-03139.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9971001&pc=GAM | 2210968-2020-03138 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE OVARIAN PEDICLE LIGATURES, DURING THE TIGHTENING OF THE NODE, THE SUTURE BROKE OVER THE ENTIRE LENGTH WITHOUT AN EXCESSIVE STRAIN. THE NODE OF THE LIGATION IS INTACT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE JV586, LOT QABBKXR0 COULD BE OBSERVED.. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS SUBMITTED VIA 2210968-2020-03136, 2210968-2020-03137, AND 2210968-2020-03139.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9971002&pc=GAM | 2210968-2020-03136 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE OVARIAN PEDICLE LIGATURES, DURING THE TIGHTENING OF THE NODE, THE SUTURE BROKE OVER THE ENTIRE LENGTH WITHOUT AN EXCESSIVE STRAIN. THE NODE OF THE LIGATION IS INTACT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE (B)(4), LOT QABBKXR0 COULD BE OBSERVED.. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS SUBMITTED VIA 2210968-2020-03136, 2210968-2020-03137, 2210968-2020-03138 AND 2210968-2020-03139.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9971014&pc=GAM | 2210968-2020-03140 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE MUSCULAR WALL SUTURES, AFTER A FEW TISSUE PASSAGES, THE TWISTED NEEDLE PULLED OFF. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE JV586, LOT QABBKXR0 COULD BE OBSERVED.. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT EVENTS SUBMITTED VIA 2210968-2020-03140 AND 2210968-2020-03141.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9971015&pc=GAM | 2210968-2020-03141 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE MUSCULAR WALL SUTURES, AFTER A FEW TISSUE PASSAGES, THE TWISTED NEEDLE PULLED OFF. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE JV586, LOT QABBKXR0 COULD BE OBSERVED. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS SUBMITTED VIA 2210968-2020-03140 AND 2210968-2020-03141.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9971051&pc=GAM | 2210968-2020-03145 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE MUSCULAR WALL SUTURE, AFTER A FEW TISSUE PASSAGES, NOT TWISTED NEEDLE PULLED OFF WITHOUT AN EXCESSIVE STRAIN. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE JV586, LOT QABBKXR0 COULD BE OBSERVED.. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS SUBMITTED VIA 2210968-2020-03142, 2210968-2020-03143, 2210968-2020-03144 AND 2210968-2020-03145.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9971055&pc=GAM | 2210968-2020-03144 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE MUSCULAR WALL SUTURE, AFTER A FEW TISSUE PASSAGES, NOT TWISTED NEEDLE PULLED OFF WITHOUT AN EXCESSIVE STRAIN. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE JV586, LOT QABBKXR0 COULD BE OBSERVED.. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT EVENTS SUBMITTED VIA 2210968-2020-03142, 2210968-2020-03143, 2210968-2020-03144 AND 2210968-2020-03145.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9971056&pc=GAM | 2210968-2020-03143 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE MUSCULAR WALL SUTURE, AFTER A FEW TISSUE PASSAGES, NOT TWISTED NEEDLE PULLED OFF WITHOUT AN EXCESSIVE STRAIN. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE JV586, LOT QABBKXR0 COULD BE OBSERVED.. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT EVENTS SUBMITTED VIA 2210968-2020-03142, 2210968-2020-03144 AND 2210968-2020-03145.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9971058&pc=GAM | 2210968-2020-03142 |  | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN ANIMAL OVARIECTOMY SURGERY IN 2020 AND THE SUTURE WAS USED. DURING THE MUSCULAR WALL SUTURE, AFTER A FEW TISSUE PASSAGES, NOT TWISTED NEEDLE PULLED OFF WITHOUT AN EXCESSIVE STRAIN. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. SUMMARY: TWO PHOTOS WERE RECEIVED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, ONE OPENED FOIL PACKET, THREE USED NEEDLE-SUTURE PIECES AND ONLY ONE OF THEM APPEARS TO BE DISTORTED OF PRODUCT CODE JV586, LOT QABBKXR0 COULD BE OBSERVED.. HOWEVER, NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. TO DATE THE DEVICE HAS NOT BEEN RETURNED. DEVICE FOLLOW-UP DEFERRED DUE TO QUARANTINE RESTRICTIONS. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. EVENTS SUBMITTED VIA 2210968-2020-03142, 2210968-2020-03143, 2210968-2020-03144 AND 2210968-2020-03145.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9971539&pc=GAM | 2210968-2020-03086 | 2020-03-23 | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL RAP UND 70CM 3/0 PS-1 PRIME MP | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ANGIOPLASTY PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. CHANGED TO ANOTHER SUTURE TO COMPLETE THE PROCEDURE WITHOUT DELAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PLBCQGP0, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9972351&pc=GAM | 2210968-2020-03160 | 2020-03-24 | Malfunction | ETHICON INC. | 2020-04-17 | GAM | VICRYL RAP UND 70CM 3/0 SH-1 PLUS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE TOO EASILY CUT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PKBDLKS0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE INITIAL PROCEDURE? WHAT WAS USED TO COMPLETE THE SURGERY? WAS THE SUTURE BROKE DURING THE PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9970607&pc=GAM | 2210968-2020-03119 |  | Malfunction | ETHICON INC. | 2020-04-16 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Tissue Damage | K022269 |  | 1 | Event Description: TITLE: A NEW SUTURE TECHNIQUE FOR PERITONEAL FLAP CLOSURE IN TAPP: A PROSPECTIVE RANDOMIZED CONTROLLED TRIAL. THIS PROSPECTIVE RANDOMIZED CONTROLLED TRIAL AIMED TO EVALUATE THE EFFICACY OF THE NEW SUTURE METHOD FOR LAPAROSCOPIC CLOSURE OF THE PERITONEAL FLAP IN TRANSABDOMINAL PREPERITONEAL (TAPP)COMPARED WITH THE CONVENTIONAL METHOD. IN (B)(6) 2019, 80 PATIENTS WITH PRIMARY UNILATERAL INGUINAL HERNIA WERE RANDOMIZED INTO THE CONVENTIONAL SUTURE METHOD GROUP AND THE NEW SUTURE GROUP FOR TAPP. IN THE PROCEDURE, BOTH TECHNIQUES UTILIZED 3-0 VICRYL ABSORBABLE SUTURE IN PERITONEAL FLAP CLOSURE. OUTCOMES INCLUDED PERITONEAL TEARING (N=15; N=11 CONVENTIONAL GROUP AND N=4 NEW SUTURE GROUP); AND SUTURE LINE BREAKING (N=11; N=8 CONVENTIONAL GROUP AND N=3 NEW SUTURE GROUP). TRADITIONALLY, THE PERITONEUM SUTURE IS STARTED BY SUTURING THE PERITONEUM VERTICALLY; IN CASE THE PERITONEUM IS VERY THIN, THE VERTICAL SUTURE FORCE OF THE NEEDLE AND THE VERTICAL PASSAGE OF THE SUTURE MATERIAL WOULD TEAR THE UPPER FLAP OF THE PERITONEUM, AND THIS COULD BE SEEN FROM OUR OBSERVATIONS. THIS NEW SUTURE METHOD IS EASY TO PERFORM, IS ASSOCIATED WITH REDUCED SUTURE TIME, AND CAUSES LESS PERITONEAL TEARING. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: ALL INFORMATION REQUESTED ARE UNKNOWN. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE COMPLAINT REFERENCE NUMBERS. DOES THE SURGEON BELIEVE THAT ETHICON PRODUCTS: VICRYL ABSORBABLE SUTURE, INVOLVED CAUSED AND/OR CONTRIBUTED TO THE ADVERSE EVENTS DESCRIBED IN THE ARTICLE, SPECIFICALLY: PERITONEAL TEARING AND SUTURE LINE BREAKING? THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: SURG LAPAROSC ENDOSC PERCUTAN TECH 2020;30:18¿21. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9963692&pc=GAM | 2210968-2020-03056 |  | Malfunction | ETHICON INC. | 2020-04-15 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Code Available | K022269 |  | 1 | Event Description: IT WAS REPORTED BY A JOURNAL ARTICLE. TITLE: REVISITING VESICOURETHRAL ANASTOMOSIS DURING OPEN RADICAL RETROPUBIC PROSTATECTOMY; A SIMPLE AND REPRODUCIBLE TECHNIQUE: A SINGLE CENTER EXPERIENCE WITH 200 CASES. THIS STUDY AIMED TO REPORT THE 8-YEAR EXPERIENCE ON A FEASIBLE AND REPRODUCIBLE SUTURING TECHNIQUE IN OPEN RADICAL PROSTATECTOMY. BETWEEN JANUARY 2009 AND JUNE 2017, A TOTAL OF 200 CONSECUTIVE PATIENTS WITH ORGAN CONFINED PROSTATE CANCER UNDERWENT OPEN RETROPUBIC RADICAL PROSTATECTOMY FOR CLINICALLY LOCALIZED PROSTATE CANCER AND WERE SELECTED FOR THE NOVEL MODIFIED VESICOURETHRAL ANASTOMOSIS TECHNIQUE. VESICOURETHRAL ANASTOMOSIS WAS PERFORMED USING AN INTERRUPTED SUTURING TECHNIQUE WITH THE USE OF SIX 3-0 VICRYL SUTURES (ETHICON) STARTING BY TIGHTENING THE ANTERIOR ONE THEN MOVING ON TO THE 10-, 8-, 6-, 4- AND 2-O'CLOCK POSITIONS. COMPLAINTS INCLUDED SUTURES DISRUPTION DURING KNOT TYING (N=2), AND URETHRAL STRICTURE (N=8). THE RESULTS PRESENTED IN THIS STUDY SHOWED THAT ALTHOUGH THE QUALITY OF THE VESICO-URETHRAL ANASTOMOSIS IS UNLIKELY TO HAVE AN IMPACT ON THE ONCOLOGICAL OUTCOME OF RADICAL PROSTATECTOMY, IT UNDOUBTEDLY AFFECTS FUNCTIONAL OUTCOME AND QUALITY OF LIFE. Manufacturer Narrative: (B)(4). THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. ADDITIONAL INFORMATION REQUESTED AND FOLLOWING WAS OBTAINED: I AM PLEASED TO ANSWER THE TWO QUESTIONS ABOUT SUTURE DISRUPTION AND URETHRAL STRICTURE IN RADICAL PROSTATECTOMY WHILE USING VICRYL SUTURE. REGARDING DISRUPTION: IT WAS NOT VICRYL DISRUPTION, ACTUALLY, IT WAS THE DETACHMENT OF BLADDER NECK FROM URETHRA THAT IS NOT UNUSUAL DURING THIS PROCEDURE AND IN THIS ARTICLE, WE HAVE PRESENTED A TECHNIQUE TO MINIMIZE THIS UNPLEASANT COMPLICATION URETHRAL STRICTURE IS ALSO A WELL-KNOWN COMPLICATION OF RADICAL PROSTATECTOMY WHICH IS NOT ONLY DEPENDENT ON THE SUTURE MATERIAL THAT WAS USED. CITATION: UROLOGY JOURNAL 2019; 16: 5, 475-477; DOI: 10.22037/UJ.V0I0.4800 . (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965525&pc=GAM | 2210968-2020-03071 |  | Malfunction | ETHICON INC. | 2020-04-15 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Conjunctivitis | K022269 |  | 1 | Event Description: IT WAS REPORTED IN A JOURNAL ARTICLE TITLE: NOVEL VICRYL RELEASABLE SUTURE TECHNIQUE TO CLOSE LEAKING SCLEROTOMIES IN A TRANSCONJUNCTIVAL VITRECTOMY THIS PROSPECTIVE STUDY AIMED TO EVALUATE A RELEASABLE SUTURE WITH VICRYL TO TREAT THE INCOMPETENT WOUND CLOSURE AND ALSO AVOID THE ADVERSE CONJUNCTIVAL REACTION (ACR) TO VICRYL. A TOTAL OF 32 PATIENTS (N=9 MALES AND N=23 FEMALES, AGE RANGE: 32-82 YEARS, MEAN AGE: 61.83 YEARS) WERE INCLUDED IN THE STUDY. DURING CLOSURE, THE RELEASABLE SUTURE WITH VICRYL 7.0 (ETHICON) WAS USED IN THE SUPERIOR TEMPORAL SCLEROTOMY. BY CONTRAST, THE PATTERN SUTURE WITH VICRYL 7.0 (ETHICON) WAS USED WITH A TRANSCONJUNCTIVAL AND TRANSSCLERAL DOUBLE KNOT IN THE INCOMPETENT NASAL SUPERIOR VITRECTOMY, AND THE INFERIOR TEMPORAL WAS SUTURELESS. THE RELEASABLE SUTURES WERE REMOVED AT THE SLIT LAMP ON DAY 6 OF THE FOLLOW-UP WITH FORCEPS, UNDER TOPICAL ANESTHESIA. COMPLAINT INCLUDED THE PROGRESSION OF ADVERSE CONJUNCTIVAL REACTION UNTIL CONJUNCTIVAL NODULE FORMATION AT THE 30TH DAY OF FOLLOW-UP IN THE PATTERN SUTURE (NONRELEASABLE) SITE (N=27). ALSO, BY DAY 6, SOME OF THE PATIENTS WERE PRESENTED WITH AN UNDONE KNOT (N=3 IN THE RELEASABLE SUTURES AND N=2 IN THE STANDARD SUTURES). IN CONCLUSION, THE AUTHORS HAVE DESCRIBED A NEW, EASILY PERFORMED RELEASABLE TRANSCONJUNCTIVAL AND TRANSSCLERAL VICRYL SUTURE TECHNIQUE FOR 23-GAUGE PARS PLANA VITRECTOMY, WHICH DEMONSTRATED AN EFFECTIVE OUTCOME IN THE TREATMENT OF POSTOPERATIVE WOUND LEAKAGE AND WHICH APPEARS TO HELP AVOID AN ACR WITH VICRYL AFTER ITS REMOVAL. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE ADDITIONAL INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE COMPLAINT REFERENCE NUMBERS. DOES THE SURGEON BELIEVE THAT ETHICON PRODUCT: VICRYL SUTURES, INVOLVED CAUSED AND/OR CONTRIBUTED TO THE ADVERSE EVENTS DESCRIBED IN THE ARTICLE, SPECIFICALLY: ADVERSE CONJUNCTIVAL REACTION AND UNDONE KNOT BY DAY 6? THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: RETINA, THE JOURNAL OF RETINAL AND VITREOUS DISEASES 2019; 39:10, S108-S111. PLEASE SEE ARTICLE ATTACHED. (B)(4).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957864&pc=GAM | 2210968-2020-03023 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-14 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OVARIECTOMY IN (B)(6) 2020 AND SUTURE WAS USED. APPROXIMATELY 5 TO 6 DAYS AFTER THE SURGERY, THE SUTURE WAS FOUND BROKEN ALONG THE LENGTH OF THE THREAD. THE KNOTS WERE INTACT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT LL8CSRQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY ALLEGED DEFICIENCY NOTED WITH THE SUTURE DURING THE SURGERY? PLEASE EXPLAIN. NO FURTHER INFORMATION AVAILABLE, ALL THE INFORMATION AVAILABLE HAVE BEEN PROVIDED IN THE INITIAL DESCRIPTION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9959472&pc=GAM | 2210968-2020-03012 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-14 | GAM | VICRYL RAP UND 70CM 3/0 PS-1 PRIME MP | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 04/14/2020. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE NAME OF THE INITIAL PROCEDURE?, WHAT IS THE CURRENT CONDITION OF THE PATIENT?, COULD YOU PLEASE PROVIDE LOT NUMBER?, HOW MANY DEVICES THAT HAVE SUTURE BREAKAGE ISSUE DURING THIS SINGLE PROCEDURE?, WHEN THIS THE EVENT OCCURRED?. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9959524&pc=GAM | 2210968-2020-03018 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-14 | GAM | VICRYL RAP UND 70CM 3/0 PS-1 PRIME MP | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE NAME OF THE INITIAL PROCEDURE? WHAT IS THE CURRENT CONDITION OF THE PATIENT? COULD YOU PLEASE PROVIDE LOT NUMBER? HOW MANY DEVICES THAT HAVE SUTURE BREAKAGE ISSUE DURING THIS SINGLE PROCEDURE? WHEN THIS THE EVENT OCCURRED? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9959630&pc=GAM | 2210968-2020-03038 | 2020-02-20 | Malfunction | ETHICON INC. | 2020-04-14 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE WHEN INSERTED INTO THE MUSCLE. THE DOCTOR MANAGED TO REMOVE IT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9959696&pc=GAM | 2210968-2020-03006 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-14 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OVARIECTOMY IN (B)(6) 2020 AND SUTURE WAS USED. APPROXIMATELY 5 TO 6 DAYS AFTER THE SURGERY, THE SUTURE WAS FOUND BROKEN ALONG THE LENGTH OF THE THREAD. THE KNOTS WERE INTACT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY ALLEGED DEFICIENCY NOTED WITH THE SUTURE DURING THE SURGERY? PLEASE EXPLAIN. NO FURTHER INFORMATION AVAILABLE, ALL THE INFORMATION AVAILABLE HAVE BEEN PROVIDED IN THE INITIAL DESCRIPTION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT LL8CSRQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9959712&pc=GAM | 2210968-2020-03011 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-14 | GAM | VICRYL RAP UND 70CM 3/0 PS-1 PRIME MP | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020; AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 04/14/2020. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE NAME OF THE INITIAL PROCEDURE?, WHAT IS THE CURRENT CONDITION OF THE PATIENT?, COULD YOU PLEASE PROVIDE LOT NUMBER?, HOW MANY DEVICES THAT HAVE SUTURE BREAKAGE ISSUE DURING THIS SINGLE PROCEDURE? WHEN THIS THE EVENT OCCURRED?. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953197&pc=GAM | 2210968-2020-02966 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-13 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OVARIECTOMY IN (B)(6) 2020 AND SUTURE WAS USED. APPROXIMATELY 5 TO 6 DAYS AFTER THE SURGERY, THE SUTURE WAS FOUND BROKEN ALONG THE LENGTH OF THE THREAD. THE KNOTS WERE INTACT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT LL8CSRQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY ALLEGED DEFICIENCY NOTED WITH THE SUTURE DURING THE SURGERY? PLEASE EXPLAIN. NO FURTHER INFORMATION AVAILABLE, ALL THE INFORMATION AVAILABLE HAVE BEEN PROVIDED IN THE INITIAL DESCRIPTION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953201&pc=GAM | 2210968-2020-02967 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-13 | GAM | VICRYL CT VIO 75CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A CAT UNDERWENT AN OVARIECTOMY IN (B)(6) 2020 AND SUTURE WAS USED. APPROXIMATELY 5 TO 6 DAYS AFTER THE SURGERY, THE SUTURE WAS FOUND BROKEN ALONG THE LENGTH OF THE THREAD. THE KNOTS WERE INTACT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT LL8CSRQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY ALLEGED DEFICIENCY NOTED WITH THE SUTURE DURING THE SURGERY? PLEASE EXPLAIN NO FURTHER INFORMATION AVAILABLE, ALL THE INFORMATION AVAILABLE HAVE BEEN PROVIDED IN THE INITIAL DESCRIPTION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944351&pc=GAM | 2210968-2020-02861 | 2020-02-13 | Malfunction | ETHICON INC. | 2020-04-09 | GAM | VICRYL RAP UND 75CM M3 | Appropriate Term/Code Not Available | Unspecified Infection; Impaired Healing | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT RESUMPTION OF TRANSTIBIAL AMPUTATION ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE REPAIR OF THE DRESSING ON (B)(6) 2020, SEVEN WEEKS AFTER THE PROCEDURE, THE SUBCUTANEOUS SUTURE WAS STILL VISIBLE, WHICH CAUSED A DELAY IN HEALING AND AN INFECTIOUS RISK. THE APPARENT SUBCUTANEOUS THREAD WAS THEREFORE REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PROCEDURE NAME? WAS THERE ANY INFECTION NOTICED WITH PATIENT? IF YES, WHERE WAS THE INFECTION NOTICED IN PATIENT? WAS THERE ANY TREATMENT ( MEDICAL OR SURGICAL INTERVENTION / ANTIBIOTICS OR PRESCRIPTION MEDICATION) PROVIDED TO PATIENT FOR THE INFECTION? WAS THERE ANY OTHER TREATMENT ( MEDICAL OR SURGICAL INTERVENTION / ANTIBIOTICS OR PRESCRIPTION MEDICATION) / RE-SUTURING PROVIDED TO PATIENT AFTER THE REMOVAL OF STITCHES POST-OP? PATIENT CURRENT CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938920&pc=GAM | 2210968-2020-02772 | 2020-02-02 | Malfunction | ETHICON INC. | 2020-04-08 | GAM | VICRYL PLUS CT VIO 90CM M4::USP1 S/A CT | Material Separation | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A LAPAROSCOPIC MYOMECTOMY SURGERY ON (B)(6)2020, AND A SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE PULLED OFF FROM THE NEEDLE. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT LK8CLLQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9939047&pc=GAM | 2210968-2020-02792 | 2020-02-04 | Malfunction | ETHICON INC. | 2020-04-08 | GAM | VICRYL*VLT PLUS 0 90CM 1-CT1 | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN MYOMECTOMY SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN THE SUTURE WAS USED, THE FILAMENT RUPTURED. THERE WERE PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ATTEMPTS TO OBTAIN THE FOLLOWING INFORMATION HAS BEEN PERFORMED AND THE FOLLOWING WAS RECEIVED. IF THE FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS USED TO COMPLETE THE PROCEDURE? A NEW SUTURE THREAD FROM THE SAME CODE THAT PRESENTED THE FAILURE WAS USED. WHAT TISSUE WAS BEING APPROACHED OR SUTURED WHEN IT RUPTURED? THE APPROACHING TISSUE WAS THE UTERUS (MYOMECTOMY SURGERY)" EVALUATION: IT WAS RECEIVED FOR ANALYSIS AN EMPTY OPENED FOIL, AN EMPTY LABELED WINDING FORMER AND TWO SUTURES PIECES OF PRODUCT CODE (B)(4), LOT AL7849. DURING VISUAL INSPECTION OF THE SUTURES PIECE, BODY FLUIDS WERE FOUND, AND THE ENDS WERE CUT THAT APPEARS TO BE BY USE OF SURGICAL INSTRUMENT. DUE TO THE CONDITION OF THE SAMPLE THE FUNCTIONAL TEST CAN¿T BE PERFORMED. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO CONDITION OF THE OPENED SAMPLE, THE ASSIGNABLE CAUSE OF PERFORMANCE BREAKAGE SUTURE SUGGESTS AN IMPROPER HANDLING AND THE COMPLAINT IS OBSERVED BY EXTERNAL CAUSE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9939232&pc=GAM | 2210968-2020-02795 |  | Malfunction | ETHICON INC. | 2020-04-08 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Break; Inadequacy of Device Shape and/or Size | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A VAGINAL DELIVERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE WAS BROKEN ON THE MIDDLE OF THE SUTURE DURING SUTURING FOR PERINEAL TEAR. A NURSE COMMENTED THAT THE SUTURE WAS THINNER THAN USUAL. THE PRODUCT INVOLVED WAS DISCARDED. THE LOT NUMBER IS UNKNOWN. FURTHER DETAILS ARE NOT PROVIDED . NO SAMPLE WILL BE RETURNED. THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT. Manufacturer Narrative: (B)(4). ATTEMPTS TO OBTAIN THE FOLLOWING INFORMATION HAS BEEN PERFORMED AND THE FOLLOWING WAS RECEIVED. IF THE FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS USED TO COMPLETE THE PROCEDURE? NO FURTHER INFORMATION IS AVAILABLE. WAS THE SURGERY COMPLETED SUCCESSFULLY? NO FURTHER INFORMATION IS AVAILABLE. DID THE NURSE COMPARE THE SUTURE WITH THE TEMPLATE IN THE PACKAGE? NO FURTHER INFORMATION IS AVAILABLE. NO FURTHER INFORMATION WILL BE PROVIDED. TO DATE THE DEVICE HAS NOT BEEN RECEIVED. IF THE FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9940519&pc=GAM | 2210968-2020-02800 | 2020-03-05 | Malfunction | ETHICON INC. | 2020-04-08 | GAM | VICRYL VIO 30CM M0.5 USP7/0 DBLE ARMED TG140-8 | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE WITHOUT ANY EXTERNAL FORCE DURING TRAUMA SUTURING PROCEDURE. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9932903&pc=GAM | 2210968-2020-02675 | 2020-03-06 | Malfunction | ETHICON INC. | 2020-04-07 | GAM | VICRYL*RAPIDE*UND 3-0 75CM V-4 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNSPECIFIED PROCEDURE ON (B)(6) 2020 USING SUTURE TO PERFORM DERMO-DERMAL OVERLOCK. AFTER FIVE OR SIX TISSUE PASSAGE, THE NEEDLE BROKE AT THE HEEL WITHOUT AN EXCESSIVE TENSION ON THE WIRE. Manufacturer Narrative: (B)(4). INVESTIGATION SUMMARY ==> IT WAS REPORTED BREAKAGE NEEDLE. THREE PHOTOS WERE PROVIDED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, ONE FOIL AND ONE BROKEN NEEDLE-SUTURE OF PRODUCT CODE VR2280, LOT AK4764 COULD BE OBSERVED. HOWEVER; NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AK4764, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928628&pc=GAM | 2210968-2020-02658 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-04-06 | GAM | VICRYL* RAPIDE* UND75 2-0 FS | Break; Degraded; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE RIGHT AT THE END OF THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928630&pc=GAM | 2210968-2020-02659 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-04-06 | GAM | VICRYL RAP UND 75CM M1.5 USP4/0 SGLE ARMED FS-2 | Break; Degraded; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE RIGHT AT THE END OF THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE NAME OF THE PROCEDURE? WHAT WAS USED TO COMPLETE THE PROCEDURE? WAS THE SURGERY COMPLETED SUCCESSFULLY? WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928823&pc=GAM | 2210968-2020-02663 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-06 | GAM | VICRYL CT VIO 75CM M2 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED ON LIGATURES OF THE OVARIAN PEDICLES, THE SEPARATE POINTS OF THE ABDOMINAL WALL AND THE SKIN OVERLOCK. TWO OR THREE DAYS AFTER THE PROCEDURE, THE SKIN OVERLOCK GOT BROKEN IN THE LENGTH, THE KNOTS WERE INTACT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PGBDDXQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928824&pc=GAM | 2210968-2020-02662 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-06 | GAM | VICRYL CT VIO 75CM M2 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED ON LIGATURES OF THE OVARIAN PEDICLES, THE SEPARATE POINTS OF THE ABDOMINAL WALL AND THE SKIN OVERLOCK. TWO OR THREE DAYS AFTER THE PROCEDURE, THE SKIN OVERLOCK GOT BROKEN IN THE LENGTH, THE KNOTS WERE INTACT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PGBDDXQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928831&pc=GAM | 2210968-2020-02661 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-06 | GAM | VICRYL CT VIO 75CM M2 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED ON LIGATURES OF THE OVARIAN PEDICLES, THE SEPARATE POINTS OF THE ABDOMINAL WALL AND THE SKIN OVERLOCK. TWO OR THREE DAYS AFTER THE PROCEDURE, THE SKIN OVERLOCK GOT BROKEN IN THE LENGTH, THE KNOTS WERE INTACT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PGBDDXQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9929278&pc=GAM | 2210968-2020-02667 | 2020-03-01 | Malfunction | ETHICON INC. | 2020-04-06 | GAM | VICRYL CT VIO 75CM M2 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT AN ANIMAL UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED ON LIGATURES OF THE OVARIAN PEDICLES, THE SEPARATE POINTS OF THE ABDOMINAL WALL AND THE SKIN OVERLOCK. TWO OR THREE DAYS AFTER THE PROCEDURE, THE SKIN OVERLOCK GOT BROKEN IN THE LENGTH, THE KNOTS WERE INTACT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PGBDDXQ0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9931785&pc=GAM | 2210968-2020-02692 | 2020-03-09 | Malfunction | ETHICON INC. | 2020-04-06 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 2-0 UND | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN OB-GYN SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. WHEN TAKING OUT THE NEEDLE-SUTURE FROM THE PACKAGE, IT WAS FOUND THAT THE NEEDLE TIP HAD BEEN CRUSHED. IT WAS NOT USED ON THE PATIENT. ANOTHER DEVICE PACKAGE WAS USED WITH NO PROBLEM. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WAS THE OUTER CASE THAT THE PACKAGE ARRIVED IN DAMAGED? NO FURTHER INFORMATION IS AVAILABLE. DO YOU HAVE ANY PHOTOS ? NO PHOTOS ARE AVAILABLE. WHERE DID YOU RECEIVE THE PRODUCT FROM (ETHICON, DISTRIBUTOR, ETC)? IF "YOU" MEANS THE HOSPITAL, IT RECEIVED THE PRODUCT FROM A DISTRIBUTOR. WAS THE SUTURE SEALS ON THE FOIL STILL INTACT? NO FURTHER INFORMATION IS AVAILABLE. DEVICE RETURN STATUS/FOLLOW UP - WE REGULARLY CONTACT WITH SALES REP ABOUT THE DEVICE RETURNING. NO FURTHER INFORMATION WILL BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9924703&pc=GAM | 2210968-2020-02624 | 2020-03-04 | Malfunction | ETHICON INC. | 2020-04-03 | GAM | VICRYL VIO 75CM M3.5::USP0 SGLE ARMED SH PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PERITONEAL SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. THE NEEDLE GOT BROKEN WHEN PULLING ON THE THREAD. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9903117&pc=GAM | 2210968-2020-02477 |  | Malfunction | ETHICON INC. | 2020-03-31 | GAM | VICRYL VIO 70CM M2 | Break; Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT THE DOG UNDERWENT AN ANIMAL SURGERY IN (B)(6) 2020 ABOUT TWO WEEKS AGO AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE AND UNRAVELED POST-OP. Manufacturer Narrative: PRODUCT COMPLAINT # ==> (B)(4). DATE SENT TO THE FDA: 04/23/2020. INVESTIGATION SUMMARY ==> IT WAS REPORTED PERFORMANCE BREAKAGE SUTURE POST-OP AND PERFORMANCE FRAYING SUTURE POST-OP. IT WAS RECEIVED FOR ANALYSIS A SEALED BOX WITH THIRTY-SIX UNOPENED SAMPLES AND OPENED BOX WITH THIRTY-ONE UNOPENED SAMPLES OF PRODUCT CODE V316H, LOT PPBBGSQ0. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF THIRTEEN UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED, FRAYING OR SUTURE BREAKAGE WERE NOTED. A FUNCTIONAL TEST WAS PERFORMED AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. PER THE CONDITION OF THE SAMPLES RECEIVED, NO PERFORMANCE ¿ BREAKAGE POST-OP SUTURE OR PERFORMANCE FRAYING SUTURE POST-OP WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9903511&pc=GAM | 2210968-2020-02483 |  | Malfunction | ETHICON INC. | 2020-03-31 | GAM | VICRYL VIO 70CM M2 | Break; Material Frayed | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ANIMAL SURGERY IN (B)(6) 2020 ABOUT TWO WEEKS AGO AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE AND UNRAVELED POST-OP. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. SUMMARY: PICTURES WERE RETURNED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES A DISPENSED NEEDLE/SUTURE AND A SEALED BOX OF THE PRODUCT COULD BE OBSERVED, AND NO CONCLUSION COULD BE REACH AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9901789&pc=GAM | 2210968-2020-02468 |  | Malfunction | ETHICON INC. | 2020-03-30 | GAM | VICRYL RAP UND 75CM M1 USP5/0 SGLE ARMED C-3 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A DENTAL SURGERY ON UNKNOWN DATE AND THE SUTURE WAS USED. THE THREAD WAS BROKEN ON THE CRIMP DURING THE PROCEDURE. ANOTHER LIKE SUTURE WAS USED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9887258&pc=GAM | 2210968-2020-02411 |  | Malfunction | ETHICON INC. | 2020-03-27 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A MYOMECTOMY ON UNKNOWN DATE AND THE SUTURE WAS USED. IT WAS REPORTED THAT ON FIRST LAYER OF THE UTERUS, THE NEEDLE BROKE, SPLIT INTO TWO. THE NEEDLE WAS RETRIEVED FROM THE PATIENT UNDER X-RAY. A BIT OF DISSECTION WAS DONE ON THE SECOND LAYER OF THE UTERUS. THE PATIENT HAS NO COMORBIDITIES, FIT AND WELL. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS WERE MADE TO OBTAIN THE DEVICE RETURN FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: COULD YOU PLEASE CONFIRM WHAT WAS THE INITIAL PROCEDURE? PATIENT PROCEDURE WAS OPEN MYOMECTOMY. WHAT IS THE CONDITION OF THE PATIENT? PATIENT HAS NO CO MORBIDITIES, FIT AND WELL. WERE X-RAYS TAKEN TO LOCATE THE NEEDLE PIECE(S)? YES. WAS THE NEEDLE PIECE(S) RETRIEVED DURING THE SAME PROCEDURE? YES. DOES A PIECE OF THE NEEDLE REMAIN IN THE PATIENT¿S TISSUE? NO, WAS RETRIEVED UNDER THE X-RAY. WHAT TISSUE STRUCTURE IS IT LOCATED? SIZE OF THE NEEDLE PIECE RETAINED? FIRST LAYER OF THE UTERUS, NEEDLE WAS SPLIT INTO TWO. ARE ANY PLANS IN PLACE TO REMOVE THE NEEDLE PIECE IN FUTURE? HAS BEEN REMOVED. IF RETAINED, WHAT IS THE SURGEON'S OPINION OF CONSEQUENCES TO THE PATIENT? NOT RETAINED. WAS THERE ANY ADDITIONAL TISSUE DAMAGE AS A RESULT OF SEARCHING FOR THE NEEDLE PIECE? A BIT OF DISSECTION DONE ON THE SECOND LAYER OF THE UTERUS. DEVICE RETURN - STAFF UNFORTUNATELY THROWN IN THE SHARPS BIN BUT TOO LATE TO PICK UP TO SEND TO COMPANY FOR ANALYSIS.¿
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9892361&pc=GAM | 2210968-2020-02443 |  | Malfunction | ETHICON INC. | 2020-03-27 | GAM | VICRYL VIO 75CM M3 | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON UNKNOWN DATE AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE CRIMP OF THE NEEDLE WAS BROKEN AND WAS LOST DURING THE CLOSURE OF THE SKIN. THE NEEDLE FELL ON THE FLOOR. THE IMAGE INTENSIFIER WAS USED TO CHECK FOR THE ABSENCE OF THE NEEDLE IN THE PATIENT. INTRAOPERATIVE RADIOSCOPY WAS NEGATIVE. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED YET. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: DID THE NEEDLE BREAK AT THE SWAGED AREA? OR NEEDLE PULLED OFF FROM THE SUTURE THREAD? THE NEEDLE WAS LOST DURING THE CLOSURE OF THE SKIN. HOW WAS CASE COMPLETED? THE IMAGE INTENSIFIER WAS USED TO CHECK FOR THE ABSENCE OF THE NEEDLE IN THE PATIENT. INTRAOPERATIVE RADIOSCOPY: NEGATIVE CONTROL.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9876287&pc=GAM | 2210968-2020-02328 |  | Malfunction | ETHICON INC. | 2020-03-25 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Ptosis | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A PTOSIS INCISION OPERATION ON AN UNKNOWN DATE IN 2019 AND SUTURE WAS USED. THERE WAS AN ABSORPTION ISSUE WITH THE INTERNAL SUTURE. NO OTHER INFORMATION IS KNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ABSORPTION ISSUE. ATTEMPTS TO OBTAIN THE FOLLOWING INFORMATION HAVE BEEN PERFORMED BUT NOT RECEIVED. IF THE FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT HOW MANY DAYS POST-OP WAS THE ABSORPTION ISSUE OCCUR? WAS THE ABSORPTION NOTED DURING REGULAR POST-OP VISIT? WAS THE WOUND INFECTED? WAS RE-SUTURING DONE? WHAT MEDICAL/SURGICAL TREATMENT WAS PROVIDED? PLEASE PROVIDE THE MEDICINE NAME, STRENGTH AND DOSE IF ADMINISTERED. WHAT IS YOUR PRESENT CONDITION? WHAT DATE DID THE PATIENT PRESENT WITH SYMPTOMS? DO YOU KNOW WHAT IS THE LOT NUMBER? TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9873470&pc=GAM | 2210968-2020-02306 | 2020-02-25 | Malfunction | ETHICON INC. | 2020-03-24 | GAM | VICRYL*RAPID UND 4-0 70CM 1PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DIAGNOSTIC LAPAROSCOPY ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, AT THE TIME OF USING THE SUTURE, RUPTURE OF THE FILAMENTS OCCURRED. ANOTHER LIKE DEVICE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. THE CURRENT CONDITION OF THE PATIENT WAS REPORTED AS GOOD. NO FURTHER INFORMATION AVAILABLE. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS THE NAME OF THE INITIAL PROCEDURE? DIAGNOSTIC LAPAROSCOPY. WAS THERE ANY CONSEQUENCE OR INJURY TO THE PATIENT? THERE WAS NO INJURY. THERE WERE NO CONSEQUENCES.. WHAT IS THE PATIENT'S CONDITION? GOOD PATIENT CONDITIONS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9873484&pc=GAM | 2210968-2020-02305 | 2020-02-25 | Malfunction | ETHICON INC. | 2020-03-24 | GAM | VICRYL*RAPID UND 4-0 70CM 1PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DIAGNOSTIC LAPAROSCOPY ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, AT THE TIME OF USING THE SUTURE, RUPTURE OF THE FILAMENTS OCCURRED. ANOTHER LIKE DEVICE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES. THE CURRENT CONDITION OF THE PATIENT WAS REPORTED AS GOOD. NO FURTHER INFORMATION AVAILABLE. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS THE NAME OF THE INITIAL PROCEDURE? DIAGNOSTIC LAPAROSCOPY. WAS THERE ANY CONSEQUENCE OR INJURY TO THE PATIENT? THERE WAS NO INJURY. THERE WERE NO CONSEQUENCES. WHAT IS THE PATIENT'S CONDITION? GOOD PATIENT CONDITIONS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9874032&pc=GAM | 2210968-2020-02312 | 2020-02-25 | Malfunction | ETHICON INC. | 2020-03-24 | GAM | VICRYL* VLT PLUS 0 70CM (1)SH | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, AT THE TIME OF USING THE SUTURE, RUPTURE OF THE FILAMENTS OCCURRED. THEREFORE, IT IS REQUIRED TO USE MORE SUTURES. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE NAME OF THE INITIAL PROCEDURE? WAS THERE ANY PATIENT CONSEQUENCE OR INJURY? WHAT IS THE CONDITION OF THE PATIENT? COULD YOU PLEASE CONFIRM IF WILL YOU RETURN DEVICE, IN PRODUCT RETURN STATUS IT SAID AVAILABILITY UNKNOWN? HOW MANY OF THE TOTAL QUANTITY WERE ACTUALLY INVOLVED FOR THE REPORTED ISSUE? WHAT DO YOU MEAN BY "RUPTURE OF FILAMENTS"? DID THE SUTURE THREAD FRAY AND BREAK INTO 2 PIECES? OR DID THE SUTURE ONLY EXHIBITED FRAYING OF THREAD?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9874035&pc=GAM | 2210968-2020-02311 | 2020-02-25 | Malfunction | ETHICON INC. | 2020-03-24 | GAM | VICRYL* VLT PLUS 0 70CM (1)SH | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, AT THE TIME OF USING THE SUTURE, RUPTURE OF THE FILAMENTS OCCURRED. THEREFORE, IT IS REQUIRED TO USE MORE SUTURES. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE NAME OF THE INITIAL PROCEDURE? WAS THERE ANY PATIENT CONSEQUENCE OR INJURY? WHAT IS THE CONDITION OF THE PATIENT? COULD YOU PLEASE CONFIRM IF WILL YOU RETURN DEVICE, IN PRODUCT RETURN STATUS IT SAID AVAILABILITY UNKNOWN? HOW MANY OF THE TOTAL QUANTITY WERE ACTUALLY INVOLVED FOR THE REPORTED ISSUE? WHAT DO YOU MEAN BY "RUPTURE OF FILAMENTS"? DID THE SUTURE THREAD FRAY AND BREAK INTO 2 PIECES? OR DID THE SUTURE ONLY EXHIBITED FRAYING OF THREAD?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9866197&pc=GAM | 2210968-2020-02283 |  | Malfunction | ETHICON INC. | 2020-03-23 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. BEFORE USE ON THE PATIENT, THE NEEDLE PULLED OFF FROM THE THREAD. THERE WERE NO PATIENT INVOLVEMENT/PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ATTEMPTS WERE MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9866198&pc=GAM | 2210968-2020-02282 |  | Malfunction | ETHICON INC. | 2020-03-23 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. BEFORE USE ON THE PATIENT, THE NEEDLE PULLED OFF FROM THE THREAD. THERE WERE NO PATIENT INVOLVEMENT/PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ATTEMPTS WERE MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9866224&pc=GAM | 2210968-2020-02284 |  | Malfunction | ETHICON INC. | 2020-03-23 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. BEFORE USE ON THE PATIENT, THE NEEDLE PULLED OFF FROM THE THREAD. THERE WERE NO PATIENT INVOLVEMENT/PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ATTEMPTS WERE MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9866264&pc=GAM | 2210968-2020-02285 |  | Malfunction | ETHICON INC. | 2020-03-23 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Adverse Event Without Identified Device or Use Problem | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. BEFORE USE ON THE PATIENT, THE NEEDLE PULLED OFF FROM THE THREAD. THERE WERE NO PATIENT INVOLVEMENT/PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). ATTEMPTS WERE MADE TO RECEIVE A DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9861513&pc=GAM | 2210968-2020-02266 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-03-20 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K132580 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE THE DEVICE WAS BROKEN EASILY DURING USE. THE PROCEDURE WAS NOT DELAYED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PE6532 BATCH NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9861516&pc=GAM | 2210968-2020-02265 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-03-20 | GAM | VICRYL PLUS SUTURE 27"(70CM) 3-0 VIO | Break | No Consequences Or Impact To Patient | K132580 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE THE DEVICE WAS BROKEN EASILY DURING USE. THE PROCEDURE WAS NOT DELAYED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PE6532 BATCH NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9832965&pc=GAM | 2210968-2020-02087 | 2020-02-19 | Malfunction | ETHICON INC. | 2020-03-14 | GAM | VICRYL RAP UND 70CM 3/0 PS-1 PRIME MP | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN SUTURE WAS BEING TIED, THE NEEDLE WAS DETACHING. THERE WAS NO ADVERSE PATIENT CONSEQUENCE REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). ATTEMPTS ARE BEING MADE TO RECEIVE DEVICE FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824454&pc=GAM | 2210968-2020-01984 |  | Malfunction | ETHICON INC. | 2020-03-12 | GAM | VICRYL* RAPIDE* UND75 2-0 FS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE PRODUCT BROKE OFF AT THE SUTURE AND NEEDLE JUNCTION. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WAS SURGERY DELAYED DUE TO THE REPORTED EVENT? NO, WAS PROCEDURE SUCCESSFULLY COMPLETED? UNKNOWN, WERE FRAGMENTS GENERATED? UNKNOWN, IF YES, WERE THEY REMOVED EASILY WITHOUT ADDITIONAL INTERVENTION? UNKNOWN, PATIENT STATUS/ OUTCOME / CONSEQUENCES. NO, WAS OTHER MEDICAL INTERVENTION (E.G. X-RAYS, ADDITIONAL PROCEDURES, PRESCRIPTIONS, OTC, REVISION) REQUIRED: UNKNOWN. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY THE EVENT; DID THE SUTURE BREAK OR DID THE NEEDLE PULL OFF THE SUTURE THREAD. DEVICE RETURN STATUS/ FOLLOW UP.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9827549&pc=GAM | 2210968-2020-01998 |  | Malfunction | ETHICON INC. | 2020-03-12 | GAM | VICRYL RAPIDE VIO 30CM M0.5 | Break; Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BROKEN OR HAS FALLEN OFF THE THREAD. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: HOW MANY SUTURES ARE BROKEN? HOW MANY SUTURES HAS FALLEN OFF THE THREAD? DEVICE RETURN STATUS/FOLLOW UP? THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9827551&pc=GAM | 2210968-2020-01999 |  | Malfunction | ETHICON INC. | 2020-03-12 | GAM | VICRYL RAPIDE VIO 30CM M0.5 | Break; Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BROKEN OR HAS FALLEN OFF THE THREAD. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: HOW MANY SUTURES ARE BROKEN? HOW MANY SUTURES HAS FALLEN OFF THE THREAD? DEVICE RETURN STATUS/FOLLOW UP? THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9827552&pc=GAM | 2210968-2020-02000 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-03-12 | GAM | VICRYL RAPIDE VIO 30CM M0.5 | Break; Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE WAS BROKEN OR HAS FALLEN OFF THE THREAD. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: HOW MANY SUTURES ARE BROKEN? HOW MANY SUTURES HAS FALLEN OFF THE THREAD? DEVICE RETURN STATUS/FOLLOW UP? THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9817897&pc=GAM | 2210968-2020-01919 | 2019-09-01 | Malfunction | ETHICON INC. | 2020-03-11 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN OOPHORECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV587; PABBKBR0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PULL OFF - SUTURE NEEDLE. ONE WINDING FORMER WITH A SUTURE OF PRODUCT CODE V587, LOT PABBKBR0 WAS RECEIVED FOR ANALYSIS. DURING VISUAL INSPECTION OF THE SAMPLE, THE SUTURE COULD NOT BE DISPENSED SINCE IT HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THIS CAUSED THAT THE NEEDLE WAS DETACHED OF THE STRAND. THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED. ADDITIONALLY, THE NEEDLE AND FOIL WERE NOT RETURNED TO DETERMINED THE ASSIGNABLE CAUSE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE SAMPLE CONDITION, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT OF: ¿THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE ¿.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9817902&pc=GAM | 2210968-2020-01920 | 2019-09-01 | Malfunction | ETHICON INC. | 2020-03-11 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN OOPHORECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE JV587; PABBKBR0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. INVESTIGATION SUMMARY: IT WAS REPORTED PULL OFF - SUTURE NEEDLE. ONE WINDING FORMER WITH A SUTURE OF PRODUCT CODE V587, LOT PABBKBR0 WAS RECEIVED FOR ANALYSIS. DURING VISUAL INSPECTION OF THE SAMPLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE WAS EXAMINED UNDER 20X MAGNIFICATION AND REMNANT SUTURE WAS OBSERVED. THE SUTURE COULD NOT BE DISPENSED SINCE IT HAS BEGUN WITH THE PROCESS OF DEGRADATION AND THIS CAUSED THAT THE NEEDLE WAS DETACHED OF THE STRAND. THE TIME OF EXPOSURE OF SUTURE TO THE ENVIRONMENT COULD NOT BE DETERMINED. ADDITIONALLY, THE FOIL WAS NOT RETURNED TO DETERMINED THE ASSIGNABLE CAUSE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE SAMPLE CONDITION, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT OF: ¿THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE.¿
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819922&pc=GAM | 2210968-2020-01931 | 2020-02-11 | Malfunction | ETHICON INC. | 2020-03-11 | GAM | VICRYL VIO 70CM M3 | Material Separation | No Consequences Or Impact To Patient; Blood Loss | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN GYNECOLOGICAL PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, WHILE SUTURING THE VAGINAL WALL, THE NEEDLE WAS NOT VISIBLE ON OTHER SIDE DUE TO BLOOD LOSS. RETRIEVED WIRE AND NEEDLE BY GENTLY PULLING THE WIRE (BY HAND, NEEDLE-HOLDER HAS NOT TOUCHED WIRE) AND OBTAINED ONLY THE WIRE WITHOUT RESISTANCE. THE NEEDLE LEFT IN THE PATIENT BUT IT WAS REMOVED COMPLETELY. A NEW SUTURE WAS TAKEN FROM THE BOX. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PPBDPBQ0, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: DID THE PATIENT LOSS MORE THAN 250 CC BLOOD? NO. WHAT IS THE CURRENT CONDITION OF THE PATIENT? UNKNOWN. PLEASE CLARIFY:"NEEDLE LEFT IN PTE. WAS ABLE TO REMOVE THE NEEDLE COMPLETELY. YES. WAS THE NEEDLE REMOVED FROM PATIENT? YES. HOW MANY PRODUCTS OF THE TOTAL QUANTITY WERE ACTUALLY INVOLVED FOR REPORTED ISSUE? UNKNOWN. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THERE ANY ADVERSE PATIENT CONSEQUENCES?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814030&pc=GAM | 2210968-2020-01894 |  | Malfunction | ETHICON INC. | 2020-03-10 | GAM | VICRYL VIO 90CM M3.5::USP0 SGLE ARMED MH PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN SURGEON PERFORMED SINGLE KNOTS AND REALIZED THAT THE SUTURE BROKE AT THE SWAGE WHILE SUTURING. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. AN EMPTY OPENED FOIL AND SIX UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITION OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF - SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS COMPLAINT WAS REPORTED WITH QUANTITY INVOLVED OF 7 SUTURES BROKEN DURING THIS SINGLE SURGERY. AND ON OUR QUESTION ¿CONFIRM SPECIFIC NUMBER OF DEVICES THAT FAILED DURING THIS REPORTED EVENT /PROCEDURE?¿ RESPONSE WAS ¿AT LEAST 10 EA. DURING DIFFERENT PROCEDURES. ¿ WERE THESE "DIFFERENT PROCEDURES" REPORTED TO ETHICON BEFORE? IF YES, PLEASE PROVIDE COMPLAINT PC NUMBERS. PLEASE SPECIFY DETAILS OF EACH PROCEDURE AND HOW MANY OF SUTURES WERE BROKE DURING EACH PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814457&pc=GAM | 2210968-2020-01901 |  | Malfunction | ETHICON INC. | 2020-03-10 | GAM | VICRYL VIO 90CM M3.5::USP0 SGLE ARMED MH PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN SURGEON PERFORMED SINGLE KNOTS AND REALIZED THAT THE SUTURE BROKE AT THE SWAGE WHILE SUTURING. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. SUMMARY: AN EMPTY OPENED FOIL AND SIX UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITION OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF - SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS COMPLAINT WAS REPORTED WITH QUANTITY INVOLVED OF 7 SUTURES BROKEN DURING THIS SINGLE SURGERY. AND ON OUR QUESTION ¿CONFIRM SPECIFIC NUMBER OF DEVICES THAT FAILED DURING THIS REPORTED EVENT /PROCEDURE?¿ RESPONSE WAS ¿AT LEAST 10 EA. DURING DIFFERENT PROCEDURES." WERE THESE " DIFFERENT PROCEDURES" REPORTED TO ETHICON BEFORE? IF YES, PLEASE PROVIDE COMPLAINT PC NUMBERS. PLEASE SPECIFY DETAILS OF EACH PROCEDURE AND HOW MANY OF SUTURES WERE BROKE DURING EACH PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814467&pc=GAM | 2210968-2020-01902 |  | Malfunction | ETHICON INC. | 2020-03-10 | GAM | VICRYL VIO 90CM M3.5::USP0 SGLE ARMED MH PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN SURGEON PERFORMED SINGLE KNOTS AND REALIZED THAT THE SUTURE BROKE AT THE SWAGE WHILE SUTURING. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. AN EMPTY OPENED FOIL AND SIX UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITION OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF - SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS COMPLAINT WAS REPORTED WITH QUANTITY INVOLVED OF 7 SUTURES BROKEN DURING THIS SINGLE SURGERY. AND ON OUR QUESTION ¿CONFIRM SPECIFIC NUMBER OF DEVICES THAT FAILED DURING THIS REPORTED EVENT /PROCEDURE?¿ RESPONSE WAS ¿AT LEAST 10 EA. DURING DIFFERENT PROCEDURES.¿ WERE THESE " DIFFERENT PROCEDURES" REPORTED TO ETHICON BEFORE? IF YES, PLEASE PROVIDE COMPLAINT PC NUMBERS. PLEASE SPECIFY DETAILS OF EACH PROCEDURE AND HOW MANY OF SUTURES WERE BROKE DURING EACH PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814728&pc=GAM | 2210968-2020-01904 |  | Malfunction | ETHICON INC. | 2020-03-10 | GAM | VICRYL VIO 90CM M3.5::USP0 SGLE ARMED MH PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN SURGEON PERFORMED SINGLE KNOTS AND REALIZED THAT THE SUTURE BROKE AT THE SWAGE WHILE SUTURING. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. SUMMARY: AN EMPTY OPENED FOIL AND SIX UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITION OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF - SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS COMPLAINT WAS REPORTED WITH QUANTITY INVOLVED OF 7 SUTURES BROKEN DURING THIS SINGLE SURGERY. AND ON OUR QUESTION ¿CONFIRM SPECIFIC NUMBER OF DEVICES THAT FAILED DURING THIS REPORTED EVENT /PROCEDURE?¿ RESPONSE WAS "AT LEAST 10 EA. DURING DIFFERENT PROCEDURES." WERE THESE "DIFFERENT PROCEDURES" REPORTED TO ETHICON BEFORE? IF YES, PLEASE PROVIDE COMPLAINT PC NUMBERS. PLEASE SPECIFY DETAILS OF EACH PROCEDURE AND HOW MANY OF SUTURES WERE BROKE DURING EACH PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814835&pc=GAM | 2210968-2020-01905 |  | Malfunction | ETHICON INC. | 2020-03-10 | GAM | VICRYL VIO 90CM M3.5::USP0 SGLE ARMED MH PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN SURGEON PERFORMED SINGLE KNOTS AND REALIZED THAT THE SUTURE BROKE AT THE SWAGE WHILE SUTURING. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. SUMMARY: AN EMPTY OPENED FOIL AND SIX UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITION OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF - SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS COMPLAINT WAS REPORTED WITH QUANTITY INVOLVED OF 7 SUTURES BROKEN DURING THIS SINGLE SURGERY. AND ON OUR QUESTION ¿CONFIRM SPECIFIC NUMBER OF DEVICES THAT FAILED DURING THIS REPORTED EVENT /PROCEDURE?¿ RESPONSE WAS ¿AT LEAST 10 EA. DURING DIFFERENT PROCEDURES. ¿ WERE THESE " DIFFERENT PROCEDURES" REPORTED TO ETHICON BEFORE? IF YES, PLEASE PROVIDE COMPLAINT PC NUMBERS. PLEASE SPECIFY DETAILS OF EACH PROCEDURE AND HOW MANY OF SUTURES WERE BROKE DURING EACH PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814907&pc=GAM | 2210968-2020-01906 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-03-10 | GAM | VICRYL VIO 90CM M3.5::USP0 SGLE ARMED MH PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN SURGEON PERFORMED SINGLE KNOTS AND REALIZED THAT THE SUTURE BROKE AT THE SWAGE WHILE SUTURING. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCES WERE IDENTIFIED. AN EMPTY OPENED FOIL AND SIX UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITION OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF - SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS COMPLAINT WAS REPORTED WITH QUANTITY INVOLVED OF 7 SUTURES BROKEN DURING THIS SINGLE SURGERY. AND ON OUR QUESTION ¿CONFIRM SPECIFIC NUMBER OF DEVICES THAT FAILED DURING THIS REPORTED EVENT /PROCEDURE?¿ RESPONSE WAS ¿AT LEAST 10 EA. DURING DIFFERENT PROCEDURES. ¿ WERE THESE " DIFFERENT PROCEDURES" REPORTED TO ETHICON BEFORE? IF YES, PLEASE PROVIDE COMPLAINT PC NUMBERS. PLEASE SPECIFY DETAILS OF EACH PROCEDURE AND HOW MANY OF SUTURES WERE BROKE DURING EACH PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814910&pc=GAM | 2210968-2020-01907 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-03-10 | GAM | VICRYL VIO 90CM M3.5::USP0 SGLE ARMED MH PLUS | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN SURGEON PERFORMED SINGLE KNOTS AND REALIZED THAT THE SUTURE BROKE AT THE SWAGE WHILE SUTURING. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE SUCCESSFULLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. AN EMPTY OPENED FOIL AND SIX UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED. DURING THE VISUAL INSPECTION OF UNOPENED SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, OR DAMAGED WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING AN INSTRON AND THE PULL FORCE WERE ABOVE THE MINIMUM REQUIREMENTS. PER THE CONDITION OF THE SAMPLES RECEIVED, NO PERFORMANCE PULL OFF - SUTURE NEEDLE WAS FOUND, AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS COMPLAINT WAS REPORTED WITH QUANTITY INVOLVED OF 7 SUTURES BROKEN DURING THIS SINGLE SURGERY. AND ON OUR QUESTION ¿CONFIRM SPECIFIC NUMBER OF DEVICES THAT FAILED DURING THIS REPORTED EVENT /PROCEDURE?¿ RESPONSE WAS ¿AT LEAST 10 EA. DURING DIFFERENT PROCEDURES. ¿ WERE THESE " DIFFERENT PROCEDURES" REPORTED TO ETHICON BEFORE? IF YES, PLEASE PROVIDE COMPLAINT PC NUMBERS. PLEASE SPECIFY DETAILS OF EACH PROCEDURE AND HOW MANY OF SUTURES WERE BROKE DURING EACH PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9809279&pc=GAM | 2210968-2020-01863 | 2020-02-11 | Malfunction | ETHICON INC. | 2020-03-09 | GAM | VICRYL RAP UND 70CM 3/0 PS-1 PRIME MP | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE NEEDLE BROKE OFF COMPLETELY. ANOTHER LIKE SUTURE WAS USED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE MADE TO OBTAIN THE DEVICE RETURN FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS REPORTED THAT " SUTURE NEEDLES BROKE OFF COMPLETELY¿ . PLEASE CLARIFY WHAT DOES " SUTURE NEEDLES BROKE OFF COMPLETELY¿ MEAN? WAS THE SUTURE DETACHED FROM THE NEEDLE DURING USE? OR THE NEEDLE BROKE IN PIECES? DEVICE RETURN STATUS/ FOLLOW UP?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9809280&pc=GAM | 2210968-2020-01864 | 2020-02-11 | Malfunction | ETHICON INC. | 2020-03-09 | GAM | VICRYL RAP UND 70CM 3/0 PS-1 PRIME MP | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE NEEDLE BROKE OFF COMPLETELY. ANOTHER LIKE SUTURE WAS USED. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE MADE TO OBTAIN THE DEVICE RETURN FOR EVALUATION. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. IT WAS REPORTED THAT " SUTURE NEEDLES BROKE OFF COMPLETELY¿ . PLEASE CLARIFY WHAT DOES " SUTURE NEEDLES BROKE OFF COMPLETELY¿ MEAN? WAS THE SUTURE DETACHED FROM THE NEEDLE DURING USE? OR THE NEEDLE BROKE IN PIECES? DEVICE RETURN STATUS/ FOLLOW UP?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9809294&pc=GAM | 2210968-2020-01861 |  | Malfunction | ETHICON INC. | 2020-03-09 | GAM | VICRYL CT VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT DURING EVALUATION, A LEGEND OF ¿NOT FOR HUMAN USE¿ PRINTING WAS OBSERVED ON THE BOX RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE, WHICH WAS RAISED TO ADDRESS THE EVENT REPORTED. THE NECESSARY ACTIONS HAVE BEEN TAKEN TO ADDRESS THE ISSUE. SUMMARY: AN OPEN BOX WITH TWELVE UNOPENED SAMPLES OF PRODUCT WAS RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE BOX, A LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED PRINTING ON BOX, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. THE SAMPLES INSIDE THE BOX WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO FINISH GOOD PRODUCT REQUIREMENTS. THE PRODUCTS MET ALL SPECIFICATIONS FOR USE ON PATIENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9809296&pc=GAM | 2210968-2020-01862 |  | Malfunction | ETHICON INC. | 2020-03-09 | GAM | VICRYL CT VIO 90CM M4 | Device Markings/Labelling Problem | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT DURING EVALUATION, A LEGEND OF ¿NOT FOR HUMAN USE¿ PRINTING WAS OBSERVED ON THE BOX RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). THE MANUFACTURING RECORD EVALUATION WAS PERFORMED AND IDENTIFIED A NONCONFORMANCE, WHICH WAS RAISED TO ADDRESS THE EVENT REPORTED. THE NECESSARY ACTIONS HAVE BEEN TAKEN TO ADDRESS THE ISSUE. SUMMARY: AN OPEN BOX WITH TWELVE UNOPENED SAMPLES OF PRODUCT WAS RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE BOX, A LEGEND OF "NOT FOR HUMAN USE" COULD BE OBSERVED PRINTING ON BOX, RESULTING IN INAPPROPRIATE INFORMATION AT THE CARDBOARD. THE SAMPLES INSIDE THE BOX WERE REVIEWED AND ALL THE INFORMATION WERE ACCORDING TO FINISH GOOD PRODUCT REQUIREMENTS. THE PRODUCTS MET ALL SPECIFICATIONS FOR USE ON PATIENTS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9794941&pc=GAM | 2210968-2020-01783 | 2020-01-24 | Malfunction | ETHICON INC. | 2020-03-05 | GAM | VICRYL VIO 90CM M4 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CESAREAN SURGICAL PROCEDURE ON (B)(6) 2020 AND THE SUTURE WAS USED. WHEN PERFORMING A STITCH OF HEMOSTASIS IN THE UTERUS OF THE PATIENT, THE SUTURE WIRE WAS DETACHED FROM THE NEEDLE. THE NEEDLE PIECE WAS REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). DATE SENT TO THE FDA: 03/05/2020. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT WAS BEING DONE WHEN THE NEEDLE PULLED-OFF (IN THE PACKAGE/ REMOVAL FROM PACKAGE / DURING PASSAGE THROUGH TISSUE¿)? WHEN PERFORMING A STITCH OF HEMOSTASIS IN THE UTERUS OF THE PATIENT. WERE X-RAYS TAKEN TO LOCATE THE NEEDLE? NO. WAS THE NEEDLE PIECE REMOVED FROM THE PATIENT DURING THE SAME PROCEDURE? YES. WAS THERE ANY PATIENT CONSEQUENCE OR INJURY? NO. WHAT IS THE CONDITION OF THE PATIENT? HOSPITALIZATION AND CARE NORMAL AFTER A CAESAREAN.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9783363&pc=GAM | 2210968-2020-01658 |  | Malfunction | ETHICON INC. | 2020-03-03 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). DATE SENT TO THE FDA: 03/03/2020. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY PATIENT EVENT PRIOR TO SYMPTOMS (COUGHING, FALLING, MOVING)? WHAT MEDICAL/SURGICAL INTERVENTION WAS PERFORMED FOR EACH OF THE PATIENT'S SYMPTOMS? WHAT WERE THE DATES OF THE SECOND PROCEDURES? CAN YOU DESCRIBE THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9775866&pc=GAM | 2210968-2020-01649 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY PATIENT EVENT PRIOR TO SYMPTOMS (COUGHING, FALLING, MOVING)? WHAT MEDICAL/SURGICAL INTERVENTION WAS PERFORMED FOR EACH OF THE PATIENT'S SYMPTOMS? WHAT WERE THE DATES OF THE SECOND PROCEDURES? CAN YOU DESCRIBE THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9776154&pc=GAM | 2210968-2020-01650 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WAS THERE ANY PATIENT EVENT PRIOR TO SYMPTOMS (COUGHING, FALLING, MOVING)? WHAT MEDICAL/SURGICAL INTERVENTION WAS PERFORMED FOR EACH OF THE PATIENT'S SYMPTOMS? WHAT WERE THE DATES OF THE SECOND PROCEDURES? CAN YOU DESCRIBE THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9776265&pc=GAM | 2210968-2020-01651 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: I DO NOT HAVE ANY ADDITIONAL INFORMATION TO ANSWER THE QUESTIONS BELOW. WAS THERE ANY PATIENT EVENT PRIOR TO SYMPTOMS (COUGHING, FALLING, MOVING)? WHAT MEDICAL/SURGICAL INTERVENTION WAS PERFORMED FOR EACH OF THE PATIENT'S SYMPTOMS? WHAT WERE THE DATES OF THE SECOND PROCEDURES? CAN YOU DESCRIBE THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9776517&pc=GAM | 2210968-2020-01653 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: REPORTER DOES NOT HAVE ANY ADDITIONAL INFORMATION TO ANSWER THE QUESTIONS BELOW. WAS THERE ANY PATIENT EVENT PRIOR TO SYMPTOMS (COUGHING, FALLING, MOVING)? WHAT MEDICAL/SURGICAL INTERVENTION WAS PERFORMED FOR EACH OF THE PATIENT'S SYMPTOMS? WHAT WERE THE DATES OF THE SECOND PROCEDURES? CAN YOU DESCRIBE THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9776661&pc=GAM | 2210968-2020-01654 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: NO ADDITIONAL INFORMATION IS KNOWN FOR THE BELOW: WHAT WAS THE DATE AND NAME OF THE PROCEDURE? WHAT TISSUE WAS THE SUTURE USED ON? WHAT DATE DID THE PATIENT EXPERIENCE DEHISCENCE? WAS ANY MEDICAL INTERVENTION INDICATED, IF SO WHAT? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777037&pc=GAM | 2210968-2020-01657 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE DATE AND NAME OF THE PROCEDURE? WHAT TISSUE WAS THE SUTURE USED ON? WHAT DATE DID THE PATIENT EXPERIENCE DEHISCENCE? WAS ANY MEDICAL INTERVENTION INDICATED, IF SO WHAT? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777159&pc=GAM | 2210968-2020-01659 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE DATE AND NAME OF THE PROCEDURE? WHAT TISSUE WAS THE SUTURE USED ON? WHAT DATE DID THE PATIENT EXPERIENCE DEHISCENCE? WAS ANY MEDICAL INTERVENTION INDICATED, IF SO WHAT? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777216&pc=GAM | 2210968-2020-01661 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE DATE AND NAME OF THE PROCEDURE? WHAT TISSUE WAS THE SUTURE USED ON? WHAT DATE DID THE PATIENT EXPERIENCE DEHISCENCE? WAS ANY MEDICAL INTERVENTION INDICATED, IF SO WHAT? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777270&pc=GAM | 2210968-2020-01663 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE DATE AND NAME OF THE PROCEDURE? WHAT TISSUE WAS THE SUTURE USED ON? WHAT DATE DID THE PATIENT EXPERIENCE DEHISCENCE? WAS ANY MEDICAL INTERVENTION INDICATED, IF SO WHAT? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777288&pc=GAM | 2210968-2020-01665 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE DATE AND NAME OF THE PROCEDURE? WHAT TISSUE WAS THE SUTURE USED ON? WHAT DATE DID THE PATIENT EXPERIENCE DEHISCENCE? WAS ANY MEDICAL INTERVENTION INDICATED, IF SO WHAT? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777314&pc=GAM | 2210968-2020-01666 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777333&pc=GAM | 2210968-2020-01667 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE DATE AND NAME OF THE PROCEDURE? WHAT TISSUE WAS THE SUTURE USED ON? WHAT DATE DID THE PATIENT EXPERIENCE DEHISCENCE? WAS ANY MEDICAL INTERVENTION INDICATED, IF SO WHAT? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777359&pc=GAM | 2210968-2020-01668 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). DATE SENT TO THE FDA: 03/02/2020. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE DATE AND NAME OF THE PROCEDURE? WHAT TISSUE WAS THE SUTURE USED ON? WHAT DATE DID THE PATIENT EXPERIENCE DEHISCENCE? WAS ANY MEDICAL INTERVENTION INDICATED, IF SO WHAT? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777663&pc=GAM | 2210968-2020-01671 |  | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | Wound Dehiscence | K022269 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE BROKE POST-OP CAUSING WOUND DEHISCENCE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # ==> (B)(4). DATE SENT TO THE FDA: 03/02/2020. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE DATE AND NAME OF THE PROCEDURE? WHAT TISSUE WAS THE SUTURE USED ON? WHAT DATE DID THE PATIENT EXPERIENCE DEHISCENCE? WAS ANY MEDICAL INTERVENTION INDICATED, IF SO WHAT? WHAT IS THE CURRENT CONDITION OF THE PATIENT?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777924&pc=GAM | 2210968-2020-01675 | 2020-02-04 | Malfunction | ETHICON INC. | 2020-03-02 | GAM | VICRYL*VLT PLUS 0 90CM 1-CT1 | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN THE SUTURE WAS USED, THE FILAMENT RUPTURED. THERE WERE PATIENT CONSEQUENCES. Manufacturer Narrative: PRODUCT COMPLAINT: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS USED TO COMPLETE THE PROCEDURE? WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9771625&pc=GAM | 2210968-2020-01603 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-02-28 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT WISDOM TEETH EXTRACTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS PULLING AWAY FROM THE SUTURE VERY EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY INVOLVED) THAT FAILED DURING THIS PROCEDURE? UNSURE SPECIFIC NUMBER OF DEVICES INVOLVED DURING THIS ONE PROCEDURE. THE REPORTER ONLY ADVISED THAT THERE WERE 6 IMPACTED SUTURE THROUGHOUT SEVERAL WISDOM TEETH EXTRACTIONS ON THE ONE DAY. NO HARM TO ANY PATIENT REPORTED. -STATUS OF PRODUCT RETURN: THE 6 IMPACTED SUTURES HAVE BEEN DISCARDED. NOTE: EVENTS REPORTED VIA MW 2210968-2020-01602, 2210968-2020-01604, 2210968-2020-01605, 2210968-2020-01606, 2210968-2020-01607.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9771629&pc=GAM | 2210968-2020-01602 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-02-28 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT WISDOM TEETH EXTRACTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS PULLING AWAY FROM THE SUTURE VERY EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY INVOLVED) THAT FAILED DURING THIS PROCEDURE? UNSURE SPECIFIC NUMBER OF DEVICES INVOLVED DURING THIS ONE PROCEDURE. THE REPORTER ONLY ADVISED THAT THERE WERE 6 IMPACTED SUTURE THROUGHOUT SEVERAL WISDOM TEETH EXTRACTIONS ON THE ONE DAY. NO HARM TO ANY PATIENT REPORTED. -STATUS OF PRODUCT RETURN: THE 6 IMPACTED SUTURES HAVE BEEN DISCARDED. NOTE: EVENTS REPORTED VIA MW 2210968-2020-01603, 2210968-2020-01604, 2210968-2020-01605, 2210968-2020-01606, 2210968-2020-01607.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9771630&pc=GAM | 2210968-2020-01604 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-02-28 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT WISDOM TEETH EXTRACTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS PULLING AWAY FROM THE SUTURE VERY EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY INVOLVED) THAT FAILED DURING THIS PROCEDURE? UNSURE SPECIFIC NUMBER OF DEVICES INVOLVED DURING THIS ONE PROCEDURE. THE REPORTER ONLY ADVISED THAT THERE WERE 6 IMPACTED SUTURE THROUGHOUT SEVERAL WISDOM TEETH EXTRACTIONS ON THE ONE DAY. NO HARM TO ANY PATIENT REPORTED. STATUS OF PRODUCT RETURN: THE 6 IMPACTED SUTURES HAVE BEEN DISCARDED. NOTE: EVENTS REPORTED VIA MW 2210968-2020-01602, 2210968-2020-01603, 2210968-2020-01605, 2210968-2020-01606, 2210968-2020-01607.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9771631&pc=GAM | 2210968-2020-01606 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-02-28 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT WISDOM TEETH EXTRACTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS PULLING AWAY FROM THE SUTURE VERY EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY INVOLVED) THAT FAILED DURING THIS PROCEDURE? UNSURE SPECIFIC NUMBER OF DEVICES INVOLVED DURING THIS ONE PROCEDURE. THE REPORTER ONLY ADVISED THAT THERE WERE 6 IMPACTED SUTURE THROUGHOUT SEVERAL WISDOM TEETH EXTRACTIONS ON THE ONE DAY. NO HARM TO ANY PATIENT REPORTED. STATUS OF PRODUCT RETURN: THE 6 IMPACTED SUTURES HAVE BEEN DISCARDED. NOTE: EVENTS REPORTED VIA MW 2210968-2020-01602, 2210968-2020-01603, 2210968-2020-01604, 2210968-2020-01605, 2210968-2020-01607.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9771634&pc=GAM | 2210968-2020-01607 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-02-28 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT WISDOM TEETH EXTRACTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS PULLING AWAY FROM THE SUTURE VERY EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY INVOLVED) THAT FAILED DURING THIS PROCEDURE? UNSURE SPECIFIC NUMBER OF DEVICES INVOLVED DURING THIS ONE PROCEDURE. THE REPORTER ONLY ADVISED THAT THERE WERE 6 IMPACTED SUTURE THROUGHOUT SEVERAL WISDOM TEETH EXTRACTIONS ON THE ONE DAY. NO HARM TO ANY PATIENT REPORTED. -STATUS OF PRODUCT RETURN: THE 6 IMPACTED SUTURES HAVE BEEN DISCARDED. NOTE: EVENTS REPORTED VIA MW 2210968-2020-01602, 2210968-2020-01603, 2210968-2020-01604, 2210968-2020-01605, 2210968-2020-01606.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9771635&pc=GAM | 2210968-2020-01605 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-02-28 | GAM | VICRYL RAP UND 75CM 5/0 FS-3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT WISDOM TEETH EXTRACTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE WAS PULLING AWAY FROM THE SUTURE VERY EASILY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY INVOLVED) THAT FAILED DURING THIS PROCEDURE? UNSURE SPECIFIC NUMBER OF DEVICES INVOLVED DURING THIS ONE PROCEDURE. THE REPORTER ONLY ADVISED THAT THERE WERE 6 IMPACTED SUTURE THROUGHOUT SEVERAL WISDOM TEETH EXTRACTIONS ON THE ONE DAY. NO HARM TO ANY PATIENT REPORTED. STATUS OF PRODUCT RETURN: THE 6 IMPACTED SUTURES HAVE BEEN DISCARDED. NOTE: EVENTS REPORTED VIA MW 2210968-2020-01602, 2210968-2020-01603, 2210968-2020-01604, 2210968-2020-01606, 2210968-2020-01607.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9771794&pc=GAM | 2210968-2020-01609 |  | Malfunction | ETHICON INC. | 2020-02-28 | GAM | VICRYL RAP UND 75CM 3-0 UX-25 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN DENTAL SURGERY ON AN UNKNOWN DATE AND SUTURE WAS USED. THE DENTIST REPORTS THAT THEY HAVE TRIED TO USE THESE SUTURES ON SEVERAL PATIENTS BUT THE SUTURES ARE DISSOLVING ON CONTACT WITH SALIVA AND NOT HOLDING THE WOUND TOGETHER. THE DENTIST RE-SUTURED THE WOUND USING AN ALTERNATIVE SUTURE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WERE THE SUTURES DISSOLVING UPON CONTACT WITH SALIVA AND NOT HOLDING TISSUE DURING SUTURING (INTRA-OP)? OR DID THE SUTURES DISSOLVED PREMATURELY POST-OP AND RE-SUTURING WAS REQUIRED? IF THE ISSUE OCCURRED POST-OP, PLEASE PROVIDE NUMBER OF POST-OP DAYS PATIENT RETURNED WITH ISSUE AND THE SYMPTOMS? IN HOW MANY PATIENTS DID THE SAME ISSUE OCCUR? FOR EACH PATIENT, PLEASE PROVIDE THE ABOVE INFORMATION INCLUDING EVENT DATES AND PRODUCT CODE/LOT NUMBER IF DIFFERENT, SEPARATELY. IS DEVICE AVAILABLE AND BEING RETURNED FOR EVALUATION? PLEASE VERIFY LOT NUMBER AJ196S AS IT APPEARS TO BE INVALID FOR THIS PRODUCT. THE CUSTOMER CONFIRMED THAT THIS WAS A HISTORIC COMPLAINT AND THE PRODUCT WAS NOT KEPT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. CAN YOU CLARIFY THE LOT NUMBER AS AJ196S? CAN YOU CLARIFY THAT THIS IS ETHICON SUTURE?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9754096&pc=GAM | 2210968-2020-01493 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-02-26 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K946173 |  | 1 | Event Description: IT WAS REPORTED BY VET THAT AN ANIMAL UNDERWENT AN OVARY-HYSTERECTOMY IN 2019 AND THE SUTURE WAS USED. 4 OR 5 DAY AFTER THE PROCEDURE, THE OVERLOCK/SUTURE OF THE ABDOMINAL WALL HAD BROKEN IN ITS LENGTH, THE NODES BEING INTACT. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9754115&pc=GAM | 2210968-2020-01494 |  | Malfunction | ETHICON INC. | 2020-02-26 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED BY VET THAT A CAT UNDERWENT AN OVARIECTOMY IN 2019 AND THE SUTURE WAS USED. 4 OR 5 DAY AFTER THE PROCEDURE, THE OVERLOCK/STITCHES OF THE ABDOMINAL WALL HAD BROKEN IN ITS LENGTH, THE NODES BEING INTACT. Manufacturer Narrative: DATE SENT TO THE FDA: 03/31/2020. ADDITIONAL SUMMARY: EIGHT UNOPENED SAMPLES OF PRODUCT WERE RETURNED FOR ANALYSIS. THE COMPLAINT SAMPLE WAS NOT RECEIVED FOR EVALUATION. DURING THE VISUAL INSPECTION OF THE SAMPLES, NO DEFECTS WERE FOUND ON THE PACKAGES. THE SAMPLES WERE OPENED, AND THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO DEFECTS, DAMAGED OR SUTURE BREAKAGE WERE OBSERVED. A FUNCTIONAL TEST WAS PERFORMED USING A INSTRON AND THE TENSILE STRENGTH FORCE WAS ABOVE THE MINIMUM REQUIREMENT. PER THE CONDITION OF THE REPRESENTATIVE SAMPLES, NO PERFORMANCE - BREAKAGE SUTURE WAS FOUND AND THE TESTED SAMPLES MET THE FINISHED GOODS REQUIREMENTS. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9756198&pc=GAW | 2210968-2020-01512 |  | Malfunction | ETHICON INC. | 2020-02-26 | GAW | VICRYL POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE CAT UNDERWENT AN ANIMAL PROCEDURE, LAPAROTOMY ALONG WITH REMOVAL OF AN ECTOPIC ABDOMINAL TESTICLE, IN 2019 AND THE SUTURE WAS USED. IT WAS REPORTED THAT ON 4TH OR 5TH DAY AFTER THE PROCEDURE, THE OVERLOCK/SUTURE OF THE ABDOMINAL WALL HAD BROKEN IN ITS LENGTH, THE NODES BEING INTACT. THE SUBCUTANEOUS OVERLOCK AND SKIN POINTS WERE GOOD. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9756974&pc=GAM | 2210968-2020-01523 | 2020-01-21 | Malfunction | ETHICON INC. | 2020-02-26 | GAM | VICRYL* UND 3-0 70CM (1)SC-20 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A UNILATERAL SIMPLE MASTECTOMY FOR DIAGNOSIS OF UNSPECIFIED MASS IN BREAST ON (B)(6) 2020 AND SUTURE WAS USED FOR CLOSURE. DURING THE PROCEDURE, WHEN THE CLOSURE WAS TO BE DONE WITH THE SUTURE, AT MAKING THE FIRST POINT AND TRY TO TAKE THE NEEDLE WITH THE NEEDLE HOLDER, ONLY THE SUTURE WAS OBSERVED. THE NEEDLE COULD NOT BE VISUALIZED. THE PATIENT WAS THOROUGHLY SEARCHED WITHOUT BEING ABLE TO SHOW IT, A PORTABLE CHEST RX WAS ORDERED SHOWING THE NEEDLE AT THE SURGICAL SITE. IT WAS POSSIBLE TO EXTRACT WITHOUT COMPLICATIONS, EVIDENCING DETACHMENT FROM THE BASE OF THE SUTURE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9744675&pc=GAM | 2210968-2020-01386 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-02-24 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE THE SUTURE BROKE FROM THE SWAGE END. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). DATE SENT TO THE FDA: 2/24/2020. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9740203&pc=GAM | 2210968-2020-01366 | 2020-01-21 | Malfunction | ETHICON INC. | 2020-02-21 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 3-0 UND | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN OB-GYN SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS ALSO REPORTED THAT THE SUTURE WAS BROKEN SHORTLY AFTER STARTING SUTURING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE AND NO NON-CONFORMANCE'S RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT WAS USED TO COMPLETE THE PROCEDURE? NO FURTHER INFORMATION IS AVAILABLE. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? NO FURTHER INFORMATION IS AVAILABLE. DEVICE RETURN STATUS/FOLLOW UP: WE REGULARLY CONTACT WITH SALES REP ABOUT THE DEVICE RETURNING. NO FURTHER INFORMATION WILL BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9740766&pc=GAM | 2210968-2020-01367 | 2020-01-20 | Malfunction | ETHICON INC. | 2020-02-21 | GAM | VICRYL RAPIDE SUTURE 36"(90CM) 3-0 UND | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN OB-GYN SURGERY ON (B)(6) 2020 AND THE SUTURE WAS USED. IT WAS ALSO REPORTED THAT WHEN OPENING THE PACKAGE, IT WAS FOUND THAT THE SUTURE HAD BEEN DETACHED FROM THE NEEDLE. IT WAS NOT A CONTROL RELEASE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: PRODUCT COMPLAINT (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WAS THE PATIENT¿S POST-OPERATIVE COURSE CHANGED AS A RESULT OF THE EXTENDED PROCEDURE? IF YES, PLEASE EXPLAIN. NO FURTHER INFORMATION IS AVAILABLE. PRODUCT RETURN STATUS: THE DEVICE WILL BE SHIPPED. NO FURTHER INFORMATION WILL BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9726510&pc=GAM | 2210968-2020-01261 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-02-19 | GAM | VICRYL ABS SUR SUT USP SYN | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE IN 2020 AND SUTURE WAS USED. THE SUTURE KEPT BREAKING WHILE USING THE SAME. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 02/18/2020. DEVICE EVALUATION SUMMARY: COMPLAINT SAMPLE NOT RECEIVED FOR INVESTIGATION. RETAIN SAMPLE INVESTIGATION: FIVE RETAIN SAMPLES OF INCIDENT CODES NW2341 AND LOT NUMBER T9010 WERE RETRIEVED FOR ANALYSIS. THE PRIMARY PACKS OF RETAIN SAMPLES WERE VISUALLY INSPECTED FOR ATTRIBUTE DEFECTS LIKE PSEUDO SEAL, PIN HOLE, PRODUCT IN SEAL, PRESS MARKS ON PACK BUT NO SUCH DEFECTS WERE OBSERVED. THE PRIMARY PACKS WERE OPENED, AND SUTURES AND NEEDLES WERE FOUND INTACT. THE SUTURES WERE PHYSICALLY INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE KINKS, WEAK SPOTS, BROKEN PIECE, FRAY, BRITTLENESS, DETACHED NEEDLES BUT NO SUCH DEFECTS WERE OBSERVED. THE NEEDLES WERE INSPECTED FOR ANY ATTRIBUTE DEFECTS LIKE BEND, CRACKED BARREL, FINS BUT NO SUCH DEFECTS WERE OBSERVED. THE RETAIN SAMPLES WERE TESTED FOR KNOT PULL TENSILE STRENGTH TEST AND FOUND TO MEET THE SPECIFICATION. BATCH RECORD REVIEW: AS A PART OF FURTHER INVESTIGATION BATCH MANUFACTURING RECORD WAS REVIEWED. THE BATCH MANUFACTURING RECORD WAS REVIEWED FOR ANY PROCESS DEVIATION, BUT NO DEVIATION WAS OBSERVED. FINISHED GOOD RECORD WAS REVIEWED FOR TENSILE STRENGTH VALUE AND NEEDLE PULL-OFF VALUE AT RELEASE AND FOUND TO MEET THE SPECIFICATION. FROM THE BATCH RECORD REVIEW, IT IS EVIDENT THAT THERE WAS NO ISSUE RELATED TO THE SUTURE QUALITY AND PROCESSING OF THIS INCIDENT LOT. KNOT PULL TENSILE STRENGTH TEST WAS PERFORMED OVER FIVE RETAIN SAMPLES TO CHECK THE BEHAVIOR OF THE SUTURE MATERIAL. THE AVERAGE KNOT PULL TENSILE STRENGTH VALUE OF THE RETAIN SAMPLE WAS FOUND TO MEETS THE USP AVERAGE KNOT PULL TENSILE STRENGTH REQUIREMENT. ALL THE INDIVIDUAL AS WELL AS AVERAGE KNOT PULL TENSILE STRENGTH VALUES WERE FOUND TO MEET THE USP SPECIFICATION REQUIREMENTS. THIS ANALYSIS SHOWS THAT THERE WAS NO ISSUE RELATED TO PROCESSING OF THE LOT. ALL THE VALUES ARE WITHIN THE CONTROL LIMITS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9720765&pc=GAM | 2210968-2020-01245 | 2019-12-28 | Malfunction | ETHICON INC. | 2020-02-18 | GAM | VICRYL VIO 90CM 0 MH | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CAESARIAN SECTION PROCEDURE ON (B)(6) 2019 AND THE SUTURE WAS USED. IT WAS REPORTED THAT A DURING SURGERY, IN THE FIRST STITCH, THE SUTURE WAS EASILY BROKEN. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 02/18/2020. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9720823&pc=GAM | 2210968-2020-01246 | 2019-07-19 | Malfunction | ETHICON INC. | 2020-02-18 | GAM | VICRYL CT VIO 75CM M4 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A TOTAL HYSTERECTOMY AND A BILATERAL SALPINGECTOMY PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE USED TO CLOSE THE VAGINA IN THREAD REMOVED FROM THE NEEDLE. THE WIRE WAS CHANGED. RISK TO LOSE THE NEEDLE ON THE TISSUES. THE NEEDLE IS NOT FALLEN INSIDE THE PATIENT BODY. THERE WAS NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: LOT #: MM6704. WHAT DO YOU MEAN BY "RISK TO LOSE THE NEEDLE ON THE TISSUES"? THE CUSTOMER MEANS THAT WHEN THE NEEDLE PULLED OFF FROM THE NEEDLE THERE IS A RISK THAT THE NEEDLE STAY INSIDE THE PATIENT. WAS THE NEEDLE REMOVED FROM THE PATIENT BODY IN INITIAL PROCEDURE ITSELF? THE NEEDLE IS NOT FALLEN INSIDE THE PATIENT BODY. IF YES, WHERE AND HOW WAS THE NEEDLE LOCATED? N/A. IF NEEDLE STILL REMAINS, ANY PLAN TO REMOVE THE NEEDLE POST-OP FROM PATIENT BODY?- ANY MEDICAL /SURGICAL INTERVENTION DONE ON PATIENT POST-OP? N/A. ANY PATIENT CONSEQUENCES/AE OUTCOME AS A RESULT OF THE EVENT REPORT? THE OPERATIVE FOLLOW-UP WAS SIMPLE, NO PATIENT CONSEQUENCES REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9712145&pc=GAM | 2210968-2020-01194 | 2019-10-24 | Malfunction | ETHICON INC. | 2020-02-14 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. NEEDLE PULLED OFF DURING THE USE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 02/14/2020. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9696252&pc=GAM | 2210968-2020-01082 |  | Malfunction | ETHICON INC. | 2020-02-11 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE OFF FROM THE NEEDLE AT THE CONNECTION POINT. THE SUTURE DID NOT BREAK IN THE FIBER. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL PATIENT EVENTS. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IS THE PRODUCT CODE AND LOT NUMBER AVAILABLE? NO PHYSICAL PRODUCT WAS GIVEN FOR INSPECTION/PRODUCT INQUIRY. NO SPECIFIC PRODUCT CODE OR LOT NUMBER WAS PROVIDED. THE CUSTOMER ONLY WANTED TO PROVIDE HIS PERSONAL FEEDBACK ON HIS EXPERIENCE WITH OUR SUTURE. STATUS OF PRODUCT RETURN: NO PHYSICAL PRODUCT WILL BE RETURNED FOR INSPECTION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9666051&pc=GAM | 2210968-2020-00920 | 2020-01-13 | Malfunction | ETHICON INC. | 2020-02-04 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NATURAL PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE WHEN SEWING THE PERINEAL SKIN. CHANGED TO ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9641677&pc=GAM | 2210968-2020-00764 | 2020-01-07 | Malfunction | ETHICON INC. | 2020-01-29 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT NATURAL BIRTH ON AN UNKNOWN DATE AND SUTURE WAS USED. THE SUTURE BROKE AT THE TIME OF KNOTTING AFTER SEWING THE PERINEAL SKIN FOR THE FIRST STITCH IN A NATURAL BIRTH SURGERY DURING WHICH THE SUTURE WAS CLAMPED BY ANY DEVICE. CHANGED TO ANOTHER DEVICE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9636230&pc=GAM | 2210968-2020-00742 | 2020-01-08 | Malfunction | ETHICON INC. | 2020-01-28 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NATURAL CHILD BIRTH ON (B)(6) AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE WHEN DRAWING THE SUTURE TO SEW THE PERINEAL SKIN. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9636259&pc=GAM | 2210968-2020-00743 | 2020-01-08 | Malfunction | ETHICON INC. | 2020-01-28 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NATURAL CHILD BIRTH ON (B)(6) 2020 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AT THE TIME OF KNOTTING AFTER SEWING THE PERINEAL SKIN FOR THE FIRST STITCH. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION COULD BE PROVIDED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A PHOTO HAS BEEN RECEIVED FOR EVALUATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9638702&pc=GAM | 2210968-2020-00758 | 2020-01-08 | Malfunction | ETHICON INC. | 2020-01-28 | GAM | VICRYL* VLT PLUS 0 70CM (1)SH | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. AT THE TIME OF USING THE SUTURE, RUPTURE OF THE FILAMENTS OCCURS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9632878&pc=GAM | 2210968-2020-00718 | 2020-01-07 | Malfunction | ETHICON INC. | 2020-01-27 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED ON THE VAGINAL MUCOSA AND SKIN. DURING THE PROCEDURE, THE SUTURE WAS BREAKING FROM MIDDLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT T8053, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS USED TO COMPLETE THE PROCEDURE? NW2777, T 9024 WERE USED TO COMPLETE THE SURGERY. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? VAGINAL MUCOSA AND SKIN WERE BEING SUTURED WHEN IT BROKE. HOW MANY OF THE TOTAL QUANTITY WERE ACTUALLY INVOLVED FOR THE REPORTED ISSUE? TOTAL QUANTITY INVOLVED WERE 3 FOILS AND REMAINING 9 FOILS WERE RETURNED WHICH WERE SENT TO YOU FOR INVESTIGATION PURPOSE. NOTE: EVENTS REPORTED IN 2210968-2020-00719, AND 2210968-2020-00720.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9632882&pc=GAM | 2210968-2020-00720 | 2020-01-07 | Malfunction | ETHICON INC. | 2020-01-27 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED ON THE VAGINAL MUCOSA AND SKIN. DURING THE PROCEDURE, THE SUTURE WAS BREAKING FROM MIDDLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT T8053, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS USED TO COMPLETE THE PROCEDURE? NW2777, T 9024 WERE USED TO COMPLETE THE SURGERY. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? VAGINAL MUCOSA AND SKIN WERE BEING SUTURED WHEN IT BROKE. HOW MANY OF THE TOTAL QUANTITY WERE ACTUALLY INVOLVED FOR THE REPORTED ISSUE? TOTAL QUANTITY INVOLVED WERE 3 FOILS AND REMAINING 9 FOILS WERE RETURNED WHICH WERE SENT TO YOU FOR INVESTIGATION PURPOSE. NOTE: EVENTS REPORTED IN 2210968-2020-00718, AND 2210968-2020-00719.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9632889&pc=GAM | 2210968-2020-00719 | 2020-01-07 | Malfunction | ETHICON INC. | 2020-01-27 | GAM | VICRYL RAPIDE ABS SUR SUT SYN | Break; Tear, Rip or Hole in Device Packaging | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED ON THE VAGINAL MUCOSA AND SKIN. DURING THE PROCEDURE, THE SUTURE WAS BREAKING FROM MIDDLE. ANOTHER LIKE DEVICE WAS USED TO COMPLETE THE PROCEDURE WITH NO ADVERSE PATIENT CONSEQUENCES. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT T8053, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS USED TO COMPLETE THE PROCEDURE? NW2777, T 9024 WERE USED TO COMPLETE THE SURGERY. WHAT TISSUE WAS BEING APPROXIMATED OR SUTURED WHEN IT BROKE? VAGINAL MUCOSA AND SKIN WERE BEING SUTURED WHEN IT BROKE. HOW MANY OF THE TOTAL QUANTITY WERE ACTUALLY INVOLVED FOR THE REPORTED ISSUE? TOTAL QUANTITY INVOLVED WERE 3 FOILS AND REMAINING 9 FOILS WERE RETURNED WHICH WERE SENT TO YOU FOR INVESTIGATION PURPOSE. NOTE: EVENTS REPORTED IN 2210968-2020-00718, AND 2210968-2020-00720.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633401&pc=GAM | 2210968-2020-00725 | 2019-12-22 | Malfunction | ETHICON INC. | 2020-01-27 | GAM | VICRYL CT VIO 70CM M3.5 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE INTO 3 PIECES AFTER 2 MINUTES OF USAGE. THE PROCEDURE WAS PROLONGED FOR UNKNOWN AMOUNT OF TIME. THE DEVICE WAS REPLACED WITH ANOTHER. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9628689&pc=GAM | 2210968-2020-00690 | 2020-01-01 | Malfunction | ETHICON INC. | 2020-01-24 | GAM | VICRYL RAP UND 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT POSTPARTUM PERINEAL SUTURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE PULLED OFF THE STRAND. THE NEEDLE WAS RETRIEVED EASILY FROM THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9628957&pc=GAM | 2210968-2020-00693 |  | Malfunction | ETHICON INC. | 2020-01-24 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT ELECTIVE CESAREAN SECTION ON AN UNKNOWN DATE AND SUTURE WAS USED. DURING THE CLOSURE OF UTERINE INCISION WITH SUTURE, THE TIP OF THE NEEDLE (ABOUT 2MM) BROKE OFF. EVERY EFFORT WAS MADE TO FIND AND RETRIEVE IT. PATIENT HAD X-RAY WHICH DID NOT SHOW THE NEEDLE TIP, POSSIBLY BECAUSE IT IS TOO SMALL TO SEE IT ON X-RAY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9623180&pc=GAM | 2210968-2020-00637 | 2020-01-05 | Malfunction | ETHICON INC. | 2020-01-23 | GAM | VICRYL RAP UND 45CM M2 3/0 X-1 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2020 AND SUTURE WAS USED. THE SUTURE GETS RIPPED AND THEN BREAKS. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PEBBBHS0 NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9590881&pc=GAM | 2210968-2020-00389 | 2019-12-20 | Malfunction | ETHICON INC. | 2020-01-15 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE FOR PERINEAL LACERATION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AT THE TIME OF KNOTTING. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AL2391, AND NO NON-CONFORMANCE WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9590885&pc=GAM | 2210968-2020-00387 | 2019-12-20 | Malfunction | ETHICON INC. | 2020-01-15 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE FOR PERINEAL LACERATION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE WHEN SEWING AT THE TIME OF KNOTTING. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT (B)(4), AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9590924&pc=GAM | 2210968-2020-00388 | 2019-12-20 | Malfunction | ETHICON INC. | 2020-01-15 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE FOR PERINEAL LACERATION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AT THE TIME OF DRAWING AFTER PIERCING. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT AL2391, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9590935&pc=GAM | 2210968-2020-00390 | 2019-12-20 | Malfunction | ETHICON INC. | 2020-01-15 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE FOR PERINEAL LACERATION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE AT THE TIME OF KNOTTING. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT (B)(4), AND NO NON-CONFORMANCES WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9591113&pc=GAM | 2210968-2020-00393 | 2019-12-19 | Malfunction | ETHICON INC. | 2020-01-15 | GAM | VICRYL RAP UD2/3-0 75CM 1A.PS2 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE FOR PERINEAL LACERATION ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE BROKE WHEN DRAWING AFTER PIERCING. CHANGED ANOTHER ONE TO COMPLETE THE SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT (B)(4), AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9591784&pc=GAM | 2210968-2020-00401 | 2019-03-04 | Malfunction | ETHICON INC. | 2020-01-15 | GAM | VICRYL CT VIO 75CM M1.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE PULLED OFF FROM THE THREAD AT THE END OF THE SUTURE. ANOTHER LIKE DEVICE WAS USED. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT MMBBDQR0, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586101&pc=GAM | 2210968-2020-00352 | 2019-01-01 | Malfunction | ETHICON INC. | 2020-01-14 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Wound Dehiscence; Scarring | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ORTHOPEDIC PROCEDURE IN 2019 AND SUTURE WAS USED. THE SUBCUTANEOUS SUTURE HAD NOT RESORBED IN ITS ENTIRETY CAUSING A DELAY FOR THE SCARRING. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ANY MEDICAL TREATMENT PROVIDED? IF YES, PLEASE PROVIDE MEDICINE NAME, STRENGTH AND DOSE WAS ANY QUALITY DEFICIENCY/NON-CONFORMANCE NOTED WITH THE SUTURE BEFORE USE, DURING USE, DURING POST-OP EXAMINATION OR DURING RE-OPERATION IF PERFORMED? WAS ANY SURGICAL INTERVENTION PERFORMED SPECIALLY RE-SUTURING? EXPLAIN. WAS DEHISCENCE NOTED? TISSUE ON WHICH USED? NAME OF PROCEDURE INITIAL PROCEDURE DATE? LOT NUMBER? WHAT IS THE PATIENT¿S CURRENT HEALTH STATUS AND CONDITION?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586166&pc=GAM | 2210968-2020-00354 | 2019-11-01 | Malfunction | ETHICON INC. | 2020-01-14 | GAM | VICRYL CT VIO 70CM M3 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN LAPAROSCOPIC PROCEDURE IN (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE IN THE ABDOMEN DURING LAPAROSCOPIC SURGERY. THEY FINALLY FOUND THE NEEDLE IN THE ABDOMEN. THE PROCEDURE WAS EXTENDED FOR AN HOUR TO X-RAY THE PATIENT TO LOCATE THE NEEDLE. THE NEEDLE PIECES WERE REMOVED DURING THE PROCEDURE. NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DATE SENT TO THE FDA: 01/14/2020. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. THE PHOTO UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WERE THERE ANY UNEXPECTED OUTCOMES OR COMPLICATIONS AS A RESULT OF THE PROLONGED SURGERY TIME?. TO MY KNOWLEDGE THERE WERE NONE OF THE ABOVE, OTHER THAN THE FACT THAT THE PATIENT NEEDED TO BE UNDER LONGER TIME OF ANESTHESIA. WAS THE PATIENT TREATMENT ALTERED IN ANYWAY DUE TO THE PROLONGED SURGERY TIME? IF YES, PLEASE EXPLAIN TO MY KNOWLEDGE THE PROLONGED SURGERY TIME DID NOT ALTER THE TREATMENT, OTHER THAN THAT THE PATIENT NEEDED TO BE IN ANESTHESIA FOR A LONGER TIME.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9581612&pc=GAM | 2210968-2020-00312 | 2019-12-16 | Malfunction | ETHICON INC. | 2020-01-13 | GAM | VICRYL CT UD 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING USE. NEEDLE BROKE IN FRONT THIRD. SUTURE ALSO BROKE. SURGERY COMPLETED WITH SAME LIKE PRODUCT. NO DELAY IN SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY INVOLVED) THAT FAILED DURING THIS PROCEDURE? 2. DEVICES CAPTURED IN MW 2210968-2020-00312 AND 2210968-2020-00313.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9581619&pc=GAM | 2210968-2020-00313 | 2019-12-16 | Malfunction | ETHICON INC. | 2020-01-13 | GAM | VICRYL CT UD 90CM M4 | Break | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A CAESAREAN SECTION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE BROKE DURING USE. NEEDLE BROKE IN FRONT THIRD. SUTURE ALSO BROKE. SURGERY COMPLETED WITH SAME LIKE PRODUCT. NO DELAY IN SURGERY. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS THE SPECIFIC NUMBER OF DEVICES (TOTAL QTY INVOLVED) THAT FAILED DURING THIS PROCEDURE? 2. DEVICES CAPTURED IN MW 2210968-2020-00312 AND 2210968-2020-00313.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9582053&pc=GAM | 2210968-2020-00318 | 2019-12-16 | Malfunction | ETHICON INC. | 2020-01-13 | GAM | VICRYL RAPIDE POLYGLACTIN 910 SUTURE | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN SURGERY ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE SUTURE WAS BROKEN DURING CONTINUOUS SUTURING ON THE PERINEUM. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9579358&pc=GAM | 2210968-2020-00277 | 2019-12-15 | Malfunction | ETHICON INC. | 2020-01-11 | GAM | VICRYL RAP UC UD 45CM M1 | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND THE SUTURE WAS USED. IT WAS REPORTED THAT WHEN SUTURING, THE SUTURE DISSOLVED/BROKE/FRAGMENTED INTO PIECES. THE KNOT WHERE UNTYING BECAUSE OF THE SUTURE MATERIAL DISSOLVED/BROKE/FRAGMENTED INTO PIECES. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: 1. WHAT IS MEANT BY 'THE SUTURE DISSOLVES WHEN THE CUSTOMER SUTURES' ? THE SUTURE DISSOLVED/BROKE/FRAGMENTED INTO PIECES WHEN SUTURING. 2. WERE THE KNOT UNTYING DURING USE? THE KNOT WHERE UNTYING BECAUSE OF THE SUTURE MATERIAL DISSOLVED/BROKE/FRAGMENTED INTO PIECES. 3. IF THERE ARE OTHER ISSUE PLEASE EXPLAIN? 4. LOT NUMBER?
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9573408&pc=GAM | 2210968-2020-00236 | 2019-09-01 | Malfunction | ETHICON INC. | 2020-01-10 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: HOLD FOR MRD 06/17/20. IT WAS REPORTED THAT A DOG UNDERWENT AN OVARIO HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ADDITIONAL DEVICE EVALUATION SUMMARY PHOTO: A PICTURE WAS PROVIDED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURE, IT WAS NOTED A WINDING FORMER WITH A SUTURE AND THE NEEDLE APPEARS TO BE DETACHED OF PRODUCT CODE JV586, LOT EH8GRSM0. ALSO, THREE NEEDLES APPEAR DETACHED AND THREE WINDING FORMERS WITH A SUTURE EACH, FOR TWO OF THESE SAMPLES THE PRODUCT CODE IS UNKNOWN, AND ONE SAMPLE BELONGS TO PRODUCT CODE JV453. HOWEVER, NO CONCLUSION COULD BE REACH ON HOW THIS HAPPEN AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. Manufacturer Narrative: ADDITIONAL SUMMARY: ONE OPENED FOLDER WITH A DETACHED NEEDLE AND THREE SMALL SUTURE PIECES OF PRODUCT WERE RECEIVED FOR ANALYSIS. DURING VISUAL INSPECTION OF THE DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE WAS EXAMINED AND REMNANT SUTURE WAS OBSERVED. THE SUTURE PIECES HAVE BEGUN WITH THE PROCESS OF DEGRADATION AND THIS CAUSED THAT THE NEEDLE WAS DETACHED OF THE STRAND. THE TIME OF EXPOSURE OF SUTURES TO THE ENVIRONMENT COULD NOT BE DETERMINED. ADDITIONALLY, THE FOIL WAS NOT RETURNED TO DETERMINE THE ASSIGNABLE CAUSE. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE SAMPLE CONDITION, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT Manufacturer Narrative: PRODUCT COMPLAINT#: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. UPON REVIEW OF PHOTO 2 IT APPEARS THAT THE SUTURE THREAD IS STILL IN THE PACKAGING, SO PLEASE CLARIFY WHEN DID THE SUTURE PULL OFF DURING USE ON ANIMAL OR IN THE PACKAGE OR DISPENSING (BEFORE USE ON ANIMAL) PLEASE ALSO VERIFY QUANTITY INVOLVED IN THIS PROCEDURE DEVICE RETURN STATUS/FOLLOW UP A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9573434&pc=GAM | 2210968-2020-00237 | 2019-09-01 | Malfunction | ETHICON INC. | 2020-01-10 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN OOPHORECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: DATE SENT TO THE FDA: 02/21/2020. ADDITIONAL SUMMARY: THREE PICTURES WERE PROVIDED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, IT WAS NOTED ONE OPEN BOX, THREE FOLDERS WITH ONE INSERT NEEDLE APPEAR DETACHED AND ONE EMPTY FOLDER OF PRODUCT. HOWEVER, NO CONCLUSION COULD BE REACH ON HOW THIS HAPPEN AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. Manufacturer Narrative: DATE SENT TO THE FDA: 03/05/2020. ADDITIONAL SUMMARY: ONE EMPTY OPENED FOIL AND ONE DETACHED NEEDLE OF PRODUCT WERE RECEIVED FOR ANALYSIS. DURING VISUAL INSPECTION OF THE DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE OF THE NEEDLE WAS EXAMINED, AND NO SUTURE REMNANT WAS NOTED. THE SUTURE WAS NOT RETURNED TO DETERMINE THE ASSIGNABLE CAUSE. IN ADDITION, SLIGHTLY MARKS WERE FOUND ON THE BODY NEEDLE APPEARS TO BE BY USE OF SURGICAL INSTRUMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE VISUAL INSPECTION, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9573437&pc=GAM | 2210968-2020-00240 | 2019-09-01 | Malfunction | ETHICON INC. | 2020-01-10 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient; No Clinical Signs, Symptoms or Conditions | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN OOPHORECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: ADDITIONAL SUMMARY: THREE PICTURES WERE PROVIDED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, IT WAS NOTED ONE OPEN BOX, THREE FOLDERS WITH ONE INSERT NEEDLE APPEAR DETACHED AND ONE EMPTY FOLDER OF PRODUCT. HOWEVER, NO CONCLUSION COULD BE REACH ON HOW THIS HAPPEN AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. Manufacturer Narrative: ADDITIONAL SUMMARY: ONE EMPTY OPENED FOIL AND ONE DETACHED NEEDLE OF PRODUCT WERE RECEIVED FOR ANALYSIS. DURING VISUAL INSPECTION OF THE DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE OF THE NEEDLE WAS EXAMINED, AND NO SUTURE REMNANT WAS NOTED. THE SUTURE WAS NOT RETURNED TO DETERMINE THE ASSIGNABLE CAUSE. IN ADDITION SLIGHTLY MARKS WERE FOUND ON THE BODY NEEDLE APPEARS TO BE BY USE OF SURGICAL INSTRUMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE VISUAL INSPECTION, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT Manufacturer Narrative: PRODUCT COMPLAINT #(B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9573440&pc=GAM | 2210968-2020-00239 | 2019-09-01 | Malfunction | ETHICON INC. | 2020-01-10 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN OOPHORECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9573441&pc=GAM | 2210968-2020-00238 | 2019-09-01 | Malfunction | ETHICON INC. | 2020-01-10 | GAM | VICRYL CT VIO 75CM M3.5 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN OOPHORECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: ADDITIONAL SUMMARY: THREE PICTURES WERE PROVIDED FOR ANALYSIS. UPON VISUAL INSPECTION OF THE PICTURES, IT WAS NOTED ONE OPEN BOX, THREE FOLDERS WITH ONE INSERT NEEDLE APPEAR DETACHED AND ONE EMPTY FOLDER OF PRODUCT. HOWEVER, NO CONCLUSION COULD BE REACH ON HOW THIS HAPPEN AS THE SAMPLE WAS NOT RETURNED FOR ANALYSIS. Manufacturer Narrative: ADDITIONAL SUMMARY: ONE EMPTY OPENED FOIL AND ONE DETACHED NEEDLE OF PRODUCT WERE RECEIVED FOR ANALYSIS. DURING VISUAL INSPECTION OF THE DETACHED NEEDLE, THE SWAGE AND ATTACHMENT AREA WERE NOTED TO BE AS EXPECTED. THE BARREL HOLE OF THE NEEDLE WAS EXAMINED, AND NO SUTURE REMNANT WAS NOTED. THE SUTURE WAS NOT RETURNED TO DETERMINE THE ASSIGNABLE CAUSE. IN ADDITION SLIGHTLY MARKS WERE FOUND ON THE BODY NEEDLE APPEARS TO BE BY USE OF SURGICAL INSTRUMENT. THE MANUFACTURING RECORDS WERE REVIEWED, AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS BATCH. ACCORDING TO THE VISUAL INSPECTION, IT COULD NOT BE DETERMINED WHAT MAY HAVE CAUSED THE REPORTED INCIDENT Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCE'S WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9573342&pc=GAM | 2210968-2020-00232 | 2019-12-09 | Malfunction | ETHICON INC. | 2020-01-09 | GAM | VICRYL POLYGLACTIN 910 SUTURE | Material Separation | No Consequences Or Impact To Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF FROM THE THREAD DURING THE USE. IT IS REMAINED INSIDE THE PATIENT. THE ACTIONS TAKEN WERE TO SEARCH AND REMOVE THE NEEDLE. NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9573376&pc=GAM | 2210968-2020-00235 | 2019-09-01 | Malfunction | ETHICON INC. | 2020-01-09 | GAM | VICRYL CT VIO 75CM M3 | Material Separation | No Consequences Or Impact To Patient | K022269 |  | 1 | Event Description: IT WAS REPORTED THAT A DOG UNDERWENT AN OVARIO HYSTERECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE PULLED OFF AT THE LEVEL OF WHITE LINE LIGATURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT, AND NO NON-CONFORMANCES WERE IDENTIFIED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9557947&pc=GAM | 2210968-2020-00152 |  | Malfunction | ETHICON INC. | 2020-01-07 | GAM | VICRYL RAP UND 45CM 4/0 FS-2S | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON AN UNKNOWN DATE AND SUTURE WAS USED. THE NEEDLE WAS INSERTED INTO THE TISSUE, THE SUTURE CAME OFF FROM THE NEEDLE. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE V2920H; MPBBSWT0 NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. THE PRODUCT UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE PRODUCT EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9558608&pc=GAM | 2210968-2020-00163 |  | Malfunction | ETHICON INC. | 2020-01-07 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Appropriate Term/Code Not Available | Pain; No Code Available | K022269 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT THAT THEY UNDERWENT A NERVE DECOMPRESSION PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. THE SUTURE HAVE NOT BEEN FULLY ABSORBED AFTER SEVEN MONTHS AND ONE IS PARTIALLY STICKING OUT AND IS CAUSING HER DISCOMFORT AND PAIN. NO INFORMATION REGARDING TREATMENT WAS INDICATED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ADDITIONAL UNKNOWN DEVICE REPORTED IN MW 2210968-2020-00164.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559269&pc=GAM | 2210968-2020-00180 | 2019-12-06 | Malfunction | ETHICON INC. | 2020-01-07 | GAM | VICRYL RAP UND 70CM 3/0 SH PLUS | Break | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A BILATERAL OVARIAN CYSTECTOMY ON (B)(6) 2019 AND SUTURE WAS USED. THE NEEDLE FELL OFF THE SUTURE WHEN OPENED INTO THE STERILE FIELD AND WAS RETRIEVED. NO DELAY TO PROCEDURE. NO REPORTED PATIENT CONSEQUENCES. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCE'S WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9561806&pc=GAM | 2210968-2020-00195 | 2019-12-11 | Malfunction | ETHICON INC. | 2020-01-07 | GAM | VICRYL RAP UND 45CM 4/0 FS-2S | Material Separation | No Consequences Or Impact To Patient | K033746 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND THE SUTURE WAS USED. IT WAS REPORTED THAT THE SUTURE STRAND PULLED OFF NEEDLE AS NEEDLE HOLDER WAS USED TO GET SUTURE READY FOR SKIN CLOSURE. ANOTHER LIKE SUTURE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO PATIENT CONSEQUENCES REPORTED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO NON-CONFORMANCES WERE IDENTIFIED. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT THE LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9554613&pc=GAM | 2210968-2020-00124 | 2019-12-10 | Malfunction | ETHICON INC. | 2020-01-06 | GAM | VICRYL PLUS CT VIO 8X45CM M4 | Break | No Consequences Or Impact To Patient | K032420 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNKNOWN PROCEDURE ON (B)(6) 2019 AND SUTURE WAS USED. DURING THE PROCEDURE, THE NEEDLE BROKE INSIDE THE PATIENT'S VAGINAL CAVITY. THE SCRUB NURSE HAD THE SMALL END OF THE NEEDLE WITH THE THREAD ATTACHED, WHILE THE SURGEON CONFIRMED THE SHARP LONG END OF THE NEEDLE WAS INSIDE THE PATIENTS VAGINAL CAVITY. THE END OF THE NEEDLE WAS LOCATED BY THE SURGEON WITHOUT THE NEED FOR X-RAY INTERVENTION. THE NEEDLE WAS REMOVED AND BOTH ENDS OF NEEDLE COMPARED WITH FULL NEEDLE TO CONFIRM THAT THERE WAS DEFINITELY NO PIECES LEFT INSIDE THE PATIENT. THERE WERE NO ADVERSE PATIENT CONSEQUENCES REPORTED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: (B)(4). A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT PHBBJDC0, AND NO NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT.
The MAUDE database houses medical device reports submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers. Each year | the FDA receives several hundred thousand medical device reports (MDRs) of suspected device-associated deaths | serious injuries and malfunctions. The FDA uses MDRs to monitor device performance | detect potential device-related safety issues | and contribute to benefit-risk assessments of these products. The MAUDE database houses MDRs submitted to the FDA by mandatory reporters (manufacturers | importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers.
Although MDRs are a valuable source of information | this passive surveillance system has limitations | including the potential submission of incomplete | inaccurate | untimely | unverified | or biased data. In addition | the incidence or prevalence of an event cannot be determined from this reporting system alone due to potential under-reporting of events and lack of information about frequency of device use. Because of this | MDRs comprise only one of the FDA's several important postmarket surveillance data sources.
MDR data alone cannot be used to establish rates of events | evaluate a change in event rates over time or compare event rates between devices. The number of reports cannot be interpreted or used in isolation to reach conclusions about the existence | severity | or frequency of problems associated with devices. Confirming whether a device actually caused a specific event can be difficult based solely on information provided in a given report. Establishing a cause-and-effect relationship is especially difficult if circumstances surrounding the event have not been verified or if the device in question has not been directly evaluated.
MAUDE data does not represent all known safety information for a reported medical device and should be interpreted in the context of other available information when making device-related or treatment decisions. Variations in trade | product | and company names affect search results. Searches only retrieve records that contain the search term(s) provided by the requester.Submission of a medical device report and the FDA's release of that information is not necessarily an admission that a product | user facility | importer | distributor | manufacturer | or medical personnel caused or contributed to the event. Certain types of report information are protected from public disclosure under the Freedom of Information Act (FOIA).
If a report contains trade secret or confidential business information | that text is replaced by "(b)(4)". If a report contains personnel or medical files information | that text is replaced by "(b)(6)". The designations "(b)(4)" and "(b)(6)" refer to the exemptions in the FOIA. For example | "(b)(4)" may be found in place of the product's composition and "(b)(6)" may be found in place of a patient's age. MAUDE is updated monthly and the search page reflects the date of the most recent update. The FDA seeks to include all reports received prior to the update but the inclusion of some reports may be delayed.
Manufacturers and importers must submit reports when they become aware of information that reasonably suggests that one of their marketed devices may have caused or contributed to a death or serious injury or has malfunctioned and the malfunction of the device or a similar device that they market would be likely to cause or contribute to a death or serious injury if the malfunction were to recur. Manufacturers must send reports of such deaths | serious injuries and malfunctions to the FDA. Importers must send reports of deaths and serious injuries to the FDA and the manufacturer | and reports of malfunctions to the manufacturer.
Device user facilities include hospitals | outpatient diagnostic or treatment facilities | nursing homes and ambulatory surgical facilities. Device user facilities must submit reports when they become aware of information that reasonably suggests that a device may have caused or contributed to a death or serious injury of a patient in their facility. Death reports must be sent to the FDA and the manufacturer | if known. Serious injury reports must be sent to the manufacturer or to the FDA | if the manufacturer is not known.